## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516913 | 25.6 | 0 | 20 | 8.512 | 0.847 | 148.17 | 247.6 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -1.2514 | 20 | 1.107 | 1.103 | 251.28 | 261.21 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.78929 | 20.2 | 1.103 | 1.098 | 261.21 | 266.24 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.47942 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516914 | 25.6 | 0 | 20 | 8.512 | 0.847 | 147.74 | 247.36 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -1.0627 | 20 | 1.107 | 1.103 | 251.01 | 260.98 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.38965 | 20.2 | 1.103 | 1.098 | 260.98 | 265.99 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.079904 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516915 | 25.6 | 0 | 20 | 8.512 | 0.847 | 148.26 | 248.14 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.80445 | 20 | 1.107 | 1.103 | 251.78 | 261.76 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.28974 | 20.2 | 1.103 | 1.098 | 261.76 | 266.83 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 1.2785 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516916 | 25.6 | 0 | 20 | 8.512 | 0.847 | 150.74 | 251.32 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.10925 | 20 | 1.107 | 1.103 | 254.98 | 264.98 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.089919 | 20.2 | 1.103 | 1.098 | 264.98 | 270.03 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.87895 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516917 | 25.6 | 0 | 20 | 8.512 | 0.847 | 148.88 | 248.64 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.92363 | 20 | 1.107 | 1.103 | 252.33 | 262.28 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.58947 | 20.2 | 1.103 | 1.098 | 262.28 | 267.26 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.51938 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
6 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516918 | 25.6 | 0 | 20 | 8.512 | 0.847 | 149.95 | 249.65 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.98322 | 20 | 1.107 | 1.103 | 253.35 | 263.35 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.089919 | 20.2 | 1.103 | 1.098 | 263.35 | 268.38 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.47942 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516919 | 25.6 | 0 | 20 | 8.512 | 0.847 | 150.7 | 251.15 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.23836 | 20 | 1.107 | 1.103 | 254.79 | 264.73 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.68938 | 20.2 | 1.103 | 1.098 | 264.73 | 269.74 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.079904 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25516920 | 25.6 | 0 | 20 | 8.512 | 0.847 | 149.1 | 248.85 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.93356 | 20 | 1.107 | 1.103 | 252.5 | 262.49 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.18983 | 20.2 | 1.103 | 1.098 | 262.49 | 267.5 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.079904 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
10 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
11 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
12 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
13 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
14 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
15 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
16 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
17 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
18 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
19 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
20 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
21 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
22 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
23 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
24 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
25 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
26 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
27 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
28 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
29 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
30 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
31 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
32 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
33 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
34 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
35 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
36 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
37 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
38 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
39 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
40 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
41 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
42 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
43 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
44 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
45 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
46 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
47 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
48 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
49 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
50 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
51 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
52 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
53 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
54 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
55 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
56 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
57 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
58 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
59 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
60 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
61 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
62 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
63 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
64 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
65 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
66 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
67 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
68 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
69 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
70 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
71 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
72 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
73 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
74 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
75 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
76 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
77 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
78 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
79 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
80 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
81 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
82 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
83 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
84 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
85 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
86 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
87 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
88 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
89 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
90 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
91 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
92 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
93 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
94 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
95 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
96 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
97 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
98 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
99 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
100 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
101 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
102 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
103 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
104 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
105 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
106 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
107 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
108 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
109 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
110 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
111 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
112 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
113 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
114 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
115 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
116 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
117 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
118 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
119 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
120 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
121 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
122 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
123 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
124 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
125 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
126 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
127 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
128 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
129 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
130 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
131 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
132 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
133 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
134 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
135 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
136 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
137 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
138 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
139 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
140 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
141 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
142 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
143 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
144 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
145 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
146 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
147 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
148 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
149 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
150 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
151 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
152 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
153 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
154 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
155 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
156 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
157 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
158 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
159 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
160 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
161 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
162 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
163 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
164 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
165 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
166 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
167 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
168 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
169 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
170 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
171 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
172 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
173 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
174 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
175 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
176 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
177 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
178 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
179 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
180 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
181 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
182 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
183 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
184 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
185 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
186 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
187 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
188 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
189 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
190 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
191 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
192 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
193 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
194 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
195 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
196 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
197 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
198 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
199 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
200 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
201 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
202 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
203 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
204 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
205 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
206 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
207 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
208 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
209 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
210 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
211 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
212 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
213 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
214 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
215 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
216 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
217 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
218 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
219 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
220 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
221 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
222 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
223 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
224 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
225 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
226 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
227 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
228 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
229 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
230 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
231 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
232 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
233 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
234 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
235 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
236 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
237 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
238 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
239 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
240 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
241 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
242 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
243 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
244 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
245 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
246 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
247 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
248 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
249 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
250 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
251 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
252 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
253 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
254 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
255 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
256 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
257 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
258 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
259 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
260 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
261 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
262 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
263 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
264 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
265 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
266 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
267 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
268 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
269 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
270 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
271 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
272 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
273 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
274 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
275 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
276 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
277 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
278 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
279 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
280 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
281 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
282 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
283 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
284 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
285 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
286 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
287 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
288 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
289 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
290 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
291 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
292 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
293 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
294 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
295 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
296 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
297 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
298 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
299 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
300 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20 | 1.107 | 1.103 |  |  | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -100 | 20.2 | 1.103 | 1.098 | 0 |  | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 3 | PA25516913 | 0 | -1.2514 | -0.78929 | 0.47942 | CONFORME | Medición
2 | 3 | PA25516914 | 0 | -1.0627 | -0.38965 | 0.079904 | CONFORME | Medición
3 | 3 | PA25516915 | 0 | -0.80445 | -0.28974 | 1.2785 | CONFORME | Medición
4 | 3 | PA25516916 | 0 | -0.10925 | -0.089919 | 0.87895 | CONFORME | Medición
5 | 3 | PA25516917 | 0 | -0.92363 | -0.58947 | -0.51938 | CONFORME | Medición
6 | 3 | PA25516918 | 0 | -0.98322 | -0.089919 | 0.47942 | CONFORME | Medición
7 | 3 | PA25516919 | 0 | -0.23836 | -0.68938 | 0.079904 | CONFORME | Medición
8 | 3 | PA25516920 | 0 | -0.93356 | -0.18983 | 0.079904 | CONFORME | Medición
9 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
10 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
11 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
12 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
13 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
14 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
15 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
16 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
17 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
18 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
19 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
20 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
21 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
22 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
23 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
24 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
25 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
26 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
27 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
28 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
29 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
30 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
31 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
32 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
33 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
34 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
35 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
36 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
37 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
38 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
39 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
40 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
41 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
42 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
43 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
44 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
45 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
46 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
47 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
48 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
49 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
50 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
51 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
52 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
53 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
54 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
55 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
56 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
57 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
58 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
59 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
60 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
61 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
62 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
63 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
64 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
65 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
66 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
67 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
68 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
69 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
70 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
71 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
72 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
73 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
74 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
75 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
76 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
77 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
78 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
79 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
80 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
81 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
82 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
83 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
84 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
85 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
86 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
87 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
88 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
89 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
90 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
91 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
92 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
93 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
94 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
95 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
96 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
97 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
98 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
99 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
100 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
101 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
102 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
103 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
104 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
105 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
106 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
107 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
108 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
109 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
110 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
111 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
112 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
113 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
114 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
115 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
116 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
117 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
118 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
119 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
120 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
121 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
122 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
123 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
124 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
125 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
126 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
127 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
128 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
129 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
130 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
131 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
132 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
133 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
134 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
135 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
136 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
137 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
138 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
139 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
140 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
141 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
142 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
143 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
144 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
145 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
146 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
147 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
148 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
149 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
150 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
151 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
152 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
153 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
154 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
155 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
156 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
157 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
158 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
159 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
160 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
161 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
162 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
163 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
164 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
165 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
166 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
167 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
168 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
169 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
170 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
171 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
172 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
173 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
174 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
175 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
176 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
177 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
178 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
179 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
180 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
181 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
182 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
183 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
184 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
185 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
186 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
187 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
188 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
189 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
190 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
191 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
192 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
193 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
194 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
195 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
196 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
197 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
198 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
199 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
200 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
201 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
202 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
203 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
204 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
205 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
206 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
207 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
208 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
209 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
210 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
211 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
212 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
213 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
214 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
215 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
216 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
217 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
218 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
219 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
220 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
221 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
222 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
223 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
224 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
225 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
226 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
227 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
228 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
229 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
230 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
231 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
232 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
233 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
234 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
235 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
236 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
237 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
238 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
239 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
240 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
241 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
242 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
243 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
244 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
245 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
246 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
247 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
248 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
249 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
250 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
251 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
252 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
253 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
254 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
255 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
256 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
257 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
258 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
259 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
260 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
261 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
262 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
263 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
264 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
265 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
266 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
267 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
268 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
269 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
270 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
271 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
272 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
273 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
274 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
275 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
276 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
277 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
278 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
279 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
280 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
281 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
282 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
283 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
284 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
285 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
286 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
287 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
288 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
289 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
290 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
291 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
292 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
293 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
294 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
295 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
296 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
297 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
298 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
299 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
300 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición