## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517041 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 244.95 | 345.19 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.62457 | 20.8 | 1.093 | 1.087 | 351 | 360.99 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.2297 | 20.9 | 1.099 | 1.096 | 360.99 | 366.02 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517042 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 249.72 | 350.17 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.41638 | 20.8 | 1.093 | 1.087 | 355.99 | 365.95 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.52931 | 20.9 | 1.099 | 1.096 | 365.95 | 370.93 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
3 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517043 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 249.14 | 349.02 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.98146 | 20.8 | 1.093 | 1.087 | 354.84 | 364.83 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.2297 | 20.9 | 1.099 | 1.096 | 364.83 | 369.86 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517044 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 252.3 | 353.44 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | 0.26767 | 20.8 | 1.093 | 1.087 | 359.3 | 369.33 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | 0.16978 | 20.9 | 1.099 | 1.096 | 369.33 | 374.4 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
5 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517045 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 251.21 | 351.11 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.96163 | 20.8 | 1.093 | 1.087 | 356.9 | 366.87 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.42944 | 20.9 | 1.099 | 1.096 | 366.87 | 371.91 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517046 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 246.86 | 347.19 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.53534 | 20.8 | 1.093 | 1.087 | 353 | 362.98 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.32957 | 20.9 | 1.099 | 1.096 | 362.98 | 367.98 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
7 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517047 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 251.26 | 351.26 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.8625 | 20.8 | 1.093 | 1.087 | 357.06 | 367.04 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.32957 | 20.9 | 1.099 | 1.096 | 367.04 | 372.02 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
8 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517048 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 247.86 | 348.27 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.45603 | 20.8 | 1.093 | 1.087 | 354.1 | 364.05 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.62918 | 20.9 | 1.099 | 1.096 | 364.05 | 369.05 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
9 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517049 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 247.31 | 347.59 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.58491 | 20.8 | 1.093 | 1.087 | 353.44 | 363.44 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.12983 | 20.9 | 1.099 | 1.096 | 363.44 | 368.45 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517050 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 245.98 | 345.95 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.89224 | 20.8 | 1.093 | 1.087 | 351.75 | 361.74 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.2297 | 20.9 | 1.099 | 1.096 | 361.74 | 366.76 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517051 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 248.16 | 348.66 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.36681 | 20.8 | 1.093 | 1.087 | 354.51 | 364.49 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.32957 | 20.9 | 1.099 | 1.096 | 364.49 | 369.51 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517052 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 245.95 | 345.8 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -1.0112 | 20.8 | 1.093 | 1.087 | 351.63 | 361.6 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.42944 | 20.9 | 1.099 | 1.096 | 361.6 | 366.59 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
13 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517053 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 278.78 | 379.26 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.38664 | 20.8 | 1.093 | 1.087 | 385.11 | 395.12 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.029961 | 20.9 | 1.099 | 1.096 | 395.12 | 400.18 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517054 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 247.89 | 348.31 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.44612 | 20.8 | 1.093 | 1.087 | 354.13 | 364.13 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.12983 | 20.9 | 1.099 | 1.096 | 364.13 | 369.18 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517055 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 249.91 | 350.03 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | -0.74353 | 20.8 | 1.093 | 1.087 | 355.84 | 365.85 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | -0.029961 | 20.9 | 1.099 | 1.096 | 365.85 | 370.93 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
16 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517056 | 25.6 | 0 | 20.8 | 8.545 | 0.626 | 250.36 | 351.3 | 100.87 | 2,29,550 | 2 | 29,550 | 0.041542 | 2428.1645 | 0.069396 | 20.8 | 1.093 | 1.087 | 357.15 | 367.21 | 10.013 | 22,41,530 | 22 | 41,530 | 0.3782 | 26.4752 | 0.46939 | 20.9 | 1.099 | 1.096 | 367.21 | 372.29 | 5.01 | 18,15,550 | 18 | 15,550 | 0.30432 | 16.463 | 1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
17 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517057 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 129.44 | 230 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.25789 | 20.9 | 1.091 | 1.087 | 233.31 | 243.32 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.089829 | 20.9 | 1.095 | 1.095 | 243.32 | 248.35 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
18 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517058 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 131.15 | 231.09 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.87284 | 20.9 | 1.091 | 1.087 | 234.4 | 244.37 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.48907 | 20.9 | 1.095 | 1.095 | 244.37 | 249.36 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | -0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
19 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517059 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 131.21 | 232.2 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | 0.16862 | 20.9 | 1.091 | 1.087 | 235.52 | 245.55 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.10979 | 20.9 | 1.095 | 1.095 | 245.55 | 250.56 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | -1 | 11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517060 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 129.3 | 229.38 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.73398 | 20.9 | 1.091 | 1.087 | 232.65 | 242.62 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.48907 | 20.9 | 1.095 | 1.095 | 242.62 | 247.64 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | -3.5386e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
21 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517061 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 130.05 | 230.49 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.37691 | 20.9 | 1.091 | 1.087 | 233.78 | 243.82 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.2096 | 20.9 | 1.095 | 1.095 | 243.82 | 248.86 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517062 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 133.19 | 234.32 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | 0.30748 | 20.9 | 1.091 | 1.087 | 237.64 | 247.66 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.009981 | 20.9 | 1.095 | 1.095 | 247.66 | 252.7 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
23 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517063 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 132.19 | 232.77 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.23805 | 20.9 | 1.091 | 1.087 | 236.11 | 246.21 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.80846 | 20.9 | 1.095 | 1.095 | 246.21 | 251.29 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517064 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 134.04 | 234.51 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.34715 | 20.9 | 1.091 | 1.087 | 237.8 | 247.81 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.089829 | 20.9 | 1.095 | 1.095 | 247.81 | 252.84 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517065 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 131.81 | 232.5 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.12894 | 20.9 | 1.091 | 1.087 | 235.8 | 245.82 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.009981 | 20.9 | 1.095 | 1.095 | 245.82 | 250.85 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517066 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 125.11 | 224.88 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -1.0415 | 20.9 | 1.091 | 1.087 | 228.2 | 238.19 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.28945 | 20.9 | 1.095 | 1.095 | 238.19 | 243.2 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
27 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517067 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 124.78 | 225.4 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.19837 | 20.9 | 1.091 | 1.087 | 228.69 | 238.74 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.30941 | 20.9 | 1.095 | 1.095 | 238.74 | 243.79 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517068 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 129.31 | 228.79 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -1.3291 | 20.9 | 1.091 | 1.087 | 232.11 | 242.07 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.58888 | 20.9 | 1.095 | 1.095 | 242.07 | 247.09 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 2.1231e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517069 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 125.85 | 226.55 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.11902 | 20.9 | 1.091 | 1.087 | 229.85 | 239.85 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.18964 | 20.9 | 1.095 | 1.095 | 239.85 | 244.87 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 2.1231e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517070 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 129.89 | 230.29 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.41658 | 20.9 | 1.091 | 1.087 | 233.58 | 243.6 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.009981 | 20.9 | 1.095 | 1.095 | 243.6 | 248.64 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517071 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 131.98 | 232.32 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | -0.4761 | 20.9 | 1.091 | 1.087 | 235.6 | 245.59 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | -0.28945 | 20.9 | 1.095 | 1.095 | 245.59 | 250.6 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
32 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517072 | 25.6 | 0 | 20.9 | 8.552 | 0.682 | 130.07 | 231.01 | 100.82 | 2,29,930 | 2 | 29,930 | 0.041647 | 2420.8097 | 0.11902 | 20.9 | 1.091 | 1.087 | 234.35 | 244.44 | 10.019 | 22,39,750 | 22 | 39,750 | 0.37771 | 26.5258 | 0.70865 | 20.9 | 1.095 | 1.095 | 244.44 | 249.5 | 5.02 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4569 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
33 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517073 | 25.6 | 0 | 21 | 8.55 | 0.676 | 158.59 | 258.95 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.43651 | 20.9 | 1.09 | 1.08 | 261.31 | 271.34 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | 0.17978 | 21.5 | 1.094 | 1.086 | 271.34 | 276.32 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
34 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517074 | 25.6 | 0 | 21 | 8.55 | 0.676 | 157.3 | 257.36 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.73413 | 20.9 | 1.09 | 1.08 | 259.67 | 269.69 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | 0.079904 | 21.5 | 1.094 | 1.086 | 269.69 | 274.75 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
35 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517075 | 25.6 | 0 | 21 | 8.55 | 0.676 | 161.57 | 261.91 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.45635 | 20.9 | 1.09 | 1.08 | 264.25 | 274.23 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.31962 | 21.5 | 1.094 | 1.086 | 274.23 | 279.22 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
36 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517076 | 25.6 | 0 | 21 | 8.55 | 0.676 | 159.91 | 260.24 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.46627 | 20.9 | 1.09 | 1.08 | 262.55 | 272.57 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | 0.079904 | 21.5 | 1.094 | 1.086 | 272.57 | 277.6 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517077 | 25.6 | 0 | 21 | 8.55 | 0.676 | 159.6 | 259.92 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.47619 | 20.9 | 1.09 | 1.08 | 262.26 | 272.24 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.31962 | 21.5 | 1.094 | 1.086 | 272.24 | 277.22 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
38 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517078 | 25.6 | 0 | 21 | 8.55 | 0.676 | 160.28 | 259.63 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -1.4385 | 20.9 | 1.09 | 1.08 | 261.8 | 271.74 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.71914 | 21.5 | 1.094 | 1.086 | 271.74 | 276.7 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
39 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517079 | 25.6 | 0 | 21 | 8.55 | 0.676 | 162.01 | 261.95 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.85317 | 20.9 | 1.09 | 1.08 | 264.26 | 274.23 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.4195 | 21.5 | 1.094 | 1.086 | 274.23 | 279.2 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
40 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517080 | 25.6 | 0 | 21 | 8.55 | 0.676 | 159.86 | 260.05 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.60516 | 20.9 | 1.09 | 1.08 | 262.4 | 272.35 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.61926 | 21.5 | 1.094 | 1.086 | 272.35 | 277.35 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
41 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517081 | 25.6 | 0 | 21 | 8.55 | 0.676 | 158.06 | 258.25 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.60516 | 20.9 | 1.09 | 1.08 | 260.54 | 270.5 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.51938 | 21.5 | 1.094 | 1.086 | 270.5 | 275.45 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
42 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517082 | 25.6 | 0 | 21 | 8.55 | 0.676 | 161.31 | 260.82 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -1.2798 | 20.9 | 1.09 | 1.08 | 263.15 | 273.09 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.71914 | 21.5 | 1.094 | 1.086 | 273.09 | 278.04 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
43 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517083 | 25.6 | 0 | 21 | 8.55 | 0.676 | 159.4 | 259.55 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.64484 | 20.9 | 1.09 | 1.08 | 261.88 | 271.83 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.61926 | 21.5 | 1.094 | 1.086 | 271.83 | 276.86 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
44 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517084 | 25.6 | 0 | 21 | 8.55 | 0.676 | 163.49 | 263.62 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.66468 | 20.9 | 1.09 | 1.08 | 265.95 | 275.9 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.61926 | 21.5 | 1.094 | 1.086 | 275.9 | 280.9 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
45 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517085 | 25.6 | 0 | 21 | 8.55 | 0.676 | 161.33 | 261.45 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.6746 | 20.9 | 1.09 | 1.08 | 263.75 | 273.7 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.61926 | 21.5 | 1.094 | 1.086 | 273.7 | 278.66 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
46 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517086 | 25.6 | 0 | 21 | 8.55 | 0.676 | 159.24 | 259.3 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.73413 | 20.9 | 1.09 | 1.08 | 261.64 | 271.6 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.51938 | 21.5 | 1.094 | 1.086 | 271.6 | 276.6 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
47 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517087 | 25.6 | 0 | 21 | 8.55 | 0.676 | 161.77 | 261.81 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -0.75397 | 20.9 | 1.09 | 1.08 | 264.14 | 274.11 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.4195 | 21.5 | 1.094 | 1.086 | 274.11 | 279.06 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
48 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517088 | 25.6 | 0 | 21 | 8.55 | 0.676 | 163.24 | 263 | 100.8 | 2,29,870 | 2 | 29,870 | 0.041631 | 2421.2985 | -1.0317 | 20.9 | 1.09 | 1.08 | 265.31 | 275.27 | 10.012 | 22,40,470 | 22 | 40,470 | 0.37791 | 26.4932 | -0.51938 | 21.5 | 1.094 | 1.086 | 275.27 | 280.29 | 5.01 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4696 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517089 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 145.75 | 245.7 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.79404 | 21.1 | 1.091 | 1.08 | 248.8 | 258.75 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.60933 | 21.6 | 1.098 | 1.087 | 258.75 | 263.77 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | 0.2196 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
50 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517090 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 142.25 | 241.95 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -1.0422 | 21.1 | 1.091 | 1.08 | 245.06 | 255.02 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.50944 | 21.6 | 1.098 | 1.087 | 255.02 | 260 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.57896 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
51 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517091 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 144.58 | 245 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.32754 | 21.1 | 1.091 | 1.08 | 248.1 | 258.1 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.10988 | 21.6 | 1.098 | 1.087 | 258.1 | 263.15 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | 0.81853 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
52 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517092 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 142.44 | 242.55 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.63524 | 21.1 | 1.091 | 1.08 | 245.64 | 255.59 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.60933 | 21.6 | 1.098 | 1.087 | 255.59 | 260.57 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.57896 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
53 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517093 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 145.19 | 244.94 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.99256 | 21.1 | 1.091 | 1.08 | 248.05 | 258 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.60933 | 21.6 | 1.098 | 1.087 | 258 | 262.99 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
54 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517094 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 144.98 | 245.34 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.3871 | 21.1 | 1.091 | 1.08 | 248.44 | 258.41 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.40955 | 21.6 | 1.098 | 1.087 | 258.41 | 263.43 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | 0.2196 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517095 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 141.59 | 241.66 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.67494 | 21.1 | 1.091 | 1.08 | 244.75 | 254.74 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.20977 | 21.6 | 1.098 | 1.087 | 254.74 | 259.73 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
56 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517096 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 143.65 | 244.25 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.14888 | 21.1 | 1.091 | 1.08 | 247.35 | 257.35 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.10988 | 21.6 | 1.098 | 1.087 | 257.35 | 262.35 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.17968 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
57 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517097 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 144.16 | 244.2 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.70471 | 21.1 | 1.091 | 1.08 | 247.27 | 257.23 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.50944 | 21.6 | 1.098 | 1.087 | 257.23 | 262.22 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
58 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517098 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 145.45 | 246.1 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.099256 | 21.1 | 1.091 | 1.08 | 249.21 | 259.2 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.20977 | 21.6 | 1.098 | 1.087 | 259.2 | 264.2 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.17968 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
59 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517099 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 147.4 | 247.81 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.33747 | 21.1 | 1.091 | 1.08 | 250.93 | 260.9 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.40955 | 21.6 | 1.098 | 1.087 | 260.9 | 265.89 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
60 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517100 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 144.36 | 244.7 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.40695 | 21.1 | 1.091 | 1.08 | 247.8 | 257.76 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.50944 | 21.6 | 1.098 | 1.087 | 257.76 | 262.73 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.7786 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
61 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517101 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 144 | 243.83 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.91315 | 21.1 | 1.091 | 1.08 | 246.91 | 256.89 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.30966 | 21.6 | 1.098 | 1.087 | 256.89 | 261.85 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.97824 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
62 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517102 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 141.77 | 241.84 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.67494 | 21.1 | 1.091 | 1.08 | 244.95 | 254.9 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.60933 | 21.6 | 1.098 | 1.087 | 254.9 | 259.92 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | 0.2196 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
63 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517103 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 147.09 | 247.15 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.68486 | 21.1 | 1.091 | 1.08 | 250.24 | 260.22 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | -0.30966 | 21.6 | 1.098 | 1.087 | 260.22 | 265.18 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | -0.97824 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
64 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517104 | 25.6 | 0 | 21.1 | 8.547 | 0.637 | 143.36 | 243.99 | 100.75 | 2,29,750 | 2 | 29,750 | 0.041597 | 2422.0367 | -0.11911 | 21.1 | 1.091 | 1.08 | 247.1 | 257.12 | 10.011 | 22,39,000 | 22 | 39,000 | 0.3775 | 26.5192 | 0.089901 | 21.6 | 1.098 | 1.087 | 257.12 | 262.15 | 5.009 | 18,14,560 | 18 | 14,560 | 0.30404 | 16.4746 | 0.41925 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
65 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517105 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 151.64 | 257.15 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.91097 | 21.1 | 1.096 | 1.091 | 261.09 | 271.08 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.23966 | 21.4 | 1.098 | 1.094 | 271.08 | 276.05 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.8182 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
66 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517106 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 151.75 | 257.24 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.92975 | 21.1 | 1.096 | 1.091 | 261.15 | 271.17 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | 0.059916 | 21.4 | 1.098 | 1.094 | 271.17 | 276.24 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | 1.1774 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
67 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517107 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 150.29 | 256.26 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.47896 | 21.1 | 1.096 | 1.091 | 260.2 | 270.2 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.1398 | 21.4 | 1.098 | 1.094 | 270.2 | 275.23 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | 0.37917 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
68 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517108 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 152.49 | 258.49 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.45079 | 21.1 | 1.096 | 1.091 | 262.41 | 272.38 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.43938 | 21.4 | 1.098 | 1.094 | 272.38 | 277.35 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.8182 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
69 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517109 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 150.95 | 256.85 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.5447 | 21.1 | 1.096 | 1.091 | 260.78 | 270.74 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.53925 | 21.4 | 1.098 | 1.094 | 270.74 | 275.74 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
70 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517110 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 150.15 | 255.55 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -1.0143 | 21.1 | 1.096 | 1.091 | 259.49 | 269.44 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.63911 | 21.4 | 1.098 | 1.094 | 269.44 | 274.39 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -1.2173 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
71 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517111 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 149.65 | 255.55 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.5447 | 21.1 | 1.096 | 1.091 | 259.5 | 269.49 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.23966 | 21.4 | 1.098 | 1.094 | 269.49 | 274.49 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
72 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517112 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 151.14 | 257.29 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.30992 | 21.1 | 1.096 | 1.091 | 261.21 | 271.18 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.43938 | 21.4 | 1.098 | 1.094 | 271.18 | 276.18 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
73 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517113 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 150.75 | 256.56 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.62923 | 21.1 | 1.096 | 1.091 | 260.51 | 270.49 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.33952 | 21.4 | 1.098 | 1.094 | 270.49 | 275.5 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
74 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517114 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 152.79 | 258.56 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.66679 | 21.1 | 1.096 | 1.091 | 262.51 | 272.51 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.1398 | 21.4 | 1.098 | 1.094 | 272.51 | 277.58 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | 1.1774 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
75 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517115 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 153.25 | 259.66 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.06574 | 21.1 | 1.096 | 1.091 | 263.62 | 273.62 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.1398 | 21.4 | 1.098 | 1.094 | 273.62 | 278.66 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
76 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517116 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 150.45 | 256.15 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.73253 | 21.1 | 1.096 | 1.091 | 260.05 | 270.01 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.53925 | 21.4 | 1.098 | 1.094 | 270.01 | 274.99 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.61864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
77 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517117 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 149.5 | 255.6 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.35687 | 21.1 | 1.096 | 1.091 | 259.55 | 269.53 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.33952 | 21.4 | 1.098 | 1.094 | 269.53 | 274.54 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
78 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517118 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 149.81 | 255 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -1.2115 | 21.1 | 1.096 | 1.091 | 258.9 | 268.84 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.73897 | 21.4 | 1.098 | 1.094 | 268.84 | 273.8 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -1.0178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
79 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517119 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 152.59 | 258.54 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | -0.49775 | 21.1 | 1.096 | 1.091 | 262.49 | 272.45 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | -0.53925 | 21.4 | 1.098 | 1.094 | 272.45 | 277.45 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
80 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517120 | 25.6 | 0 | 21.3 | 8.544 | 0.685 | 151.24 | 257.9 | 106.48 | 2,38,310 | 2 | 38,310 | 0.043975 | 2421.3758 | 0.16905 | 21.1 | 1.096 | 1.091 | 261.86 | 271.95 | 10.014 | 22,40,330 | 22 | 40,330 | 0.37787 | 26.5012 | 0.75894 | 21.4 | 1.098 | 1.094 | 271.95 | 277.01 | 5.011 | 18,17,550 | 18 | 17,550 | 0.30488 | 16.4362 | 0.97785 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
81 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517121 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 172.62 | 272.51 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.81422 | 21.3 | 1.087 | 1.079 | 275.38 | 285.37 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.21974 | 22.1 | 1.09 | 1.073 | 285.37 | 290.35 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.61864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
82 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517122 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 173.81 | 274.39 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.12908 | 21.3 | 1.087 | 1.079 | 277.26 | 287.22 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.51938 | 22.1 | 1.09 | 1.073 | 287.22 | 292.18 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -1.0178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
83 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517123 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 173.4 | 273.31 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.79436 | 21.3 | 1.087 | 1.079 | 276.15 | 286.14 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.21974 | 22.1 | 1.09 | 1.073 | 286.14 | 291.15 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
84 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517124 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 171.89 | 271.6 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.99295 | 21.3 | 1.087 | 1.079 | 274.46 | 284.4 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.71914 | 22.1 | 1.09 | 1.073 | 284.4 | 289.35 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -1.2173 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
85 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517125 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 173.49 | 273.41 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.78443 | 21.3 | 1.087 | 1.079 | 276.25 | 286.22 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.4195 | 22.1 | 1.09 | 1.073 | 286.22 | 291.22 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
86 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517126 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 170.85 | 270.26 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -1.2908 | 21.3 | 1.087 | 1.079 | 273.09 | 283.04 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.61926 | 22.1 | 1.09 | 1.073 | 283.04 | 288.01 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.8182 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
87 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517127 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 172.09 | 271.56 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -1.2313 | 21.3 | 1.087 | 1.079 | 274.43 | 284.39 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.51938 | 22.1 | 1.09 | 1.073 | 284.39 | 289.32 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -1.6164 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
88 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517128 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 171.3 | 270.49 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -1.5093 | 21.3 | 1.087 | 1.079 | 273.34 | 283.32 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.31962 | 22.1 | 1.09 | 1.073 | 283.32 | 288.34 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | 0.1796 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
89 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517129 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 174.64 | 274.45 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.89366 | 21.3 | 1.087 | 1.079 | 277.29 | 287.28 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.21974 | 22.1 | 1.09 | 1.073 | 287.28 | 292.25 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.8182 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
90 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517130 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 172.61 | 273.06 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.25817 | 21.3 | 1.087 | 1.079 | 275.9 | 285.91 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.019976 | 22.1 | 1.09 | 1.073 | 285.91 | 290.95 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
91 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517131 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 176.65 | 276.95 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.40711 | 21.3 | 1.087 | 1.079 | 279.8 | 289.85 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | 0.37954 | 22.1 | 1.09 | 1.073 | 289.85 | 294.86 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
92 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517132 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 177.4 | 278.34 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | 0.22838 | 21.3 | 1.087 | 1.079 | 281.19 | 291.21 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | 0.079904 | 22.1 | 1.09 | 1.073 | 291.21 | 296.25 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
93 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517133 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 175.31 | 275.45 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.56598 | 21.3 | 1.087 | 1.079 | 278.29 | 288.29 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.11986 | 22.1 | 1.09 | 1.073 | 288.29 | 293.3 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
94 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517134 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 175.92 | 276.3 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.32767 | 21.3 | 1.087 | 1.079 | 279.14 | 289.12 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | -0.31962 | 22.1 | 1.09 | 1.073 | 289.12 | 294.11 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
95 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517135 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 177.49 | 278.19 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.0099295 | 21.3 | 1.087 | 1.079 | 281.04 | 291.09 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | 0.37954 | 22.1 | 1.09 | 1.073 | 291.09 | 296.15 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | 0.97785 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
96 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517136 | 25.6 | 0 | 21.3 | 8.558 | 0.687 | 175.41 | 275.98 | 100.71 | 2,29,750 | 2 | 29,750 | 0.041597 | 2421.0751 | -0.13901 | 21.3 | 1.087 | 1.079 | 278.84 | 288.92 | 10.012 | 22,41,360 | 22 | 41,360 | 0.37816 | 26.4759 | 0.67918 | 22.1 | 1.09 | 1.073 | 288.92 | 293.98 | 5.011 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4586 | 0.97785 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
97 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517137 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 148.06 | 248.49 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.36706 | 21.4 | 1.093 | 1.083 | 250.8 | 260.76 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.51938 | 21.4 | 1.09 | 1.093 | 260.76 | 265.75 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
98 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517138 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 146.31 | 245.98 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -1.121 | 21.4 | 1.093 | 1.083 | 248.15 | 258.12 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.4195 | 21.4 | 1.09 | 1.093 | 258.12 | 263.09 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
99 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517139 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 146.62 | 246.5 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.9127 | 21.4 | 1.093 | 1.083 | 248.76 | 258.76 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.11986 | 21.4 | 1.09 | 1.093 | 258.76 | 263.69 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -1.5968 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
100 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517140 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 145.94 | 245.5 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -1.2302 | 21.4 | 1.093 | 1.083 | 247.78 | 257.74 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.51938 | 21.4 | 1.09 | 1.093 | 257.74 | 262.71 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
101 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517141 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 147.91 | 248.3 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.40675 | 21.4 | 1.093 | 1.083 | 250.59 | 260.55 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.51938 | 21.4 | 1.09 | 1.093 | 260.55 | 265.54 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
102 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517142 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 142.59 | 242.59 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.79365 | 21.4 | 1.093 | 1.083 | 244.87 | 254.82 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.61926 | 21.4 | 1.09 | 1.093 | 254.82 | 259.79 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
103 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517143 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 143.74 | 244.05 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.48611 | 21.4 | 1.093 | 1.083 | 246.39 | 256.4 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.019976 | 21.4 | 1.09 | 1.093 | 256.4 | 261.44 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
104 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517144 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 142.96 | 242.54 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -1.2103 | 21.4 | 1.093 | 1.083 | 244.74 | 254.76 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | 0.079904 | 21.4 | 1.09 | 1.093 | 254.76 | 259.68 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -1.7964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
105 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517145 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 142.54 | 242.45 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.88294 | 21.4 | 1.093 | 1.083 | 244.74 | 254.72 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.31962 | 21.4 | 1.09 | 1.093 | 254.72 | 259.71 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
106 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517146 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 147.56 | 247.76 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.59524 | 21.4 | 1.093 | 1.083 | 250.1 | 260.05 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.61926 | 21.4 | 1.09 | 1.093 | 260.05 | 265.05 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
107 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517147 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 148.8 | 249 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.59524 | 21.4 | 1.093 | 1.083 | 251.31 | 261.27 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.51938 | 21.4 | 1.09 | 1.093 | 261.27 | 266.22 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
108 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517148 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 144.96 | 245.63 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.12897 | 21.4 | 1.093 | 1.083 | 247.95 | 257.93 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.31962 | 21.4 | 1.09 | 1.093 | 257.93 | 262.95 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
109 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517149 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 144.79 | 245.19 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.39683 | 21.4 | 1.093 | 1.083 | 247.49 | 257.46 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.4195 | 21.4 | 1.09 | 1.093 | 257.46 | 262.49 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
110 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517150 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 147.66 | 248 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.45635 | 21.4 | 1.093 | 1.083 | 250.31 | 260.27 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.51938 | 21.4 | 1.09 | 1.093 | 260.27 | 265.29 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
111 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517151 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 148.86 | 249.05 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.60516 | 21.4 | 1.093 | 1.083 | 251.36 | 261.29 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.81902 | 21.4 | 1.09 | 1.093 | 261.29 | 266.3 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
112 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517152 | 25.6 | 0 | 21.3 | 8.556 | 0.656 | 146.1 | 246.36 | 100.8 | 2,30,050 | 2 | 30,050 | 0.041681 | 2418.3939 | -0.53571 | 21.4 | 1.093 | 1.083 | 248.65 | 258.65 | 10.012 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5022 | -0.11986 | 21.4 | 1.09 | 1.093 | 258.65 | 263.73 | 5.01 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.4846 | 1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
113 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517153 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.11 | 238 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.90278 | 21.5 | 1.091 | 1.081 | 240.66 | 250.6 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.7786 | 21.5 | 1.094 | 1.092 | 250.6 | 255.56 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -1.0178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
114 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517154 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 140.11 | 239.81 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -1.0913 | 21.5 | 1.091 | 1.081 | 242.53 | 252.49 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.57896 | 21.5 | 1.094 | 1.092 | 252.49 | 257.49 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
115 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517155 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 136.75 | 236.86 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.68452 | 21.5 | 1.091 | 1.081 | 239.6 | 249.57 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.47914 | 21.5 | 1.094 | 1.092 | 249.57 | 254.57 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
116 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517156 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.09 | 238.24 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.64484 | 21.5 | 1.091 | 1.081 | 240.96 | 250.94 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.37932 | 21.5 | 1.094 | 1.092 | 250.94 | 255.96 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | 0.1796 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
117 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517157 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.91 | 238.7 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -1.002 | 21.5 | 1.091 | 1.081 | 241.45 | 251.4 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.67878 | 21.5 | 1.094 | 1.092 | 251.4 | 256.41 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
118 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517158 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 142.29 | 242.31 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.77381 | 21.5 | 1.091 | 1.081 | 245 | 254.56 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -4.5718 | 21.5 | 1.094 | 1.092 | 254.56 | 258.91 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -13.191 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
119 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517159 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 135.84 | 235.58 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -1.0516 | 21.5 | 1.091 | 1.081 | 238.32 | 248.3 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.37932 | 21.5 | 1.094 | 1.092 | 248.3 | 253.2 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -2.2151 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
120 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517160 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.61 | 238.39 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -1.0119 | 21.5 | 1.091 | 1.081 | 241.09 | 250.99 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -1.1779 | 21.5 | 1.094 | 1.092 | 250.99 | 255.94 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -1.2173 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
121 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517161 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 136.82 | 236.64 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.97222 | 21.5 | 1.091 | 1.081 | 239.39 | 249.34 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.67878 | 21.5 | 1.094 | 1.092 | 249.34 | 254.3 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -1.0178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
122 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517162 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.96 | 238.91 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.84325 | 21.5 | 1.091 | 1.081 | 241.64 | 251.62 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.37932 | 21.5 | 1.094 | 1.092 | 251.62 | 256.66 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
123 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517163 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 139.9 | 240.59 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.10913 | 21.5 | 1.091 | 1.081 | 243.36 | 253.35 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.2795 | 21.5 | 1.094 | 1.092 | 253.35 | 258.39 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
124 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517164 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.9 | 240.01 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | 0.30754 | 21.5 | 1.091 | 1.081 | 242.74 | 252.72 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.37932 | 21.5 | 1.094 | 1.092 | 252.72 | 257.75 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | 0.37917 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
125 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517165 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 134.05 | 234.2 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.64484 | 21.5 | 1.091 | 1.081 | 236.88 | 246.86 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.37932 | 21.5 | 1.094 | 1.092 | 246.86 | 251.86 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
126 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517166 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 138.75 | 239.15 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.39683 | 21.5 | 1.091 | 1.081 | 241.87 | 251.83 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.57896 | 21.5 | 1.094 | 1.092 | 251.83 | 256.81 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -0.61864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
127 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517167 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 142.05 | 242.39 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -0.45635 | 21.5 | 1.091 | 1.081 | 245.11 | 255.08 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -0.47914 | 21.5 | 1.094 | 1.092 | 255.08 | 260.08 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
128 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517168 | 25.6 | 0 | 21.6 | 8.559 | 0.633 | 141.36 | 241.09 | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -1.0615 | 21.5 | 1.091 | 1.081 | 243.8 | 253.83 | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | 0.11978 | 21.5 | 1.094 | 1.092 | 253.83 | 258.87 | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
129 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517169 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517170 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517171 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517172 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517173 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517174 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517175 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517176 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517177 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517178 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517179 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517180 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517181 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517182 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517183 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517184 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517185 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517186 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517187 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517188 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517189 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517190 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517191 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517192 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517193 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517194 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517195 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517196 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517197 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517198 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517199 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517200 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517201 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517202 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517203 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517204 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517205 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517206 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517207 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517208 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517209 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517210 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517211 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517212 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517213 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517214 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517215 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517216 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517217 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517218 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517219 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517220 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517221 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517222 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517223 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517224 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517225 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517226 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517227 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517228 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517229 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517230 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517231 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517232 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517233 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517234 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517235 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517236 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517237 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517238 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517239 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517240 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517241 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517242 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517243 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517244 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517245 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517246 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517247 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517248 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517249 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517250 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517251 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517252 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517253 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517254 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517255 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517256 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517257 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517258 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517259 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517260 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517261 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517262 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517263 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517264 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517265 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517266 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517267 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517268 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517269 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517270 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517271 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517272 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517273 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517274 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517275 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517276 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517277 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517278 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517279 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517280 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517281 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517282 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517283 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517284 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517285 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517286 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517287 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517288 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517289 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517290 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517291 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517292 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517293 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517294 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517295 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517296 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517297 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517298 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517299 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517300 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517301 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517302 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517303 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517304 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517305 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517306 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517307 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517308 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517309 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517310 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517311 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517312 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517313 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517314 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517315 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517316 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517317 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517318 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517319 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517320 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517321 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517322 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517323 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517324 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517325 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517326 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517327 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517328 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517329 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-19 | 2025-09-19 | 3 | 9 |  | PA25517330 | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-19 | 2025-09-19 | 3 | 9 |  |  | 25.6 | 0 | 21.6 | 8.559 | 0.633 |  |  | 100.8 | 2,29,880 | 2 | 29,880 | 0.041633 | 2421.1369 | -100 | 21.5 | 1.091 | 1.081 |  |  | 10.018 | 22,39,470 | 22 | 39,470 | 0.37763 | 26.5286 | -100 | 21.5 | 1.094 | 1.092 | 0 |  | 5.011 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4294 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 3 | PA25517041 | 0 | -0.62457 | -0.2297 | 0.3992 | CONFORME | Medición
2 | 3 | PA25517042 | 0 | -0.41638 | -0.52931 | -0.5988 | CONFORME | Medición
3 | 3 | PA25517043 | 0 | -0.98146 | -0.2297 | 0.3992 | CONFORME | Medición
4 | 3 | PA25517044 | 0 | 0.26767 | 0.16978 | 1.1976 | CONFORME | Medición
5 | 3 | PA25517045 | 0 | -0.96163 | -0.42944 | 0.5988 | CONFORME | Medición
6 | 3 | PA25517046 | 0 | -0.53534 | -0.32957 | -0.1996 | CONFORME | Medición
7 | 3 | PA25517047 | 0 | -0.8625 | -0.32957 | -0.5988 | CONFORME | Medición
8 | 3 | PA25517048 | 0 | -0.45603 | -0.62918 | -0.1996 | CONFORME | Medición
9 | 3 | PA25517049 | 0 | -0.58491 | -0.12983 | -1.7728e-13 | CONFORME | Medición
10 | 3 | PA25517050 | 0 | -0.89224 | -0.2297 | 0.1996 | CONFORME | Medición
11 | 3 | PA25517051 | 0 | -0.36681 | -0.32957 | 0.1996 | CONFORME | Medición
12 | 3 | PA25517052 | 0 | -1.0112 | -0.42944 | -0.3992 | CONFORME | Medición
13 | 3 | PA25517053 | 0 | -0.38664 | -0.029961 | 0.998 | CONFORME | Medición
14 | 3 | PA25517054 | 0 | -0.44612 | -0.12983 | 0.7984 | CONFORME | Medición
15 | 3 | PA25517055 | 0 | -0.74353 | -0.029961 | 1.3972 | CONFORME | Medición
16 | 3 | PA25517056 | 0 | 0.069396 | 0.46939 | 1.3972 | CONFORME | Medición
17 | 3 | PA25517057 | 0 | -0.25789 | -0.089829 | 0.1992 | CONFORME | Medición
18 | 3 | PA25517058 | 0 | -0.87284 | -0.48907 | -0.59761 | CONFORME | Medición
19 | 3 | PA25517059 | 0 | 0.16862 | 0.10979 | -0.1992 | CONFORME | Medición
20 | 3 | PA25517060 | 0 | -0.73398 | -0.48907 | -3.5386e-13 | CONFORME | Medición
21 | 3 | PA25517061 | 0 | -0.37691 | 0.2096 | 0.39841 | CONFORME | Medición
22 | 3 | PA25517062 | 0 | 0.30748 | 0.009981 | 0.39841 | CONFORME | Medición
23 | 3 | PA25517063 | 0 | -0.23805 | 0.80846 | 1.1952 | CONFORME | Medición
24 | 3 | PA25517064 | 0 | -0.34715 | -0.089829 | 0.1992 | CONFORME | Medición
25 | 3 | PA25517065 | 0 | -0.12894 | 0.009981 | 0.1992 | CONFORME | Medición
26 | 3 | PA25517066 | 0 | -1.0415 | -0.28945 | -0.1992 | CONFORME | Medición
27 | 3 | PA25517067 | 0 | -0.19837 | 0.30941 | 0.59761 | CONFORME | Medición
28 | 3 | PA25517068 | 0 | -1.3291 | -0.58888 | 2.1231e-13 | CONFORME | Medición
29 | 3 | PA25517069 | 0 | -0.11902 | -0.18964 | 2.1231e-13 | CONFORME | Medición
30 | 3 | PA25517070 | 0 | -0.41658 | 0.009981 | 0.39841 | CONFORME | Medición
31 | 3 | PA25517071 | 0 | -0.4761 | -0.28945 | -0.1992 | CONFORME | Medición
32 | 3 | PA25517072 | 0 | 0.11902 | 0.70865 | 0.79681 | CONFORME | Medición
33 | 3 | PA25517073 | 0 | -0.43651 | 0.17978 | -0.5988 | CONFORME | Medición
34 | 3 | PA25517074 | 0 | -0.73413 | 0.079904 | 0.998 | CONFORME | Medición
35 | 3 | PA25517075 | 0 | -0.45635 | -0.31962 | -0.3992 | CONFORME | Medición
36 | 3 | PA25517076 | 0 | -0.46627 | 0.079904 | 0.3992 | CONFORME | Medición
37 | 3 | PA25517077 | 0 | -0.47619 | -0.31962 | -0.5988 | CONFORME | Medición
38 | 3 | PA25517078 | 0 | -1.4385 | -0.71914 | -0.998 | CONFORME | Medición
39 | 3 | PA25517079 | 0 | -0.85317 | -0.4195 | -0.7984 | CONFORME | Medición
40 | 3 | PA25517080 | 0 | -0.60516 | -0.61926 | -0.1996 | CONFORME | Medición
41 | 3 | PA25517081 | 0 | -0.60516 | -0.51938 | -1.1976 | CONFORME | Medición
42 | 3 | PA25517082 | 0 | -1.2798 | -0.71914 | -1.1976 | CONFORME | Medición
43 | 3 | PA25517083 | 0 | -0.64484 | -0.61926 | 0.3992 | CONFORME | Medición
44 | 3 | PA25517084 | 0 | -0.66468 | -0.61926 | -0.1996 | CONFORME | Medición
45 | 3 | PA25517085 | 0 | -0.6746 | -0.61926 | -0.998 | CONFORME | Medición
46 | 3 | PA25517086 | 0 | -0.73413 | -0.51938 | -0.1996 | CONFORME | Medición
47 | 3 | PA25517087 | 0 | -0.75397 | -0.4195 | -1.1976 | CONFORME | Medición
48 | 3 | PA25517088 | 0 | -1.0317 | -0.51938 | 0.1996 | CONFORME | Medición
49 | 3 | PA25517089 | 0 | -0.79404 | -0.60933 | 0.2196 | CONFORME | Medición
50 | 3 | PA25517090 | 0 | -1.0422 | -0.50944 | -0.57896 | CONFORME | Medición
51 | 3 | PA25517091 | 0 | -0.32754 | -0.10988 | 0.81853 | CONFORME | Medición
52 | 3 | PA25517092 | 0 | -0.63524 | -0.60933 | -0.57896 | CONFORME | Medición
53 | 3 | PA25517093 | 0 | -0.99256 | -0.60933 | -0.37932 | CONFORME | Medición
54 | 3 | PA25517094 | 0 | -0.3871 | -0.40955 | 0.2196 | CONFORME | Medición
55 | 3 | PA25517095 | 0 | -0.67494 | -0.20977 | -0.37932 | CONFORME | Medición
56 | 3 | PA25517096 | 0 | -0.14888 | -0.10988 | -0.17968 | CONFORME | Medición
57 | 3 | PA25517097 | 0 | -0.70471 | -0.50944 | -0.37932 | CONFORME | Medición
58 | 3 | PA25517098 | 0 | -0.099256 | -0.20977 | -0.17968 | CONFORME | Medición
59 | 3 | PA25517099 | 0 | -0.33747 | -0.40955 | -0.37932 | CONFORME | Medición
60 | 3 | PA25517100 | 0 | -0.40695 | -0.50944 | -0.7786 | CONFORME | Medición
61 | 3 | PA25517101 | 0 | -0.91315 | -0.30966 | -0.97824 | CONFORME | Medición
62 | 3 | PA25517102 | 0 | -0.67494 | -0.60933 | 0.2196 | CONFORME | Medición
63 | 3 | PA25517103 | 0 | -0.68486 | -0.30966 | -0.97824 | CONFORME | Medición
64 | 3 | PA25517104 | 0 | -0.11911 | 0.089901 | 0.41925 | CONFORME | Medición
65 | 3 | PA25517105 | 0 | -0.91097 | -0.23966 | -0.8182 | CONFORME | Medición
66 | 3 | PA25517106 | 0 | -0.92975 | 0.059916 | 1.1774 | CONFORME | Medición
67 | 3 | PA25517107 | 0 | -0.47896 | -0.1398 | 0.37917 | CONFORME | Medición
68 | 3 | PA25517108 | 0 | -0.45079 | -0.43938 | -0.8182 | CONFORME | Medición
69 | 3 | PA25517109 | 0 | -0.5447 | -0.53925 | -0.21952 | CONFORME | Medición
70 | 3 | PA25517110 | 0 | -1.0143 | -0.63911 | -1.2173 | CONFORME | Medición
71 | 3 | PA25517111 | 0 | -0.5447 | -0.23966 | -0.21952 | CONFORME | Medición
72 | 3 | PA25517112 | 0 | -0.30992 | -0.43938 | -0.21952 | CONFORME | Medición
73 | 3 | PA25517113 | 0 | -0.62923 | -0.33952 | -0.019956 | CONFORME | Medición
74 | 3 | PA25517114 | 0 | -0.66679 | -0.1398 | 1.1774 | CONFORME | Medición
75 | 3 | PA25517115 | 0 | -0.06574 | -0.1398 | 0.57873 | CONFORME | Medición
76 | 3 | PA25517116 | 0 | -0.73253 | -0.53925 | -0.61864 | CONFORME | Medición
77 | 3 | PA25517117 | 0 | -0.35687 | -0.33952 | -0.019956 | CONFORME | Medición
78 | 3 | PA25517118 | 0 | -1.2115 | -0.73897 | -1.0178 | CONFORME | Medición
79 | 3 | PA25517119 | 0 | -0.49775 | -0.53925 | -0.21952 | CONFORME | Medición
80 | 3 | PA25517120 | 0 | 0.16905 | 0.75894 | 0.97785 | CONFORME | Medición
81 | 3 | PA25517121 | 0 | -0.81422 | -0.21974 | -0.61864 | CONFORME | Medición
82 | 3 | PA25517122 | 0 | -0.12908 | -0.51938 | -1.0178 | CONFORME | Medición
83 | 3 | PA25517123 | 0 | -0.79436 | -0.21974 | -0.019956 | CONFORME | Medición
84 | 3 | PA25517124 | 0 | -0.99295 | -0.71914 | -1.2173 | CONFORME | Medición
85 | 3 | PA25517125 | 0 | -0.78443 | -0.4195 | -0.21952 | CONFORME | Medición
86 | 3 | PA25517126 | 0 | -1.2908 | -0.61926 | -0.8182 | CONFORME | Medición
87 | 3 | PA25517127 | 0 | -1.2313 | -0.51938 | -1.6164 | CONFORME | Medición
88 | 3 | PA25517128 | 0 | -1.5093 | -0.31962 | 0.1796 | CONFORME | Medición
89 | 3 | PA25517129 | 0 | -0.89366 | -0.21974 | -0.8182 | CONFORME | Medición
90 | 3 | PA25517130 | 0 | -0.25817 | -0.019976 | 0.57873 | CONFORME | Medición
91 | 3 | PA25517131 | 0 | -0.40711 | 0.37954 | -0.019956 | CONFORME | Medición
92 | 3 | PA25517132 | 0 | 0.22838 | 0.079904 | 0.57873 | CONFORME | Medición
93 | 3 | PA25517133 | 0 | -0.56598 | -0.11986 | -0.019956 | CONFORME | Medición
94 | 3 | PA25517134 | 0 | -0.32767 | -0.31962 | -0.41908 | CONFORME | Medición
95 | 3 | PA25517135 | 0 | -0.0099295 | 0.37954 | 0.97785 | CONFORME | Medición
96 | 3 | PA25517136 | 0 | -0.13901 | 0.67918 | 0.97785 | CONFORME | Medición
97 | 3 | PA25517137 | 0 | -0.36706 | -0.51938 | -0.3992 | CONFORME | Medición
98 | 3 | PA25517138 | 0 | -1.121 | -0.4195 | -0.7984 | CONFORME | Medición
99 | 3 | PA25517139 | 0 | -0.9127 | -0.11986 | -1.5968 | CONFORME | Medición
100 | 3 | PA25517140 | 0 | -1.2302 | -0.51938 | -0.7984 | CONFORME | Medición
101 | 3 | PA25517141 | 0 | -0.40675 | -0.51938 | -0.3992 | CONFORME | Medición
102 | 3 | PA25517142 | 0 | -0.79365 | -0.61926 | -0.7984 | CONFORME | Medición
103 | 3 | PA25517143 | 0 | -0.48611 | -0.019976 | 0.5988 | CONFORME | Medición
104 | 3 | PA25517144 | 0 | -1.2103 | 0.079904 | -1.7964 | CONFORME | Medición
105 | 3 | PA25517145 | 0 | -0.88294 | -0.31962 | -0.3992 | CONFORME | Medición
106 | 3 | PA25517146 | 0 | -0.59524 | -0.61926 | -0.1996 | CONFORME | Medición
107 | 3 | PA25517147 | 0 | -0.59524 | -0.51938 | -1.1976 | CONFORME | Medición
108 | 3 | PA25517148 | 0 | -0.12897 | -0.31962 | 0.1996 | CONFORME | Medición
109 | 3 | PA25517149 | 0 | -0.39683 | -0.4195 | 0.3992 | CONFORME | Medición
110 | 3 | PA25517150 | 0 | -0.45635 | -0.51938 | 0.1996 | CONFORME | Medición
111 | 3 | PA25517151 | 0 | -0.60516 | -0.81902 | -1.7728e-13 | CONFORME | Medición
112 | 3 | PA25517152 | 0 | -0.53571 | -0.11986 | 1.3972 | CONFORME | Medición
113 | 3 | PA25517153 | 0 | -0.90278 | -0.7786 | -1.0178 | CONFORME | Medición
114 | 3 | PA25517154 | 0 | -1.0913 | -0.57896 | -0.21952 | CONFORME | Medición
115 | 3 | PA25517155 | 0 | -0.68452 | -0.47914 | -0.21952 | CONFORME | Medición
116 | 3 | PA25517156 | 0 | -0.64484 | -0.37932 | 0.1796 | CONFORME | Medición
117 | 3 | PA25517157 | 0 | -1.002 | -0.67878 | -0.019956 | CONFORME | Medición
118 | 3 | PA25517158 | 0 | -0.77381 | -4.5718 | -13.191 | NO CONFORME | Medición
119 | 3 | PA25517159 | 0 | -1.0516 | -0.37932 | -2.2151 | CONFORME | Medición
120 | 3 | PA25517160 | 0 | -1.0119 | -1.1779 | -1.2173 | CONFORME | Medición
121 | 3 | PA25517161 | 0 | -0.97222 | -0.67878 | -1.0178 | CONFORME | Medición
122 | 3 | PA25517162 | 0 | -0.84325 | -0.37932 | 0.57873 | CONFORME | Medición
123 | 3 | PA25517163 | 0 | -0.10913 | -0.2795 | 0.57873 | CONFORME | Medición
124 | 3 | PA25517164 | 0 | 0.30754 | -0.37932 | 0.37917 | CONFORME | Medición
125 | 3 | PA25517165 | 0 | -0.64484 | -0.37932 | -0.21952 | CONFORME | Medición
126 | 3 | PA25517166 | 0 | -0.39683 | -0.57896 | -0.61864 | CONFORME | Medición
127 | 3 | PA25517167 | 0 | -0.45635 | -0.47914 | -0.21952 | CONFORME | Medición
128 | 3 | PA25517168 | 0 | -1.0615 | 0.11978 | 0.57873 | CONFORME | Medición
129 | 3 | PA25517169 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 3 | PA25517170 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 3 | PA25517171 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 3 | PA25517172 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 3 | PA25517173 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 3 | PA25517174 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 3 | PA25517175 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 3 | PA25517176 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 3 | PA25517177 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 3 | PA25517178 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 3 | PA25517179 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 3 | PA25517180 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 3 | PA25517181 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 3 | PA25517182 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 3 | PA25517183 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 3 | PA25517184 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 3 | PA25517185 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 3 | PA25517186 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 3 | PA25517187 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 3 | PA25517188 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 3 | PA25517189 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 3 | PA25517190 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 3 | PA25517191 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 3 | PA25517192 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 3 | PA25517193 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 3 | PA25517194 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 3 | PA25517195 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 3 | PA25517196 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 3 | PA25517197 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 3 | PA25517198 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 3 | PA25517199 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 3 | PA25517200 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 3 | PA25517201 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 3 | PA25517202 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 3 | PA25517203 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 3 | PA25517204 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 3 | PA25517205 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 3 | PA25517206 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 3 | PA25517207 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 3 | PA25517208 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 3 | PA25517209 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 3 | PA25517210 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 3 | PA25517211 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 3 | PA25517212 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 3 | PA25517213 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 3 | PA25517214 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 3 | PA25517215 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 3 | PA25517216 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 3 | PA25517217 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 3 | PA25517218 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 3 | PA25517219 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 3 | PA25517220 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 3 | PA25517221 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 3 | PA25517222 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 3 | PA25517223 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 3 | PA25517224 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 3 | PA25517225 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 3 | PA25517226 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 3 | PA25517227 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 3 | PA25517228 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 3 | PA25517229 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 3 | PA25517230 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 3 | PA25517231 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 3 | PA25517232 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 3 | PA25517233 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 3 | PA25517234 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 3 | PA25517235 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 3 | PA25517236 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 3 | PA25517237 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 3 | PA25517238 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 3 | PA25517239 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 3 | PA25517240 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 3 | PA25517241 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 3 | PA25517242 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 3 | PA25517243 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 3 | PA25517244 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 3 | PA25517245 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 3 | PA25517246 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 3 | PA25517247 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 3 | PA25517248 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 3 | PA25517249 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 3 | PA25517250 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 3 | PA25517251 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 3 | PA25517252 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 3 | PA25517253 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 3 | PA25517254 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 3 | PA25517255 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 3 | PA25517256 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 3 | PA25517257 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 3 | PA25517258 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 3 | PA25517259 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 3 | PA25517260 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 3 | PA25517261 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 3 | PA25517262 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 3 | PA25517263 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 3 | PA25517264 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 3 | PA25517265 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 3 | PA25517266 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 3 | PA25517267 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 3 | PA25517268 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 3 | PA25517269 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 3 | PA25517270 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 3 | PA25517271 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 3 | PA25517272 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 3 | PA25517273 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 3 | PA25517274 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 3 | PA25517275 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 3 | PA25517276 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 3 | PA25517277 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 3 | PA25517278 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 3 | PA25517279 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 3 | PA25517280 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 3 | PA25517281 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 3 | PA25517282 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 3 | PA25517283 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 3 | PA25517284 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 3 | PA25517285 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 3 | PA25517286 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 3 | PA25517287 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 3 | PA25517288 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 3 | PA25517289 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 3 | PA25517290 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 3 | PA25517291 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 3 | PA25517292 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 3 | PA25517293 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 3 | PA25517294 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 3 | PA25517295 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 3 | PA25517296 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 3 | PA25517297 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 3 | PA25517298 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 3 | PA25517299 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 3 | PA25517300 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 3 | PA25517301 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 3 | PA25517302 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 3 | PA25517303 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 3 | PA25517304 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 3 | PA25517305 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 3 | PA25517306 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 3 | PA25517307 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 3 | PA25517308 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 3 | PA25517309 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 3 | PA25517310 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 3 | PA25517311 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 3 | PA25517312 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 3 | PA25517313 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 3 | PA25517314 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 3 | PA25517315 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 3 | PA25517316 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 3 | PA25517317 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 3 | PA25517318 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 3 | PA25517319 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 3 | PA25517320 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 3 | PA25517321 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 3 | PA25517322 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 3 | PA25517323 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 3 | PA25517324 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 3 | PA25517325 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 3 | PA25517326 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 3 | PA25517327 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 3 | PA25517328 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 3 | PA25517329 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 3 | PA25517330 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición