## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516865 | 25.6 | 0 | 21 | 8.49 | 0.549 | 244.48 | 344.05 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -1.1123 | 20.5 | 1.102 | 1.094 | 371.08 | 381.03 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.5994 | 21 | 1.098 | 1.088 | 381.03 | 386.02 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516866 | 25.6 | 0 | 21 | 8.49 | 0.549 | 205.75 | 305.34 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -1.0925 | 20.5 | 1.102 | 1.094 | 332.45 | 342.46 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -8.8729e-14 | 21 | 1.098 | 1.088 | 342.46 | 347.51 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.81853 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516867 | 25.6 | 0 | 21 | 8.49 | 0.549 | 206.81 | 307 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.49657 | 20.5 | 1.102 | 1.094 | 334.16 | 344.16 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.0999 | 21 | 1.098 | 1.088 | 344.16 | 349.19 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.41925 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516868 | 25.6 | 0 | 21 | 8.49 | 0.549 | 246.14 | 346.37 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.45685 | 20.5 | 1.102 | 1.094 | 373.53 | 383.5 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.3996 | 21 | 1.098 | 1.088 | 383.5 | 388.51 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.019964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516869 | 25.6 | 0 | 21 | 8.49 | 0.549 | 288.5 | 388.1 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -1.0825 | 20.5 | 1.102 | 1.094 | 415.15 | 425.13 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.2997 | 21 | 1.098 | 1.088 | 425.13 | 430.15 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.2196 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516870 | 25.6 | 0 | 21 | 8.49 | 0.549 | 245.27 | 344.66 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -1.2911 | 20.5 | 1.102 | 1.094 | 371.74 | 381.67 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.7992 | 21 | 1.098 | 1.088 | 381.67 | 386.64 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -0.7786 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516871 | 25.6 | 0 | 21 | 8.49 | 0.549 | 205.17 | 305.02 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.83424 | 20.5 | 1.102 | 1.094 | 332.18 | 342.18 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.0999 | 21 | 1.098 | 1.088 | 342.18 | 347.21 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.41925 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516872 | 25.6 | 0 | 21 | 8.49 | 0.549 | 209.25 | 309.78 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.1589 | 20.5 | 1.102 | 1.094 | 336.96 | 347.03 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | 0.5994 | 21 | 1.098 | 1.088 | 347.03 | 352.03 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -0.17968 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516873 | 25.6 | 0 | 21 | 8.49 | 0.549 | 209.35 | 309.4 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.63561 | 20.5 | 1.102 | 1.094 | 336.54 | 346.42 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -1.2987 | 21 | 1.098 | 1.088 | 346.42 | 351.35 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -1.5772 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516874 | 25.6 | 0 | 21 | 8.49 | 0.549 | 208.55 | 308.29 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.94349 | 20.5 | 1.102 | 1.094 | 335.35 | 345.28 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.7992 | 21 | 1.098 | 1.088 | 345.28 | 350.3 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.2196 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516875 | 25.6 | 0 | 21 | 8.49 | 0.549 | 208.27 | 308.79 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.16884 | 20.5 | 1.102 | 1.094 | 335.94 | 345.89 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.5994 | 21 | 1.098 | 1.088 | 345.89 | 350.88 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516876 | 25.6 | 0 | 21 | 8.49 | 0.549 | 247.8 | 347.95 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.5363 | 20.5 | 1.102 | 1.094 | 375.03 | 384.95 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.8991 | 21 | 1.098 | 1.088 | 384.95 | 389.9 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -1.1779 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516877 | 25.6 | 0 | 21 | 8.49 | 0.549 | 248.29 | 348.55 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.42705 | 20.5 | 1.102 | 1.094 | 375.53 | 385.41 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -1.2987 | 21 | 1.098 | 1.088 | 385.41 | 390.37 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -0.97824 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516878 | 25.6 | 0 | 21 | 8.49 | 0.549 | 207.53 | 307.39 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.82431 | 20.5 | 1.102 | 1.094 | 334.49 | 344.44 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.5994 | 21 | 1.098 | 1.088 | 344.44 | 349.45 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.019964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516879 | 25.6 | 0 | 21 | 8.49 | 0.549 | 246.77 | 346.77 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -0.68527 | 20.5 | 1.102 | 1.094 | 373.78 | 383.67 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -1.1988 | 21 | 1.098 | 1.088 | 383.67 | 388.63 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | -0.97824 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516880 | 25.6 | 0 | 21 | 8.49 | 0.549 | 206.75 | 306.42 | 100.69 | 2,30,160 | 2 | 30,160 | 0.041711 | 2413.9851 | -1.013 | 20.5 | 1.102 | 1.094 | 333.48 | 343.43 | 10.01 | 22,55,230 | 22 | 55,230 | 0.38201 | 26.2036 | -0.5994 | 21 | 1.098 | 1.088 | 343.43 | 348.44 | 5.009 | 18,15,540 | 18 | 15,540 | 0.30432 | 16.4598 | 0.019964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516881 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 100.7 | 199.56 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -1.4455 | 20.7 | 1.095 | 1.086 | 224.84 | 234.76 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.8793 | 21.3 | 1.096 | 1.087 | 234.76 | 239.72 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.8991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516882 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 141.29 | 240.55 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -1.0468 | 20.7 | 1.095 | 1.086 | 265.99 | 275.99 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.079936 | 21.3 | 1.096 | 1.087 | 275.99 | 281.02 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516883 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 103.19 | 202.51 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.98694 | 20.7 | 1.095 | 1.086 | 227.95 | 237.9 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.57954 | 21.3 | 1.096 | 1.087 | 237.9 | 242.94 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516884 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 140.7 | 240.36 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.64799 | 20.7 | 1.095 | 1.086 | 265.75 | 275.69 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.67946 | 21.3 | 1.096 | 1.087 | 275.69 | 280.68 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516885 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 102.85 | 202.66 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.49845 | 20.7 | 1.095 | 1.086 | 228.26 | 238.27 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | 0.019984 | 21.3 | 1.096 | 1.087 | 238.27 | 243.29 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516886 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 103.05 | 203.68 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | 0.31901 | 20.7 | 1.095 | 1.086 | 229.3 | 239.33 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | 0.21982 | 21.3 | 1.096 | 1.087 | 239.33 | 244.39 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 1.0989 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516887 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 101.11 | 200.56 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.85734 | 20.7 | 1.095 | 1.086 | 225.98 | 235.94 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.47962 | 21.3 | 1.096 | 1.087 | 235.94 | 240.97 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516888 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 102.55 | 201.48 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -1.3757 | 20.7 | 1.095 | 1.086 | 226.7 | 236.61 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.97922 | 21.3 | 1.096 | 1.087 | 236.61 | 241.56 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -1.0989 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516889 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 142.43 | 242.08 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.65796 | 20.7 | 1.095 | 1.086 | 267.48 | 277.41 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.77938 | 21.3 | 1.096 | 1.087 | 277.41 | 282.42 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516890 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 101.45 | 200.55 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -1.2063 | 20.7 | 1.095 | 1.086 | 225.88 | 235.79 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.97922 | 21.3 | 1.096 | 1.087 | 235.79 | 240.79 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
27 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516891 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 100.33 | 200.09 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.5483 | 20.7 | 1.095 | 1.086 | 225.52 | 235.45 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.77938 | 21.3 | 1.096 | 1.087 | 235.45 | 240.44 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
28 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516892 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 100.31 | 200.06 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.55827 | 20.7 | 1.095 | 1.086 | 225.44 | 235.37 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.77938 | 21.3 | 1.096 | 1.087 | 235.37 | 240.36 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
29 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516893 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 99.76 | 199.32 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.74768 | 20.7 | 1.095 | 1.086 | 224.65 | 234.58 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.77938 | 21.3 | 1.096 | 1.087 | 234.58 | 239.58 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
30 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516894 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 100.45 | 200.63 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.1296 | 20.7 | 1.095 | 1.086 | 226.17 | 236.14 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.3797 | 21.3 | 1.096 | 1.087 | 236.14 | 241.19 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.8991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516895 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 101.2 | 200.65 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.85734 | 20.7 | 1.095 | 1.086 | 226.02 | 235.99 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.3797 | 21.3 | 1.096 | 1.087 | 235.99 | 241.03 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | 0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516896 | 25.6 | 0 | 20.5 | 8.487 | 0.592 | 100.15 | 199.64 | 100.31 | 2,30,140 | 2 | 30,140 | 0.041706 | 2405.1952 | -0.81747 | 20.7 | 1.095 | 1.086 | 225.02 | 234.99 | 10.008 | 22,32,060 | 22 | 32,060 | 0.37557 | 26.6473 | -0.3797 | 21.3 | 1.096 | 1.087 | 234.99 | 239.96 | 5.005 | 18,09,850 | 18 | 09,850 | 0.30274 | 16.5326 | -0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
33 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516897 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 109.21 | 209.35 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.33838 | 21.3 | 1.096 | 1.088 | 214.75 | 224.69 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.6993 | 21.5 | 1.097 | 1.089 | 224.69 | 229.67 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
34 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516898 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 150.3 | 250 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.77627 | 21.3 | 1.096 | 1.088 | 255.4 | 265.35 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.5994 | 21.5 | 1.097 | 1.089 | 265.35 | 270.37 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
35 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516899 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 108.58 | 208.08 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.97532 | 21.3 | 1.096 | 1.088 | 213.48 | 223.42 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.6993 | 21.5 | 1.097 | 1.089 | 223.42 | 228.44 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
36 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516900 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 109.51 | 209.6 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.38814 | 21.3 | 1.096 | 1.088 | 215 | 225.01 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -8.8729e-14 | 21.5 | 1.097 | 1.089 | 225.01 | 230.09 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516901 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 111.31 | 211.64 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.14928 | 21.3 | 1.096 | 1.088 | 217.05 | 227.08 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | 0.1998 | 21.5 | 1.097 | 1.089 | 227.08 | 232.14 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516902 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 109.2 | 208.8 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.8758 | 21.3 | 1.096 | 1.088 | 214.19 | 224.13 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.6993 | 21.5 | 1.097 | 1.089 | 224.13 | 229.18 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
39 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516903 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 150.79 | 250.93 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.33838 | 21.3 | 1.096 | 1.088 | 256.34 | 266.28 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.6993 | 21.5 | 1.097 | 1.089 | 266.28 | 271.29 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 9.5732e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
40 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516904 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 109.99 | 209.84 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.62699 | 21.3 | 1.096 | 1.088 | 215.24 | 225.22 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.2997 | 21.5 | 1.097 | 1.089 | 225.22 | 230.28 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
41 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516905 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 110.91 | 210.69 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.69666 | 21.3 | 1.096 | 1.088 | 216.05 | 225.95 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -1.0989 | 21.5 | 1.097 | 1.089 | 225.95 | 230.9 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
42 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516906 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 111.19 | 211.29 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.37818 | 21.3 | 1.096 | 1.088 | 216.67 | 226.6 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.7992 | 21.5 | 1.097 | 1.089 | 226.6 | 231.6 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
43 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516907 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 150.38 | 250.2 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.65685 | 21.3 | 1.096 | 1.088 | 255.6 | 265.54 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.6993 | 21.5 | 1.097 | 1.089 | 265.54 | 270.54 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516908 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 151.56 | 251.86 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.17914 | 21.3 | 1.096 | 1.088 | 257.29 | 267.2 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.999 | 21.5 | 1.097 | 1.089 | 267.2 | 272.17 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
45 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516909 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 112.35 | 212.11 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.71656 | 21.3 | 1.096 | 1.088 | 217.53 | 227.54 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -8.8729e-14 | 21.5 | 1.097 | 1.089 | 227.54 | 232.54 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
46 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516910 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 111.75 | 211.69 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.53742 | 21.3 | 1.096 | 1.088 | 217.13 | 227.12 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.1998 | 21.5 | 1.097 | 1.089 | 227.12 | 232.17 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
47 | 2025-09-19 | 2025-09-19 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | [5 empty] | 21.3 | 1.096 | 1.088
48 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516911 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 111.16 | 210.13 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -1.5028 | 21.3 | 1.096 | 1.088 | 215.53 | 225.48 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.5994 | 21.5 | 1.097 | 1.089 | 225.48 | 230.51 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-20 | 2025-09-20 | 3 | 8 |  | PA25516912 | 25.6 | 0 | 20.9 | 8.473 | 0.521 | 110.2 | 210 | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -0.67675 | 21.3 | 1.096 | 1.088 | 215.41 | 225.34 | 10.01 | 22,44,380 | 22 | 44,380 | 0.37899 | 26.412 | -0.7992 | 21.5 | 1.097 | 1.089 | 225.34 | 230.35 | 5.01 | 18,16,100 | 18 | 16,100 | 0.30447 | 16.4547 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
50 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
51 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
52 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
53 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
54 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
55 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
56 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
57 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
58 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
59 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
60 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
61 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.018 | 11,12,670 | 11 | 12,670 | 0.18685 | 26.8554 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
62 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
63 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
64 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
65 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
66 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
67 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
68 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
69 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
70 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
71 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
72 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
73 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
74 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
75 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
76 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
77 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
78 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
79 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
80 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
81 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
82 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
83 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
84 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
85 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
86 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
87 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
88 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
89 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
90 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
91 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
92 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
93 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
94 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
95 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
96 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
97 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
98 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
99 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
100 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
101 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
102 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
103 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
104 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
105 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
106 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
107 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
108 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
109 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
110 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
111 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
112 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
113 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
114 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
115 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
116 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
117 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
118 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
119 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
120 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
121 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
301 | 2025-09-20 | 2025-09-20 | 3 | 8 |  |  | 25.6 | 0 | 20.9 | 8.473 | 0.521 |  |  | 100.48 | 2,29,610 | 2 | 29,610 | 0.041558 | 2417.8063 | -100 | 21.3 | 1.096 | 1.088 |  |  | 5.01 | 11,15,880 | 11 | 15,880 | 0.18774 | 26.6852 | -100 | 18.4 | 1.098 | 1.087 | 0 |  | 2.005 | 7,19,050 | 7, | 19,050 | 0.12196 | 16.44 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 3 | PA25516865 | 0 | -1.1123 | -0.5994 | -0.37932 | CONFORME | Medición
2 | 3 | PA25516866 | 0 | -1.0925 | -8.8729e-14 | 0.81853 | CONFORME | Medición
3 | 3 | PA25516867 | 0 | -0.49657 | -0.0999 | 0.41925 | CONFORME | Medición
4 | 3 | PA25516868 | 0 | -0.45685 | -0.3996 | 0.019964 | CONFORME | Medición
5 | 3 | PA25516869 | 0 | -1.0825 | -0.2997 | 0.2196 | CONFORME | Medición
6 | 3 | PA25516870 | 0 | -1.2911 | -0.7992 | -0.7786 | CONFORME | Medición
7 | 3 | PA25516871 | 0 | -0.83424 | -0.0999 | 0.41925 | CONFORME | Medición
8 | 3 | PA25516872 | 0 | -0.1589 | 0.5994 | -0.17968 | CONFORME | Medición
9 | 3 | PA25516873 | 0 | -0.63561 | -1.2987 | -1.5772 | CONFORME | Medición
10 | 3 | PA25516874 | 0 | -0.94349 | -0.7992 | 0.2196 | CONFORME | Medición
11 | 3 | PA25516875 | 0 | -0.16884 | -0.5994 | -0.37932 | CONFORME | Medición
12 | 3 | PA25516876 | 0 | -0.5363 | -0.8991 | -1.1779 | CONFORME | Medición
13 | 3 | PA25516877 | 0 | -0.42705 | -1.2987 | -0.97824 | CONFORME | Medición
14 | 3 | PA25516878 | 0 | -0.82431 | -0.5994 | 0.019964 | CONFORME | Medición
15 | 3 | PA25516879 | 0 | -0.68527 | -1.1988 | -0.97824 | CONFORME | Medición
16 | 3 | PA25516880 | 0 | -1.013 | -0.5994 | 0.019964 | CONFORME | Medición
17 | 3 | PA25516881 | 0 | -1.4455 | -0.8793 | -0.8991 | CONFORME | Medición
18 | 3 | PA25516882 | 0 | -1.0468 | -0.079936 | 0.4995 | CONFORME | Medición
19 | 3 | PA25516883 | 0 | -0.98694 | -0.57954 | 0.6993 | CONFORME | Medición
20 | 3 | PA25516884 | 0 | -0.64799 | -0.67946 | -0.2997 | CONFORME | Medición
21 | 3 | PA25516885 | 0 | -0.49845 | 0.019984 | 0.2997 | CONFORME | Medición
22 | 3 | PA25516886 | 0 | 0.31901 | 0.21982 | 1.0989 | CONFORME | Medición
23 | 3 | PA25516887 | 0 | -0.85734 | -0.47962 | 0.4995 | CONFORME | Medición
24 | 3 | PA25516888 | 0 | -1.3757 | -0.97922 | -1.0989 | CONFORME | Medición
25 | 3 | PA25516889 | 0 | -0.65796 | -0.77938 | 0.0999 | CONFORME | Medición
26 | 3 | PA25516890 | 0 | -1.2063 | -0.97922 | -0.0999 | CONFORME | Medición
27 | 3 | PA25516891 | 0 | -0.5483 | -0.77938 | -0.2997 | CONFORME | Medición
28 | 3 | PA25516892 | 0 | -0.55827 | -0.77938 | -0.2997 | CONFORME | Medición
29 | 3 | PA25516893 | 0 | -0.74768 | -0.77938 | -0.0999 | CONFORME | Medición
30 | 3 | PA25516894 | 0 | -0.1296 | -0.3797 | 0.8991 | CONFORME | Medición
31 | 3 | PA25516895 | 0 | -0.85734 | -0.3797 | 0.6993 | CONFORME | Medición
32 | 3 | PA25516896 | 0 | -0.81747 | -0.3797 | -0.6993 | CONFORME | Medición
33 | 3 | PA25516897 | 0 | -0.33838 | -0.6993 | -0.5988 | CONFORME | Medición
34 | 3 | PA25516898 | 0 | -0.77627 | -0.5994 | 0.1996 | CONFORME | Medición
35 | 3 | PA25516899 | 0 | -0.97532 | -0.6993 | 0.1996 | CONFORME | Medición
36 | 3 | PA25516900 | 0 | -0.38814 | -8.8729e-14 | 1.3972 | CONFORME | Medición
37 | 3 | PA25516901 | 0 | -0.14928 | 0.1998 | 0.998 | CONFORME | Medición
38 | 3 | PA25516902 | 0 | -0.8758 | -0.6993 | 0.7984 | CONFORME | Medición
39 | 3 | PA25516903 | 0 | -0.33838 | -0.6993 | 9.5732e-13 | CONFORME | Medición
40 | 3 | PA25516904 | 0 | -0.62699 | -0.2997 | 0.998 | CONFORME | Medición
41 | 3 | PA25516905 | 0 | -0.69666 | -1.0989 | -1.1976 | CONFORME | Medición
42 | 3 | PA25516906 | 0 | -0.37818 | -0.7992 | -0.1996 | CONFORME | Medición
43 | 3 | PA25516907 | 0 | -0.65685 | -0.6993 | -0.1996 | CONFORME | Medición
44 | 3 | PA25516908 | 0 | -0.17914 | -0.999 | -0.7984 | CONFORME | Medición
45 | 3 | PA25516909 | 0 | -0.71656 | -8.8729e-14 | -0.1996 | CONFORME | Medición
46 | 3 | PA25516910 | 0 | -0.53742 | -0.1998 | 0.7984 | CONFORME | Medición
48 | 3 | PA25516911 | 0 | -1.5028 | -0.5994 | 0.3992 | CONFORME | Medición
49 | 3 | PA25516912 | 0 | -0.67675 | -0.7992 | -1.7728e-13 | CONFORME | Medición
50 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
51 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
52 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
53 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
54 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
55 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
56 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
57 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
58 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
59 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
60 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
61 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
62 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
63 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
64 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
65 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
66 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
67 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
68 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
69 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
70 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
71 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
72 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
73 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
74 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
75 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
76 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
77 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
78 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
79 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
80 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
81 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
82 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
83 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
84 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
85 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
86 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
87 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
88 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
89 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
90 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
91 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
92 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
93 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
94 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
95 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
96 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
97 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
98 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
99 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
100 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
101 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
102 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
103 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
104 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
105 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
106 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
107 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
108 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
109 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
110 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
111 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
112 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
113 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
114 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
115 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
116 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
117 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
118 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
119 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
120 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
121 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
301 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición