## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517016 | 25.6 | 0 | 20 | 7.898 | 0.557 | 193.7 | 293.49 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.85445 | 20.2 | 1.509 | 1.507 | 299.6 | 309.65 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | 0.31943 | 20.3 | 1.507 | 1.503 | 309.65 | 314.59 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | -1.5936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517017 | 25.6 | 0 | 20 | 7.898 | 0.557 | 156.2 | 256.05 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.79483 | 20.2 | 1.509 | 1.507 | 262.15 | 272.1 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.67878 | 20.3 | 1.507 | 1.503 | 272.1 | 277.15 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517018 | 25.6 | 0 | 20 | 7.898 | 0.557 | 156.7 | 256.6 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.74516 | 20.2 | 1.509 | 1.507 | 262.72 | 272.74 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | 0.019964 | 20.3 | 1.507 | 1.503 | 272.74 | 277.79 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517019 | 25.6 | 0 | 20 | 7.898 | 0.557 | 156.55 | 256.5 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.69548 | 20.2 | 1.509 | 1.507 | 262.72 | 272.7 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.37932 | 20.3 | 1.507 | 1.503 | 272.7 | 277.75 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517020 | 25.6 | 0 | 20 | 7.898 | 0.557 | 153.52 | 253.2 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.96374 | 20.2 | 1.509 | 1.507 | 259.35 | 269.34 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.2795 | 20.3 | 1.507 | 1.503 | 269.34 | 274.35 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517021 | 25.6 | 0 | 20 | 7.898 | 0.557 | 154.92 | 255.12 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.44709 | 20.2 | 1.509 | 1.507 | 261.35 | 271.32 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.47914 | 20.3 | 1.507 | 1.503 | 271.32 | 276.37 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517022 | 25.6 | 0 | 20 | 7.898 | 0.557 | 153.75 | 253.52 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.87432 | 20.2 | 1.509 | 1.507 | 259.6 | 269.54 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.7786 | 20.3 | 1.507 | 1.503 | 269.54 | 274.58 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517023 | 25.6 | 0 | 20 | 7.898 | 0.557 | 154.7 | 254.76 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.58619 | 20.2 | 1.509 | 1.507 | 260.92 | 271.01 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | 0.71871 | 20.3 | 1.507 | 1.503 | 271.01 | 276.12 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 1.7928 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517024 | 25.6 | 0 | 20 | 7.898 | 0.557 | 157.2 | 256.74 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -1.1028 | 20.2 | 1.509 | 1.507 | 262.85 | 272.85 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.17968 | 20.3 | 1.507 | 1.503 | 272.85 | 277.86 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517025 | 25.6 | 0 | 20 | 7.898 | 0.557 | 200.4 | 300.02 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -1.0233 | 20.2 | 1.509 | 1.507 | 306.72 | 316.7 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.37932 | 20.3 | 1.507 | 1.503 | 316.7 | 321.75 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517026 | 25.6 | 0 | 20 | 7.898 | 0.557 | 158.65 | 259.02 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.27819 | 20.2 | 1.509 | 1.507 | 265.12 | 275.14 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | 0.019964 | 20.3 | 1.507 | 1.503 | 275.14 | 280.16 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 7.7848e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517027 | 25.6 | 0 | 20 | 7.898 | 0.557 | 157.85 | 258.35 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.14903 | 20.2 | 1.509 | 1.507 | 264.47 | 274.4 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.87842 | 20.3 | 1.507 | 1.503 | 274.4 | 279.39 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | -0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517028 | 25.6 | 0 | 20 | 7.898 | 0.557 | 239.37 | 339.46 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.55638 | 20.2 | 1.509 | 1.507 | 345.56 | 355.44 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -1.3775 | 20.3 | 1.507 | 1.503 | 355.44 | 360.4 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | -1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517029 | 25.6 | 0 | 20 | 7.898 | 0.557 | 156.85 | 257.3 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.19871 | 20.2 | 1.509 | 1.507 | 263.5 | 273.45 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | -0.67878 | 20.3 | 1.507 | 1.503 | 273.45 | 278.5 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517030 | 25.6 | 0 | 20 | 7.898 | 0.557 | 198.7 | 298.99 | 100.65 | 2,29,920 | 2 | 29,920 | 0.041644 | 2416.889 | -0.35768 | 20.2 | 1.509 | 1.507 | 305.15 | 315.2 | 10.018 | 22,21,430 | 22 | 21,430 | 0.37262 | 26.8853 | 0.31943 | 20.3 | 1.507 | 1.503 | 315.2 | 320.29 | 5.02 | 17,59,000 | 17 | 59,000 | 0.29972 | 16.7488 | 1.3944 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517031 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 159.7 | 259.75 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -0.57637 | 20.4 | 1.507 | 1.504 | 265.52 | 275.47 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -0.68869 | 20.5 | 1.505 | 1.503 | 275.47 | 280.5 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.23914 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517032 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 198.9 | 298.6 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -0.92418 | 20.4 | 1.507 | 1.504 | 304.37 | 314.27 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -1.1877 | 20.5 | 1.505 | 1.503 | 314.27 | 319.37 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 1.6341 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517033 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 118.42 | 217.8 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -1.2422 | 20.4 | 1.507 | 1.504 | 223.6 | 233.5 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -1.1877 | 20.5 | 1.505 | 1.503 | 233.5 | 238.54 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.43842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517034 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 156.82 | 256.36 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -1.0832 | 20.4 | 1.507 | 1.504 | 262.05 | 272.02 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -0.48907 | 20.5 | 1.505 | 1.503 | 272.02 | 277.06 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.43842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517035 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 199.21 | 299.26 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -0.57637 | 20.4 | 1.507 | 1.504 | 305 | 315.01 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -0.089829 | 20.5 | 1.505 | 1.503 | 315.01 | 320.06 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.6377 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517036 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 120.72 | 220.75 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -0.59624 | 20.4 | 1.507 | 1.504 | 226.5 | 236.4 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -1.1877 | 20.5 | 1.505 | 1.503 | 236.4 | 241.45 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.6377 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517037 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 116.2 | 215.75 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -1.0732 | 20.4 | 1.507 | 1.504 | 221.44 | 231.35 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -1.0879 | 20.5 | 1.505 | 1.503 | 231.35 | 236.37 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.039857 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517038 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 117.64 | 217.4 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -0.86455 | 20.4 | 1.507 | 1.504 | 223.06 | 233 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -0.7885 | 20.5 | 1.505 | 1.503 | 233 | 237.95 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -1.3551 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517039 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 114.86 | 213.6 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -1.8782 | 20.4 | 1.507 | 1.504 | 219.35 | 229.27 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -0.98812 | 20.5 | 1.505 | 1.503 | 229.27 | 234.27 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -0.35871 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 6 | 15 |  | PA25517040 | 25.6 | 0 | 20.1 | 7.898 | 0.554 | 156.65 | 256.85 | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -0.42731 | 20.4 | 1.507 | 1.504 | 262.62 | 272.59 | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -0.48907 | 20.5 | 1.505 | 1.503 | 272.59 | 277.63 | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | 0.43842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
27 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
28 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
29 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
30 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
31 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
32 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
33 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
34 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
35 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
36 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
37 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
38 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
39 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
40 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
41 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
42 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
43 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
44 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
45 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
46 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
47 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
48 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
49 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
50 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
51 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
52 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
53 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
54 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
55 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
56 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
57 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
58 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
59 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
60 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
61 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
62 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
63 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
64 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
65 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
66 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
67 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
68 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
69 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
70 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
71 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
72 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
73 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
74 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
75 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
76 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
77 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
78 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
79 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
80 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
81 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
82 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
83 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
84 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
85 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
86 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
87 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
88 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
89 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
90 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
91 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
92 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
93 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
94 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
95 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
96 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
97 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
98 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
99 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
100 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
101 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
102 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
103 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
104 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
105 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
106 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
107 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
108 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
109 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
110 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
111 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
112 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
113 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
114 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
115 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
116 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
117 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
118 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
119 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
120 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
121 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-20 | 2025-09-20 | 6 | 15 |  |  | 25.6 | 0 | 20.1 | 7.898 | 0.554 |  |  | 100.63 | 2,29,980 | 2 | 29,980 | 0.041661 | 2415.4421 | -100 | 20.4 | 1.507 | 1.504 |  |  | 10.019 | 22,13,560 | 22 | 13,560 | 0.37043 | 27.0467 | -100 | 20.5 | 1.505 | 1.503 | 0 |  | 5.018 | 18,00,000 | 18 | 00,000 | 0.3 | 16.7267 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 6 | PA25517016 | 0 | -0.85445 | 0.31943 | -1.5936 | CONFORME | Medición
2 | 6 | PA25517017 | 0 | -0.79483 | -0.67878 | 0.59761 | CONFORME | Medición
3 | 6 | PA25517018 | 0 | -0.74516 | 0.019964 | 0.59761 | CONFORME | Medición
4 | 6 | PA25517019 | 0 | -0.69548 | -0.37932 | 0.59761 | CONFORME | Medición
5 | 6 | PA25517020 | 0 | -0.96374 | -0.2795 | -0.1992 | CONFORME | Medición
6 | 6 | PA25517021 | 0 | -0.44709 | -0.47914 | 0.59761 | CONFORME | Medición
7 | 6 | PA25517022 | 0 | -0.87432 | -0.7786 | 0.39841 | CONFORME | Medición
8 | 6 | PA25517023 | 0 | -0.58619 | 0.71871 | 1.7928 | CONFORME | Medición
9 | 6 | PA25517024 | 0 | -1.1028 | -0.17968 | -0.1992 | CONFORME | Medición
10 | 6 | PA25517025 | 0 | -1.0233 | -0.37932 | 0.59761 | CONFORME | Medición
11 | 6 | PA25517026 | 0 | -0.27819 | 0.019964 | 7.7848e-13 | CONFORME | Medición
12 | 6 | PA25517027 | 0 | -0.14903 | -0.87842 | -0.59761 | CONFORME | Medición
13 | 6 | PA25517028 | 0 | -0.55638 | -1.3775 | -1.1952 | CONFORME | Medición
14 | 6 | PA25517029 | 0 | -0.19871 | -0.67878 | 0.59761 | CONFORME | Medición
15 | 6 | PA25517030 | 0 | -0.35768 | 0.31943 | 1.3944 | CONFORME | Medición
16 | 6 | PA25517031 | 0 | -0.57637 | -0.68869 | 0.23914 | CONFORME | Medición
17 | 6 | PA25517032 | 0 | -0.92418 | -1.1877 | 1.6341 | CONFORME | Medición
18 | 6 | PA25517033 | 0 | -1.2422 | -1.1877 | 0.43842 | CONFORME | Medición
19 | 6 | PA25517034 | 0 | -1.0832 | -0.48907 | 0.43842 | CONFORME | Medición
20 | 6 | PA25517035 | 0 | -0.57637 | -0.089829 | 0.6377 | CONFORME | Medición
21 | 6 | PA25517036 | 0 | -0.59624 | -1.1877 | 0.6377 | CONFORME | Medición
22 | 6 | PA25517037 | 0 | -1.0732 | -1.0879 | 0.039857 | CONFORME | Medición
23 | 6 | PA25517038 | 0 | -0.86455 | -0.7885 | -1.3551 | CONFORME | Medición
24 | 6 | PA25517039 | 0 | -1.8782 | -0.98812 | -0.35871 | CONFORME | Medición
25 | 6 | PA25517040 | 0 | -0.42731 | -0.48907 | 0.43842 | CONFORME | Medición
26 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
27 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
28 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
29 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
30 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
31 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
32 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
33 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
34 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
35 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
36 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
37 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
38 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
39 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
40 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
41 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
42 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
43 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
44 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
45 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
46 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
47 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
48 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
49 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
50 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
51 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
52 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
53 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
54 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
55 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
56 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
57 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
58 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
59 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
60 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
61 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
62 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
63 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
64 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
65 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
66 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
67 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
68 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
69 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
70 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
71 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
72 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
73 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
74 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
75 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
76 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
77 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
78 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
79 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
80 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
81 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
82 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
83 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
84 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
85 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
86 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
87 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
88 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
89 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
90 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
91 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
92 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
93 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
94 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
95 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
96 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
97 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
98 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
99 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
100 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
101 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
102 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
103 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
104 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
105 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
106 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
107 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
108 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
109 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
110 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
111 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
112 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
113 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
114 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
115 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
116 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
117 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
118 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
119 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
120 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
121 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición