## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517413 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 143.41 | 243.74 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.72234 | 19.7 | 1.105 | 1.102 | 253.88 | 263.8 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -1.0178 | 19.9 | 1.107 | 1.1 | 263.8 | 268.81 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
2 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517414 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 143.29 | 243.16 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -1.1775 | 19.7 | 1.105 | 1.102 | 253.25 | 263.16 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -1.1175 | 19.9 | 1.107 | 1.1 | 263.16 | 268.2 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517415 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 143.26 | 243.99 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.32654 | 19.7 | 1.105 | 1.102 | 254.14 | 264.17 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | 0.079824 | 19.9 | 1.107 | 1.1 | 264.17 | 269.27 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517416 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 143.82 | 244.53 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.34633 | 19.7 | 1.105 | 1.102 | 254.66 | 264.7 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | 0.1796 | 19.9 | 1.107 | 1.1 | 264.7 | 269.8 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517417 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 143.34 | 243.79 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.6036 | 19.7 | 1.105 | 1.102 | 253.91 | 263.93 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -0.019956 | 19.9 | 1.107 | 1.1 | 263.93 | 269 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517418 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 143.36 | 243.95 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.46507 | 19.7 | 1.105 | 1.102 | 254.13 | 264.13 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -0.21952 | 19.9 | 1.107 | 1.1 | 264.13 | 269.16 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517419 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 142.64 | 243.36 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.33643 | 19.7 | 1.105 | 1.102 | 253.51 | 263.58 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | 0.47895 | 19.9 | 1.107 | 1.1 | 263.58 | 268.72 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 2.513 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-22 | 2025-09-22 | 4 | 3 |  | PA25517420 | 25.6 | 0 | 19.8 | 8.311 | 0.524 | 142.41 | 242.92 | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -0.54423 | 19.7 | 1.105 | 1.102 | 253.06 | 263.11 | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | 0.27939 | 19.9 | 1.107 | 1.1 | 263.11 | 268.2 | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
10 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
11 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
12 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
13 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
14 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
15 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
16 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
17 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
18 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
19 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
20 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
21 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
22 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
23 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
24 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
25 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
26 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
27 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
28 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
29 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
30 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
31 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
32 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
33 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
34 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
35 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
36 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
37 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
38 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
39 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
40 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
41 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
42 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
43 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
44 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
45 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
46 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
47 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
48 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
49 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
50 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
51 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
52 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
53 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
54 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
55 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
56 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
57 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
58 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
59 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
60 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
61 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
62 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
63 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
64 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
65 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
66 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
67 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
68 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
69 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
70 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
71 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
72 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
73 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
74 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
75 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
76 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
77 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
78 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
79 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
80 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
81 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
82 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
83 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
84 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
85 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
86 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
87 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
88 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
89 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
90 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
91 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
92 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
93 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
94 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
95 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
96 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
97 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
98 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
99 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
100 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
101 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
102 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
103 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
104 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
105 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
106 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
107 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
108 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
109 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
110 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
111 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
112 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
113 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
114 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
115 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
116 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
117 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
118 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
119 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
120 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
121 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-22 | 2025-09-22 | 4 | 3 |  |  | 25.6 | 0 | 19.8 | 8.311 | 0.524 |  |  | 101.06 | 2,30,960 | 2 | 30,960 | 0.041933 | 2410.0159 | -100 | 19.7 | 1.105 | 1.102 |  |  | 10.022 | 22,56,360 | 22 | 56,360 | 0.38232 | 26.2135 | -100 | 19.9 | 1.107 | 1.1 | 0 |  | 5.014 | 18,31,150 | 18 | 31,150 | 0.30865 | 16.2448 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 4 | PA25517413 | 0 | -0.72234 | -1.0178 | -0.079777 | CONFORME | Medición
2 | 4 | PA25517414 | 0 | -1.1775 | -1.1175 | 0.51855 | CONFORME | Medición
3 | 4 | PA25517415 | 0 | -0.32654 | 0.079824 | 1.7152 | CONFORME | Medición
4 | 4 | PA25517416 | 0 | -0.34633 | 0.1796 | 1.7152 | CONFORME | Medición
5 | 4 | PA25517417 | 0 | -0.6036 | -0.019956 | 1.1169 | CONFORME | Medición
6 | 4 | PA25517418 | 0 | -0.46507 | -0.21952 | 0.31911 | CONFORME | Medición
7 | 4 | PA25517419 | 0 | -0.33643 | 0.47895 | 2.513 | CONFORME | Medición
8 | 4 | PA25517420 | 0 | -0.54423 | 0.27939 | 1.5158 | CONFORME | Medición
9 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
10 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
11 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
12 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
13 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
14 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
15 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
16 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
17 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
18 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
19 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
20 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
21 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
22 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
23 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
24 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
25 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
26 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
27 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
28 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
29 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
30 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
31 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
32 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
33 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
34 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
35 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
36 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
37 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
38 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
39 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
40 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
41 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
42 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
43 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
44 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
45 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
46 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
47 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
48 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
49 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
50 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
51 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
52 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
53 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
54 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
55 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
56 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
57 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
58 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
59 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
60 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
61 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
62 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
63 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
64 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
65 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
66 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
67 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
68 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
69 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
70 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
71 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
72 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
73 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
74 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
75 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
76 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
77 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
78 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
79 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
80 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
81 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
82 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
83 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
84 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
85 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
86 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
87 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
88 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
89 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
90 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
91 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
92 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
93 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
94 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
95 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
96 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
97 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
98 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
99 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
100 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
101 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
102 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
103 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
104 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
105 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
106 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
107 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
108 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
109 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
110 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
111 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
112 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
113 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
114 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
115 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
116 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
117 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
118 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
119 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
120 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
121 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición