## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517766 | 25.6 | 0 | 20 | 7.67 | 0.525 | 300.61 | 401.49 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.1688 | 19.2 | 1.51 | 1.507 | 404.12 | 414.14 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | 0.029949 | 19.2 | 1.515 | 1.514 | 414.14 | 419.19 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
2 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517767 | 25.6 | 0 | 20 | 7.67 | 0.525 | 148.76 | 249.63 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.15887 | 19.2 | 1.51 | 1.507 | 252.3 | 262.3 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.16971 | 19.2 | 1.515 | 1.514 | 262.3 | 267.36 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517768 | 25.6 | 0 | 20 | 7.67 | 0.525 | 147.36 | 247.59 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | -0.47662 | 19.2 | 1.51 | 1.507 | 250.23 | 260.24 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.069881 | 19.2 | 1.515 | 1.514 | 260.24 | 265.27 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517769 | 25.6 | 0 | 20 | 7.67 | 0.525 | 188.25 | 290.06 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 1.0922 | 19.2 | 1.51 | 1.507 | 292.71 | 302.84 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | 1.1281 | 19.2 | 1.515 | 1.514 | 302.84 | 307.93 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
5 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517770 | 25.6 | 0 | 20 | 7.67 | 0.525 | 187.02 | 287.57 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | -0.15887 | 19.2 | 1.51 | 1.507 | 290.24 | 300.22 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.36937 | 19.2 | 1.515 | 1.514 | 300.22 | 305.29 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517771 | 25.6 | 0 | 20 | 7.67 | 0.525 | 184.86 | 285.75 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.17873 | 19.2 | 1.51 | 1.507 | 288.44 | 298.44 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.16971 | 19.2 | 1.515 | 1.514 | 298.44 | 303.46 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517772 | 25.6 | 0 | 20 | 7.67 | 0.525 | 144.44 | 244.66 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | -0.48655 | 19.2 | 1.51 | 1.507 | 247.31 | 257.3 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.26954 | 19.2 | 1.515 | 1.514 | 257.3 | 262.35 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517773 | 25.6 | 0 | 20 | 7.67 | 0.525 | 144.96 | 245.64 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | -0.029789 | 19.2 | 1.51 | 1.507 | 248.29 | 258.29 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.16971 | 19.2 | 1.515 | 1.514 | 258.29 | 263.35 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517774 | 25.6 | 0 | 20 | 7.67 | 0.525 | 148.74 | 249.56 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.10922 | 19.2 | 1.51 | 1.507 | 252.22 | 262.2 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.36937 | 19.2 | 1.515 | 1.514 | 262.2 | 267.26 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
10 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517775 | 25.6 | 0 | 20 | 7.67 | 0.525 | 188.02 | 288.55 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | -0.17873 | 19.2 | 1.51 | 1.507 | 291.2 | 301.16 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.56903 | 19.2 | 1.515 | 1.514 | 301.16 | 306.18 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517776 | 25.6 | 0 | 20 | 7.67 | 0.525 | 185.51 | 286.45 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.22838 | 19.2 | 1.51 | 1.507 | 289.1 | 299.07 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | -0.4692 | 19.2 | 1.515 | 1.514 | 299.07 | 304.07 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | -0.27922 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517777 | 25.6 | 0 | 20 | 7.67 | 0.525 | 188.54 | 289.59 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.3376 | 19.2 | 1.51 | 1.507 | 292.28 | 302.34 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | 0.42927 | 19.2 | 1.515 | 1.514 | 302.34 | 307.4 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
13 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517778 | 25.6 | 0 | 20 | 7.67 | 0.525 | 186.85 | 287.44 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | -0.11915 | 19.2 | 1.51 | 1.507 | 290.06 | 300.09 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | 0.12978 | 19.2 | 1.515 | 1.514 | 300.09 | 305.15 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517779 | 25.6 | 0 | 20 | 7.67 | 0.525 | 148.44 | 249.15 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 1.4111e-14 | 19.2 | 1.51 | 1.507 | 251.78 | 261.84 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | 0.42927 | 19.2 | 1.515 | 1.514 | 261.84 | 266.95 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 1.9146 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
15 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517780 | 25.6 | 0 | 20 | 7.67 | 0.525 | 147.66 | 248.54 | 100.71 | 2,31,670 | 2 | 31,670 | 0.042131 | 2390.4266 | 0.1688 | 19.2 | 1.51 | 1.507 | 251.21 | 261.23 | 10.017 | 22,39,740 | 22 | 39,740 | 0.37771 | 26.5207 | 0.029949 | 19.2 | 1.515 | 1.514 | 261.23 | 266.29 | 5.014 | 18,18,010 | 18 | 18,010 | 0.305 | 16.4392 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
16 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517781 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 101.48 | 201.7 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.52605 | 19.3 | 1.514 | 1.51 | 204.45 | 214.45 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.15974 | 20.1 | 1.517 | 1.51 | 214.45 | 219.49 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517782 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 102.19 | 202.67 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.26799 | 19.3 | 1.514 | 1.51 | 205.44 | 215.39 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.65895 | 20.1 | 1.517 | 1.51 | 215.39 | 220.4 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
18 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517783 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 103.97 | 205.09 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.36725 | 19.3 | 1.514 | 1.51 | 207.84 | 217.85 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.059904 | 20.1 | 1.517 | 1.51 | 217.85 | 222.93 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517784 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 100.76 | 201.91 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.39702 | 19.3 | 1.514 | 1.51 | 204.69 | 214.69 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.15974 | 20.1 | 1.517 | 1.51 | 214.69 | 219.74 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517785 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 100.16 | 200.46 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.44665 | 19.3 | 1.514 | 1.51 | 203.21 | 213.21 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.15974 | 20.1 | 1.517 | 1.51 | 213.21 | 218.28 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
21 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517786 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 103.11 | 203.82 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.039702 | 19.3 | 1.514 | 1.51 | 206.59 | 216.6 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.059904 | 20.1 | 1.517 | 1.51 | 216.6 | 221.66 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517787 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 103.01 | 203.6 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.15881 | 19.3 | 1.514 | 1.51 | 206.37 | 216.35 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.35942 | 20.1 | 1.517 | 1.51 | 216.35 | 221.36 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
23 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517788 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 103.8 | 204.84 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.28784 | 19.3 | 1.514 | 1.51 | 207.59 | 217.56 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.45927 | 20.1 | 1.517 | 1.51 | 217.56 | 222.6 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517789 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 104.96 | 205.39 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.31762 | 19.3 | 1.514 | 1.51 | 208.11 | 218.11 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.15974 | 20.1 | 1.517 | 1.51 | 218.11 | 223.18 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517790 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 102.06 | 203.29 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.47643 | 19.3 | 1.514 | 1.51 | 206.01 | 216.06 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | 0.33946 | 20.1 | 1.517 | 1.51 | 216.06 | 221.17 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 1.7928 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
26 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517791 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 138.62 | 239.45 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.079404 | 19.3 | 1.514 | 1.51 | 242.21 | 252.26 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | 0.33946 | 20.1 | 1.517 | 1.51 | 252.26 | 257.34 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
27 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517792 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 138.7 | 239.01 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.43672 | 19.3 | 1.514 | 1.51 | 241.74 | 251.71 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.45927 | 20.1 | 1.517 | 1.51 | 251.71 | 256.76 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517793 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 140.06 | 240.52 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | -0.28784 | 19.3 | 1.514 | 1.51 | 243.26 | 253.25 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.25958 | 20.1 | 1.517 | 1.51 | 253.25 | 258.31 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517794 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 145.05 | 246.15 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.34739 | 19.3 | 1.514 | 1.51 | 248.89 | 258.89 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | -0.15974 | 20.1 | 1.517 | 1.51 | 258.89 | 263.95 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517795 | 25.6 | 0 | 19.3 | 7.666 | 0.54 | 144.45 | 245.31 | 100.75 | 2,32,340 | 2 | 32,340 | 0.042317 | 2380.8586 | 0.10918 | 19.3 | 1.514 | 1.51 | 248.05 | 258.11 | 10.016 | 22,39,220 | 22 | 39,220 | 0.37756 | 26.5282 | 0.4393 | 20.1 | 1.517 | 1.51 | 258.11 | 263.18 | 5.02 | 18,18,260 | 18 | 18,260 | 0.30507 | 16.4551 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
31 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517796 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 144.71 | 245.09 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.29797 | 19.6 | 1.514 | 1.507 | 247.6 | 257.56 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.54918 | 20 | 1.511 | 1.509 | 257.56 | 262.59 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517797 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 144.17 | 244.77 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.07946 | 19.6 | 1.514 | 1.507 | 247.32 | 257.34 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.049925 | 20 | 1.511 | 1.509 | 257.34 | 262.4 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
33 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517798 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 144.01 | 245.05 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.35757 | 19.6 | 1.514 | 1.507 | 247.59 | 257.63 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.24963 | 20 | 1.511 | 1.509 | 257.63 | 262.65 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
34 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517799 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 144.44 | 244.19 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.92372 | 19.6 | 1.514 | 1.507 | 246.69 | 256.65 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.54918 | 20 | 1.511 | 1.509 | 256.65 | 261.69 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
35 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517800 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 145.6 | 246.05 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.22845 | 19.6 | 1.514 | 1.507 | 248.55 | 258.6 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.34948 | 20 | 1.511 | 1.509 | 258.6 | 263.68 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
36 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517801 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 144.51 | 245.2 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.0099325 | 19.6 | 1.514 | 1.507 | 247.7 | 257.66 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.54918 | 20 | 1.511 | 1.509 | 257.66 | 262.67 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517802 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 145.65 | 246.42 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.089392 | 19.6 | 1.514 | 1.507 | 248.94 | 258.94 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.14978 | 20 | 1.511 | 1.509 | 258.94 | 264.01 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517803 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 146.44 | 247.09 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.029797 | 19.6 | 1.514 | 1.507 | 249.6 | 259.61 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.049925 | 20 | 1.511 | 1.509 | 259.61 | 264.68 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
39 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517804 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 145.46 | 246.78 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.63568 | 19.6 | 1.514 | 1.507 | 249.35 | 259.43 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.64903 | 20 | 1.511 | 1.509 | 259.43 | 264.51 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
40 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517805 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 145.55 | 246.4 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.16885 | 19.6 | 1.514 | 1.507 | 248.91 | 258.9 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.24963 | 20 | 1.511 | 1.509 | 258.9 | 263.96 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
41 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517806 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 146.34 | 246.56 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.45689 | 19.6 | 1.514 | 1.507 | 249.11 | 259.12 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.049925 | 20 | 1.511 | 1.509 | 259.12 | 264.21 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
42 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517807 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 144.72 | 245.69 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.28804 | 19.6 | 1.514 | 1.507 | 248.25 | 258.3 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.34948 | 20 | 1.511 | 1.509 | 258.3 | 263.4 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
43 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517808 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 147.78 | 248.55 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | 0.089392 | 19.6 | 1.514 | 1.507 | 251.1 | 261.16 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.44933 | 20 | 1.511 | 1.509 | 261.16 | 266.25 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517809 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 142.73 | 243.18 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.22845 | 19.6 | 1.514 | 1.507 | 245.69 | 255.66 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | -0.44933 | 20 | 1.511 | 1.509 | 255.66 | 260.7 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
45 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517810 | 25.6 | 0 | 19.6 | 7.651 | 0.523 | 145.34 | 245.34 | 100.68 | 2,32,120 | 2 | 32,120 | 0.042256 | 2382.6453 | -0.67541 | 19.6 | 1.514 | 1.507 | 247.84 | 257.86 | 10.015 | 22,40,950 | 22 | 40,950 | 0.37804 | 26.4918 | 0.049925 | 20 | 1.511 | 1.509 | 257.86 | 262.91 | 5.014 | 18,16,500 | 18 | 16,500 | 0.30458 | 16.4618 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
46 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517811 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 98.64 | 199.2 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.14894 | 19.7 | 1.511 | 1.509 | 201.62 | 211.65 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | 0.17978 | 20.2 | 1.516 | 1.511 | 211.65 | 216.69 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
47 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517812 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 96.09 | 195.97 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.82415 | 19.7 | 1.511 | 1.509 | 198.4 | 208.31 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -1.0188 | 20.2 | 1.516 | 1.511 | 208.31 | 213.31 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
48 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517813 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 99.1 | 199.86 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | 0.049648 | 19.7 | 1.511 | 1.509 | 202.35 | 212.34 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.21974 | 20.2 | 1.516 | 1.511 | 212.34 | 217.34 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517814 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 96.91 | 197.36 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.25817 | 19.7 | 1.511 | 1.509 | 199.78 | 209.79 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.019976 | 20.2 | 1.516 | 1.511 | 209.79 | 214.82 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
50 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517815 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 98.54 | 199.54 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | 0.28796 | 19.7 | 1.511 | 1.509 | 201.99 | 211.96 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.4195 | 20.2 | 1.516 | 1.511 | 211.96 | 217 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
51 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517816 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 96.35 | 196.51 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.54612 | 19.7 | 1.511 | 1.509 | 198.94 | 208.87 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.81902 | 20.2 | 1.516 | 1.511 | 208.87 | 213.85 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
52 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517817 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 97.61 | 197.55 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.76457 | 19.7 | 1.511 | 1.509 | 199.99 | 209.94 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.61926 | 20.2 | 1.516 | 1.511 | 209.94 | 214.96 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
53 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517818 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 98.8 | 199.46 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.049648 | 19.7 | 1.511 | 1.509 | 201.86 | 211.79 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.81902 | 20.2 | 1.516 | 1.511 | 211.79 | 216.79 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
54 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517819 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 99.71 | 200.7 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | 0.27803 | 19.7 | 1.511 | 1.509 | 203.15 | 213.12 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.4195 | 20.2 | 1.516 | 1.511 | 213.12 | 218.14 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517820 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 97.7 | 198.25 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.15887 | 19.7 | 1.511 | 1.509 | 200.65 | 210.69 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | 0.27966 | 20.2 | 1.516 | 1.511 | 210.69 | 215.75 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
56 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517821 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 98.95 | 198.93 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.72485 | 19.7 | 1.511 | 1.509 | 201.39 | 211.39 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.11986 | 20.2 | 1.516 | 1.511 | 211.39 | 216.46 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
57 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517822 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 99.44 | 200.6 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | 0.44683 | 19.7 | 1.511 | 1.509 | 203.05 | 213.05 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.11986 | 20.2 | 1.516 | 1.511 | 213.05 | 218.15 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
58 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517823 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 96.5 | 197.35 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | 0.13901 | 19.7 | 1.511 | 1.509 | 199.76 | 209.76 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -0.11986 | 20.2 | 1.516 | 1.511 | 209.76 | 214.81 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
59 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517824 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 98.35 | 198.49 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.56598 | 19.7 | 1.511 | 1.509 | 200.86 | 210.71 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | -1.6181 | 20.2 | 1.516 | 1.511 | 210.71 | 215.61 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | -2.2931 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
60 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517825 | 25.6 | 0 | 19.7 | 7.671 | 0.52 | 98.9 | 199.45 | 100.71 | 2,32,050 | 2 | 32,050 | 0.042236 | 2384.4525 | -0.15887 | 19.7 | 1.511 | 1.509 | 201.87 | 211.93 | 10.012 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.5146 | 0.47942 | 20.2 | 1.516 | 1.511 | 211.93 | 217 | 5.015 | 18,12,460 | 18 | 12,460 | 0.30346 | 16.526 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
61 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517826 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 104.19 | 204.62 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.26812 | 19.8 | 1.516 | 1.51 | 207.14 | 217.1 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.52931 | 20.4 | 1.517 | 1.512 | 217.1 | 222.17 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
62 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517827 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 102.95 | 203.16 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.48659 | 19.8 | 1.516 | 1.51 | 205.64 | 215.6 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.52931 | 20.4 | 1.517 | 1.512 | 215.6 | 220.64 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
63 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517828 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 138.87 | 239.55 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.019861 | 19.8 | 1.516 | 1.51 | 242.05 | 252.09 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | 0.26965 | 20.4 | 1.517 | 1.512 | 252.09 | 257.14 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
64 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517829 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 140.55 | 241.14 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.10924 | 19.8 | 1.516 | 1.51 | 243.61 | 253.59 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.32957 | 20.4 | 1.517 | 1.512 | 253.59 | 258.63 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
65 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517830 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 143.25 | 243.95 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -1.4112e-14 | 19.8 | 1.516 | 1.51 | 246.48 | 256.51 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | 0.16978 | 20.4 | 1.517 | 1.512 | 256.51 | 261.56 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
66 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517831 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 101.61 | 202.4 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | 0.089374 | 19.8 | 1.516 | 1.51 | 204.88 | 214.91 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | 0.16978 | 20.4 | 1.517 | 1.512 | 214.91 | 220 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 1.4753 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
67 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517832 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 139.98 | 240.49 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.18868 | 19.8 | 1.516 | 1.51 | 242.98 | 253 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | 0.069909 | 20.4 | 1.517 | 1.512 | 253 | 258.11 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 1.874 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
68 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517833 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 141.25 | 241.92 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.029791 | 19.8 | 1.516 | 1.51 | 244.45 | 254.41 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.52931 | 20.4 | 1.517 | 1.512 | 254.41 | 259.45 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
69 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517834 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 102.7 | 203.85 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | 0.44687 | 19.8 | 1.516 | 1.51 | 206.38 | 216.4 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | 0.069909 | 20.4 | 1.517 | 1.512 | 216.4 | 221.45 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
70 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517835 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 105.54 | 206.31 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | 0.069513 | 19.8 | 1.516 | 1.51 | 208.79 | 218.76 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.42944 | 20.4 | 1.517 | 1.512 | 218.76 | 223.79 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.27911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
71 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517836 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 103.7 | 204.35 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.049652 | 19.8 | 1.516 | 1.51 | 206.84 | 216.86 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | 0.069909 | 20.4 | 1.517 | 1.512 | 216.86 | 221.94 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 1.2759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
72 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517837 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 105.81 | 206.95 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | 0.43694 | 19.8 | 1.516 | 1.51 | 209.49 | 219.46 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.42944 | 20.4 | 1.517 | 1.512 | 219.46 | 224.51 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
73 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517838 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 105.65 | 206.55 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | 0.19861 | 19.8 | 1.516 | 1.51 | 209.05 | 219 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.62918 | 20.4 | 1.517 | 1.512 | 219 | 224 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
74 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517839 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 141.92 | 242.69 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | 0.069513 | 19.8 | 1.516 | 1.51 | 245.22 | 255.22 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.12983 | 20.4 | 1.517 | 1.512 | 255.22 | 260.28 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | 0.87719 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
75 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517840 | 25.6 | 0 | 19.8 | 7.674 | 0.529 | 136.78 | 237.18 | 100.7 | 2,32,090 | 2 | 32,090 | 0.042247 | 2383.5887 | -0.29791 | 19.8 | 1.516 | 1.51 | 239.65 | 249.6 | 10.013 | 22,41,420 | 22 | 41,420 | 0.37817 | 26.4774 | -0.62918 | 20.4 | 1.517 | 1.512 | 249.6 | 254.6 | 5.016 | 18,18,140 | 18 | 18,140 | 0.30504 | 16.4438 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
76 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517841 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 141.61 | 241.94 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.41687 | 19.9 | 1.519 | 1.508 | 244.72 | 254.74 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | 0.059916 | 20.5 | 1.513 | 1.512 | 254.74 | 259.79 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
77 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517842 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 140.78 | 241.65 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | 0.11911 | 19.9 | 1.519 | 1.508 | 244.45 | 254.49 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | 0.25964 | 20.5 | 1.513 | 1.512 | 254.49 | 259.55 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
78 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517843 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 144.49 | 244.99 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.24814 | 19.9 | 1.519 | 1.508 | 247.76 | 257.79 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | 0.15978 | 20.5 | 1.513 | 1.512 | 257.79 | 262.81 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
79 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517844 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 142.91 | 243.9 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | 0.23821 | 19.9 | 1.519 | 1.508 | 246.71 | 256.7 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.23966 | 20.5 | 1.513 | 1.512 | 256.7 | 261.69 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
80 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517845 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 142 | 242.3 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.44665 | 19.9 | 1.519 | 1.508 | 245.06 | 255.06 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.1398 | 20.5 | 1.513 | 1.512 | 255.06 | 260.14 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
81 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517846 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 179.21 | 280.11 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | 0.14888 | 19.9 | 1.519 | 1.508 | 282.9 | 292.98 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | 0.65908 | 20.5 | 1.513 | 1.512 | 292.98 | 298.06 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
82 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517847 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 100.99 | 201.48 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.25806 | 19.9 | 1.519 | 1.508 | 204.27 | 214.22 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.63911 | 20.5 | 1.513 | 1.512 | 214.22 | 219.24 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
83 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517848 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 100.7 | 201.15 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.29777 | 19.9 | 1.519 | 1.508 | 203.94 | 213.87 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.83883 | 20.5 | 1.513 | 1.512 | 213.87 | 218.87 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
84 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517849 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 100.15 | 200.66 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.23821 | 19.9 | 1.519 | 1.508 | 203.49 | 213.41 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.93869 | 20.5 | 1.513 | 1.512 | 213.41 | 218.44 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
85 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517850 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 139.55 | 239.75 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.54591 | 19.9 | 1.519 | 1.508 | 242.54 | 252.54 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.1398 | 20.5 | 1.513 | 1.512 | 252.54 | 257.57 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
86 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517851 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 103.94 | 204.41 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.27792 | 19.9 | 1.519 | 1.508 | 207.21 | 217.14 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.83883 | 20.5 | 1.513 | 1.512 | 217.14 | 222.14 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
87 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517852 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 101.65 | 202.58 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | 0.17866 | 19.9 | 1.519 | 1.508 | 205.4 | 215.37 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.43938 | 20.5 | 1.513 | 1.512 | 215.37 | 220.39 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
88 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517853 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 101.89 | 202.59 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | -0.049628 | 19.9 | 1.519 | 1.508 | 205.4 | 215.32 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.93869 | 20.5 | 1.513 | 1.512 | 215.32 | 220.35 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
89 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517854 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 101.41 | 202.4 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | 0.23821 | 19.9 | 1.519 | 1.508 | 205.19 | 215.19 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.1398 | 20.5 | 1.513 | 1.512 | 215.19 | 220.21 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
90 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517855 | 25.6 | 0 | 19.9 | 7.665 | 0.543 | 103.54 | 204.49 | 100.75 | 2,32,060 | 2 | 32,060 | 0.042239 | 2385.2427 | 0.19851 | 19.9 | 1.519 | 1.508 | 207.29 | 217.29 | 10.014 | 22,39,560 | 22 | 39,560 | 0.37766 | 26.5162 | -0.1398 | 20.5 | 1.513 | 1.512 | 217.29 | 222.3 | 5.015 | 18,18,050 | 18 | 18,050 | 0.30501 | 16.4419 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
91 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517856 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 100.5 | 201.04 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.19853 | 20 | 1.512 | 1.51 | 203.61 | 213.65 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | 0.26965 | 20.6 | 1.51 | 1.512 | 213.65 | 218.71 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
92 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517857 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 99.75 | 200.6 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.10919 | 20 | 1.512 | 1.51 | 203.22 | 213.18 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.52931 | 20.6 | 1.51 | 1.512 | 213.18 | 218.24 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
93 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517858 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 99.93 | 200.64 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.02978 | 20 | 1.512 | 1.51 | 203.24 | 213.19 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.62918 | 20.6 | 1.51 | 1.512 | 213.19 | 218.21 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
94 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517859 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 100.25 | 200.51 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.47647 | 20 | 1.512 | 1.51 | 203.09 | 213.05 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.52931 | 20.6 | 1.51 | 1.512 | 213.05 | 218.1 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
95 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517860 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 101 | 201.38 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.35736 | 20 | 1.512 | 1.51 | 203.95 | 213.93 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.32957 | 20.6 | 1.51 | 1.512 | 213.93 | 219 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
96 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517861 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 99.64 | 200.35 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.02978 | 20 | 1.512 | 1.51 | 202.91 | 212.91 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.12983 | 20.6 | 1.51 | 1.512 | 212.91 | 217.94 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
97 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517862 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 101.16 | 202.06 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.15882 | 20 | 1.512 | 1.51 | 204.65 | 214.68 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | 0.16978 | 20.6 | 1.51 | 1.512 | 214.68 | 219.7 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
98 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517863 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 103.5 | 204.31 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.069486 | 20 | 1.512 | 1.51 | 206.89 | 216.91 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | 0.069909 | 20.6 | 1.51 | 1.512 | 216.91 | 221.99 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
99 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517864 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 104.5 | 205.65 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.40699 | 20 | 1.512 | 1.51 | 208.29 | 218.29 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.12983 | 20.6 | 1.51 | 1.512 | 218.29 | 223.31 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
100 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517865 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 99.45 | 200.22 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.02978 | 20 | 1.512 | 1.51 | 202.8 | 212.81 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.029961 | 20.6 | 1.51 | 1.512 | 212.81 | 217.86 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
101 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517866 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 101.55 | 203.36 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 1.0621 | 20 | 1.512 | 1.51 | 205.95 | 215.99 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | 0.26965 | 20.6 | 1.51 | 1.512 | 215.99 | 221.05 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
102 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517867 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 101.7 | 202.16 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.27794 | 20 | 1.512 | 1.51 | 204.71 | 214.7 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.2297 | 20.6 | 1.51 | 1.512 | 214.7 | 219.71 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
103 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517868 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 100.85 | 201.91 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.31765 | 20 | 1.512 | 1.51 | 204.53 | 214.55 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | 0.069909 | 20.6 | 1.51 | 1.512 | 214.55 | 219.61 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
104 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517869 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 215.86 | 316.26 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | -0.3375 | 20 | 1.512 | 1.51 | 318.82 | 328.75 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | -0.82892 | 20.6 | 1.51 | 1.512 | 328.75 | 333.76 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
105 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517870 | 25.6 | 0 | 19.9 | 7.652 | 0.516 | 102.99 | 203.77 | 100.74 | 2,32,110 | 2 | 32,110 | 0.042253 | 2384.2219 | 0.039706 | 20 | 1.512 | 1.51 | 206.39 | 216.45 | 10.013 | 22,40,700 | 22 | 40,700 | 0.37797 | 26.4914 | 0.46939 | 20.6 | 1.51 | 1.512 | 216.45 | 221.51 | 5.014 | 18,14,170 | 18 | 14,170 | 0.30394 | 16.4969 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
106 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517871 | 25.6 | 0 | 20 | 7.66 | 0.541 | 100.09 | 200.92 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | 0.14899 | 20 | 1.516 | 1.509 | 204 | 214.05 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | 0.36952 | 20.5 | 1.514 | 1.507 | 214.05 | 219.11 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.87719 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
107 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517872 | 25.6 | 0 | 20 | 7.66 | 0.541 | 98.58 | 198.77 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.48669 | 20 | 1.516 | 1.509 | 201.84 | 211.74 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -1.1285 | 20.5 | 1.514 | 1.507 | 211.74 | 216.75 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -0.11962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
108 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517873 | 25.6 | 0 | 20 | 7.66 | 0.541 | 101.05 | 201.75 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | 0.019865 | 20 | 1.516 | 1.509 | 204.82 | 214.78 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.52931 | 20.5 | 1.514 | 1.507 | 214.78 | 219.83 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
109 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517874 | 25.6 | 0 | 20 | 7.66 | 0.541 | 143.64 | 243.62 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.69527 | 20 | 1.516 | 1.509 | 246.66 | 256.68 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | 0.069909 | 20.5 | 1.514 | 1.507 | 256.68 | 261.75 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
110 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517875 | 25.6 | 0 | 20 | 7.66 | 0.541 | 142.84 | 242.88 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.63568 | 20 | 1.516 | 1.509 | 245.94 | 255.89 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.62918 | 20.5 | 1.514 | 1.507 | 255.89 | 260.92 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.27911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
111 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517876 | 25.6 | 0 | 20 | 7.66 | 0.541 | 102.64 | 202.84 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.47676 | 20 | 1.516 | 1.509 | 205.89 | 215.85 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.52931 | 20.5 | 1.514 | 1.507 | 215.85 | 220.89 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
112 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517877 | 25.6 | 0 | 20 | 7.66 | 0.541 | 101.62 | 200.91 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -1.3806 | 20 | 1.516 | 1.509 | 203.92 | 213.84 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.92879 | 20.5 | 1.514 | 1.507 | 213.84 | 218.84 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
113 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517878 | 25.6 | 0 | 20 | 7.66 | 0.541 | 101.51 | 200.75 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -1.4303 | 20 | 1.516 | 1.509 | 203.78 | 213.7 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.92879 | 20.5 | 1.514 | 1.507 | 213.7 | 218.7 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
114 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517879 | 25.6 | 0 | 20 | 7.66 | 0.541 | 102.92 | 203.19 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.40723 | 20 | 1.516 | 1.509 | 206.25 | 216.19 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.72905 | 20.5 | 1.514 | 1.507 | 216.19 | 221.19 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
115 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517880 | 25.6 | 0 | 20 | 7.66 | 0.541 | 102.55 | 202.6 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.62574 | 20 | 1.516 | 1.509 | 205.65 | 215.6 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.62918 | 20.5 | 1.514 | 1.507 | 215.6 | 220.65 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
116 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517881 | 25.6 | 0 | 20 | 7.66 | 0.541 | 100.85 | 201.39 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.13905 | 20 | 1.516 | 1.509 | 204.44 | 214.39 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.62918 | 20.5 | 1.514 | 1.507 | 214.39 | 219.4 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -0.11962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
117 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517882 | 25.6 | 0 | 20 | 7.66 | 0.541 | 101.45 | 201.87 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.25824 | 20 | 1.516 | 1.509 | 204.95 | 214.91 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.52931 | 20.5 | 1.514 | 1.507 | 214.91 | 219.94 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.27911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
118 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517883 | 25.6 | 0 | 20 | 7.66 | 0.541 | 101.83 | 201.79 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.71514 | 20 | 1.516 | 1.509 | 204.84 | 214.75 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -1.0287 | 20.5 | 1.514 | 1.507 | 214.75 | 219.75 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
119 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517884 | 25.6 | 0 | 20 | 7.66 | 0.541 | 99.71 | 199.74 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.64561 | 20 | 1.516 | 1.509 | 202.46 | 212.39 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.82892 | 20.5 | 1.514 | 1.507 | 212.39 | 217.44 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
120 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517885 | 25.6 | 0 | 20 | 7.66 | 0.541 | 103.12 | 203.5 | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -0.29797 | 20 | 1.516 | 1.509 | 206.55 | 216.5 | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -0.62918 | 20.5 | 1.514 | 1.507 | 216.5 | 221.53 | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | 0.27911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
121 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517886 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517887 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517888 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517889 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517890 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517891 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517892 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517893 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517894 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517895 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517896 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517897 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517898 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517899 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517900 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517901 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517902 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517903 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517904 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517905 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517906 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517907 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517908 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517909 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517910 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517911 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517912 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517913 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517914 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517915 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517916 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517917 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517918 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517919 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517920 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517921 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517922 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517923 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517924 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517925 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517926 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517927 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517928 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517929 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517930 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517931 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517932 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517933 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517934 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517935 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517936 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517937 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517938 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517939 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517940 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517941 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517942 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517943 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517944 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517945 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517946 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517947 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517948 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517949 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517950 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517951 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517952 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517953 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517954 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517955 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517956 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517957 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517958 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517959 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517960 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517961 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517962 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517963 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517964 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517965 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517966 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517967 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517968 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517969 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517970 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517971 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517972 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517973 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517974 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517975 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517976 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517977 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517978 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517979 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517980 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517981 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517982 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517983 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517984 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517985 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517986 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517987 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517988 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517989 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517990 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517991 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517992 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517993 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517994 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517995 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517996 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517997 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517998 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25517999 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518000 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518001 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518002 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518003 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518004 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518005 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518006 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518007 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518008 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518009 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518010 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518011 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518012 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518013 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518014 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518015 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518016 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518017 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518018 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518019 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518020 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518021 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518022 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518023 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518024 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518025 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518026 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518027 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518028 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518029 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518030 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518031 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518032 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518033 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518034 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518035 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518036 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518037 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518038 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518039 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518040 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518041 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518042 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518043 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518044 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518045 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518046 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518047 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518048 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518049 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518050 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518051 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518052 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518053 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518054 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518055 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518056 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518057 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518058 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-22 | 2025-09-22 | 5 | 9 |  | PA25518059 | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-22 | 2025-09-22 | 5 | 9 |  |  | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-22 | 2025-09-22 | 5 | 9 |  |  | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-22 | 2025-09-22 | 5 | 9 |  |  | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-22 | 2025-09-22 | 5 | 9 |  |  | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-22 | 2025-09-22 | 5 | 9 |  |  | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-22 | 2025-09-22 | 5 | 9 |  |  | 25.6 | 0 | 20 | 7.66 | 0.541 |  |  | 100.68 | 2,32,790 | 2 | 32,790 | 0.042442 | 2372.1971 | -100 | 20 | 1.516 | 1.509 |  |  | 10.013 | 22,39,400 | 22 | 39,400 | 0.37761 | 26.5167 | -100 | 20.5 | 1.514 | 1.507 | 0 |  | 5.016 | 18,11,540 | 18 | 11,540 | 0.30321 | 16.5432 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 5 | PA25517766 | 0 | 0.1688 | 0.029949 | 0.71799 | CONFORME | Medición
2 | 5 | PA25517767 | 0 | 0.15887 | -0.16971 | 0.91743 | CONFORME | Medición
3 | 5 | PA25517768 | 0 | -0.47662 | -0.069881 | 0.31911 | CONFORME | Medición
4 | 5 | PA25517769 | 0 | 1.0922 | 1.1281 | 1.5158 | CONFORME | Medición
5 | 5 | PA25517770 | 0 | -0.15887 | -0.36937 | 1.1169 | CONFORME | Medición
6 | 5 | PA25517771 | 0 | 0.17873 | -0.16971 | 0.11966 | CONFORME | Medición
7 | 5 | PA25517772 | 0 | -0.48655 | -0.26954 | 0.71799 | CONFORME | Medición
8 | 5 | PA25517773 | 0 | -0.029789 | -0.16971 | 0.91743 | CONFORME | Medición
9 | 5 | PA25517774 | 0 | 0.10922 | -0.36937 | 0.91743 | CONFORME | Medición
10 | 5 | PA25517775 | 0 | -0.17873 | -0.56903 | 0.11966 | CONFORME | Medición
11 | 5 | PA25517776 | 0 | 0.22838 | -0.4692 | -0.27922 | CONFORME | Medición
12 | 5 | PA25517777 | 0 | 0.3376 | 0.42927 | 0.91743 | CONFORME | Medición
13 | 5 | PA25517778 | 0 | -0.11915 | 0.12978 | 0.91743 | CONFORME | Medición
14 | 5 | PA25517779 | 0 | 1.4111e-14 | 0.42927 | 1.9146 | CONFORME | Medición
15 | 5 | PA25517780 | 0 | 0.1688 | 0.029949 | 0.91743 | CONFORME | Medición
16 | 5 | PA25517781 | 0 | -0.52605 | -0.15974 | 0.39841 | CONFORME | Medición
17 | 5 | PA25517782 | 0 | -0.26799 | -0.65895 | -0.1992 | CONFORME | Medición
18 | 5 | PA25517783 | 0 | 0.36725 | -0.059904 | 1.1952 | CONFORME | Medición
19 | 5 | PA25517784 | 0 | 0.39702 | -0.15974 | 0.59761 | CONFORME | Medición
20 | 5 | PA25517785 | 0 | -0.44665 | -0.15974 | 0.99602 | CONFORME | Medición
21 | 5 | PA25517786 | 0 | -0.039702 | -0.059904 | 0.79681 | CONFORME | Medición
22 | 5 | PA25517787 | 0 | -0.15881 | -0.35942 | -0.1992 | CONFORME | Medición
23 | 5 | PA25517788 | 0 | 0.28784 | -0.45927 | 0.39841 | CONFORME | Medición
24 | 5 | PA25517789 | 0 | -0.31762 | -0.15974 | 0.99602 | CONFORME | Medición
25 | 5 | PA25517790 | 0 | 0.47643 | 0.33946 | 1.7928 | CONFORME | Medición
26 | 5 | PA25517791 | 0 | 0.079404 | 0.33946 | 1.1952 | CONFORME | Medición
27 | 5 | PA25517792 | 0 | -0.43672 | -0.45927 | 0.59761 | CONFORME | Medición
28 | 5 | PA25517793 | 0 | -0.28784 | -0.25958 | 0.79681 | CONFORME | Medición
29 | 5 | PA25517794 | 0 | 0.34739 | -0.15974 | 0.79681 | CONFORME | Medición
30 | 5 | PA25517795 | 0 | 0.10918 | 0.4393 | 0.99602 | CONFORME | Medición
31 | 5 | PA25517796 | 0 | -0.29797 | -0.54918 | 0.31911 | CONFORME | Medición
32 | 5 | PA25517797 | 0 | -0.07946 | 0.049925 | 0.91743 | CONFORME | Medición
33 | 5 | PA25517798 | 0 | 0.35757 | 0.24963 | 0.11966 | CONFORME | Medición
34 | 5 | PA25517799 | 0 | -0.92372 | -0.54918 | 0.51855 | CONFORME | Medición
35 | 5 | PA25517800 | 0 | -0.22845 | 0.34948 | 1.3163 | CONFORME | Medición
36 | 5 | PA25517801 | 0 | 0.0099325 | -0.54918 | -0.079777 | CONFORME | Medición
37 | 5 | PA25517802 | 0 | 0.089392 | -0.14978 | 1.1169 | CONFORME | Medición
38 | 5 | PA25517803 | 0 | -0.029797 | -0.049925 | 1.1169 | CONFORME | Medición
39 | 5 | PA25517804 | 0 | 0.63568 | 0.64903 | 1.3163 | CONFORME | Medición
40 | 5 | PA25517805 | 0 | 0.16885 | -0.24963 | 0.91743 | CONFORME | Medición
41 | 5 | PA25517806 | 0 | -0.45689 | -0.049925 | 1.5158 | CONFORME | Medición
42 | 5 | PA25517807 | 0 | 0.28804 | 0.34948 | 1.7152 | CONFORME | Medición
43 | 5 | PA25517808 | 0 | 0.089392 | 0.44933 | 1.5158 | CONFORME | Medición
44 | 5 | PA25517809 | 0 | -0.22845 | -0.44933 | 0.51855 | CONFORME | Medición
45 | 5 | PA25517810 | 0 | -0.67541 | 0.049925 | 0.71799 | CONFORME | Medición
46 | 5 | PA25517811 | 0 | -0.14894 | 0.17978 | 0.4985 | CONFORME | Medición
47 | 5 | PA25517812 | 0 | -0.82415 | -1.0188 | -0.2991 | CONFORME | Medición
48 | 5 | PA25517813 | 0 | 0.049648 | -0.21974 | -0.2991 | CONFORME | Medición
49 | 5 | PA25517814 | 0 | -0.25817 | -0.019976 | 0.2991 | CONFORME | Medición
50 | 5 | PA25517815 | 0 | 0.28796 | -0.4195 | 0.4985 | CONFORME | Medición
51 | 5 | PA25517816 | 0 | -0.54612 | -0.81902 | -0.69791 | CONFORME | Medición
52 | 5 | PA25517817 | 0 | -0.76457 | -0.61926 | 0.099701 | CONFORME | Medición
53 | 5 | PA25517818 | 0 | -0.049648 | -0.81902 | -0.2991 | CONFORME | Medición
54 | 5 | PA25517819 | 0 | 0.27803 | -0.4195 | 0.099701 | CONFORME | Medición
55 | 5 | PA25517820 | 0 | -0.15887 | 0.27966 | 0.89731 | CONFORME | Medición
56 | 5 | PA25517821 | 0 | -0.72485 | -0.11986 | 1.0967 | CONFORME | Medición
57 | 5 | PA25517822 | 0 | 0.44683 | -0.11986 | 1.6949 | CONFORME | Medición
58 | 5 | PA25517823 | 0 | 0.13901 | -0.11986 | 0.69791 | CONFORME | Medición
59 | 5 | PA25517824 | 0 | -0.56598 | -1.6181 | -2.2931 | CONFORME | Medición
60 | 5 | PA25517825 | 0 | -0.15887 | 0.47942 | 1.0967 | CONFORME | Medición
61 | 5 | PA25517826 | 0 | -0.26812 | -0.52931 | 1.0766 | CONFORME | Medición
62 | 5 | PA25517827 | 0 | -0.48659 | -0.52931 | 0.47847 | CONFORME | Medición
63 | 5 | PA25517828 | 0 | -0.019861 | 0.26965 | 0.67783 | CONFORME | Medición
64 | 5 | PA25517829 | 0 | -0.10924 | -0.32957 | 0.47847 | CONFORME | Medición
65 | 5 | PA25517830 | 0 | -1.4112e-14 | 0.16978 | 0.67783 | CONFORME | Medición
66 | 5 | PA25517831 | 0 | 0.089374 | 0.16978 | 1.4753 | CONFORME | Medición
67 | 5 | PA25517832 | 0 | -0.18868 | 0.069909 | 1.874 | CONFORME | Medición
68 | 5 | PA25517833 | 0 | -0.029791 | -0.52931 | 0.47847 | CONFORME | Medición
69 | 5 | PA25517834 | 0 | 0.44687 | 0.069909 | 0.67783 | CONFORME | Medición
70 | 5 | PA25517835 | 0 | 0.069513 | -0.42944 | 0.27911 | CONFORME | Medición
71 | 5 | PA25517836 | 0 | -0.049652 | 0.069909 | 1.2759 | CONFORME | Medición
72 | 5 | PA25517837 | 0 | 0.43694 | -0.42944 | 0.67783 | CONFORME | Medición
73 | 5 | PA25517838 | 0 | 0.19861 | -0.62918 | -0.31898 | CONFORME | Medición
74 | 5 | PA25517839 | 0 | 0.069513 | -0.12983 | 0.87719 | CONFORME | Medición
75 | 5 | PA25517840 | 0 | -0.29791 | -0.62918 | -0.31898 | CONFORME | Medición
76 | 5 | PA25517841 | 0 | -0.41687 | 0.059916 | 0.69791 | CONFORME | Medición
77 | 5 | PA25517842 | 0 | 0.11911 | 0.25964 | 0.89731 | CONFORME | Medición
78 | 5 | PA25517843 | 0 | -0.24814 | 0.15978 | 0.099701 | CONFORME | Medición
79 | 5 | PA25517844 | 0 | 0.23821 | -0.23966 | -0.4985 | CONFORME | Medición
80 | 5 | PA25517845 | 0 | -0.44665 | -0.1398 | 1.2961 | CONFORME | Medición
81 | 5 | PA25517846 | 0 | 0.14888 | 0.65908 | 1.2961 | CONFORME | Medición
82 | 5 | PA25517847 | 0 | -0.25806 | -0.63911 | 0.099701 | CONFORME | Medición
83 | 5 | PA25517848 | 0 | -0.29777 | -0.83883 | -0.2991 | CONFORME | Medición
84 | 5 | PA25517849 | 0 | -0.23821 | -0.93869 | 0.2991 | CONFORME | Medición
85 | 5 | PA25517850 | 0 | -0.54591 | -0.1398 | 0.2991 | CONFORME | Medición
86 | 5 | PA25517851 | 0 | -0.27792 | -0.83883 | -0.2991 | CONFORME | Medición
87 | 5 | PA25517852 | 0 | 0.17866 | -0.43938 | 0.099701 | CONFORME | Medición
88 | 5 | PA25517853 | 0 | -0.049628 | -0.93869 | 0.2991 | CONFORME | Medición
89 | 5 | PA25517854 | 0 | 0.23821 | -0.1398 | 0.099701 | CONFORME | Medición
90 | 5 | PA25517855 | 0 | 0.19851 | -0.1398 | -0.099701 | CONFORME | Medición
91 | 5 | PA25517856 | 0 | -0.19853 | 0.26965 | 0.91743 | CONFORME | Medición
92 | 5 | PA25517857 | 0 | 0.10919 | -0.52931 | 0.91743 | CONFORME | Medición
93 | 5 | PA25517858 | 0 | -0.02978 | -0.62918 | 0.11966 | CONFORME | Medición
94 | 5 | PA25517859 | 0 | -0.47647 | -0.52931 | 0.71799 | CONFORME | Medición
95 | 5 | PA25517860 | 0 | -0.35736 | -0.32957 | 1.1169 | CONFORME | Medición
96 | 5 | PA25517861 | 0 | -0.02978 | -0.12983 | 0.31911 | CONFORME | Medición
97 | 5 | PA25517862 | 0 | 0.15882 | 0.16978 | 0.11966 | CONFORME | Medición
98 | 5 | PA25517863 | 0 | 0.069486 | 0.069909 | 1.3163 | CONFORME | Medición
99 | 5 | PA25517864 | 0 | 0.40699 | -0.12983 | 0.11966 | CONFORME | Medición
100 | 5 | PA25517865 | 0 | 0.02978 | -0.029961 | 0.71799 | CONFORME | Medición
101 | 5 | PA25517866 | 0 | 1.0621 | 0.26965 | 0.91743 | CONFORME | Medición
102 | 5 | PA25517867 | 0 | -0.27794 | -0.2297 | -0.079777 | CONFORME | Medición
103 | 5 | PA25517868 | 0 | 0.31765 | 0.069909 | 0.91743 | CONFORME | Medición
104 | 5 | PA25517869 | 0 | -0.3375 | -0.82892 | -0.079777 | CONFORME | Medición
105 | 5 | PA25517870 | 0 | 0.039706 | 0.46939 | 0.91743 | CONFORME | Medición
106 | 5 | PA25517871 | 0 | 0.14899 | 0.36952 | 0.87719 | CONFORME | Medición
107 | 5 | PA25517872 | 0 | -0.48669 | -1.1285 | -0.11962 | CONFORME | Medición
108 | 5 | PA25517873 | 0 | 0.019865 | -0.52931 | 0.67783 | CONFORME | Medición
109 | 5 | PA25517874 | 0 | -0.69527 | 0.069909 | 1.0766 | CONFORME | Medición
110 | 5 | PA25517875 | 0 | -0.63568 | -0.62918 | 0.27911 | CONFORME | Medición
111 | 5 | PA25517876 | 0 | -0.47676 | -0.52931 | 0.47847 | CONFORME | Medición
112 | 5 | PA25517877 | 0 | -1.3806 | -0.92879 | -0.31898 | CONFORME | Medición
113 | 5 | PA25517878 | 0 | -1.4303 | -0.92879 | -0.31898 | CONFORME | Medición
114 | 5 | PA25517879 | 0 | -0.40723 | -0.72905 | -0.31898 | CONFORME | Medición
115 | 5 | PA25517880 | 0 | -0.62574 | -0.62918 | 0.67783 | CONFORME | Medición
116 | 5 | PA25517881 | 0 | -0.13905 | -0.62918 | -0.11962 | CONFORME | Medición
117 | 5 | PA25517882 | 0 | -0.25824 | -0.52931 | 0.27911 | CONFORME | Medición
118 | 5 | PA25517883 | 0 | -0.71514 | -1.0287 | -0.31898 | CONFORME | Medición
119 | 5 | PA25517884 | 0 | -0.64561 | -0.82892 | 0.67783 | CONFORME | Medición
120 | 5 | PA25517885 | 0 | -0.29797 | -0.62918 | 0.27911 | CONFORME | Medición
121 | 5 | PA25517886 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 5 | PA25517887 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 5 | PA25517888 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 5 | PA25517889 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 5 | PA25517890 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 5 | PA25517891 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 5 | PA25517892 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 5 | PA25517893 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 5 | PA25517894 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 5 | PA25517895 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 5 | PA25517896 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 5 | PA25517897 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 5 | PA25517898 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 5 | PA25517899 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 5 | PA25517900 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 5 | PA25517901 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 5 | PA25517902 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 5 | PA25517903 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 5 | PA25517904 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 5 | PA25517905 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 5 | PA25517906 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 5 | PA25517907 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 5 | PA25517908 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 5 | PA25517909 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 5 | PA25517910 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 5 | PA25517911 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 5 | PA25517912 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 5 | PA25517913 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 5 | PA25517914 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 5 | PA25517915 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 5 | PA25517916 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 5 | PA25517917 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 5 | PA25517918 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 5 | PA25517919 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 5 | PA25517920 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 5 | PA25517921 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 5 | PA25517922 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 5 | PA25517923 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 5 | PA25517924 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 5 | PA25517925 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 5 | PA25517926 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 5 | PA25517927 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 5 | PA25517928 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 5 | PA25517929 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 5 | PA25517930 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 5 | PA25517931 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 5 | PA25517932 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 5 | PA25517933 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 5 | PA25517934 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 5 | PA25517935 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 5 | PA25517936 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 5 | PA25517937 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 5 | PA25517938 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 5 | PA25517939 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 5 | PA25517940 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 5 | PA25517941 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 5 | PA25517942 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 5 | PA25517943 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 5 | PA25517944 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 5 | PA25517945 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 5 | PA25517946 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 5 | PA25517947 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 5 | PA25517948 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 5 | PA25517949 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 5 | PA25517950 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 5 | PA25517951 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 5 | PA25517952 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 5 | PA25517953 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 5 | PA25517954 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 5 | PA25517955 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 5 | PA25517956 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 5 | PA25517957 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 5 | PA25517958 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 5 | PA25517959 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 5 | PA25517960 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 5 | PA25517961 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 5 | PA25517962 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 5 | PA25517963 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 5 | PA25517964 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 5 | PA25517965 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 5 | PA25517966 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 5 | PA25517967 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 5 | PA25517968 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 5 | PA25517969 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 5 | PA25517970 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 5 | PA25517971 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 5 | PA25517972 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 5 | PA25517973 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 5 | PA25517974 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 5 | PA25517975 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 5 | PA25517976 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 5 | PA25517977 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 5 | PA25517978 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 5 | PA25517979 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 5 | PA25517980 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 5 | PA25517981 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 5 | PA25517982 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 5 | PA25517983 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 5 | PA25517984 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 5 | PA25517985 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 5 | PA25517986 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 5 | PA25517987 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 5 | PA25517988 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 5 | PA25517989 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 5 | PA25517990 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 5 | PA25517991 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 5 | PA25517992 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 5 | PA25517993 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 5 | PA25517994 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 5 | PA25517995 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 5 | PA25517996 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 5 | PA25517997 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 5 | PA25517998 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 5 | PA25517999 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 5 | PA25518000 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 5 | PA25518001 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 5 | PA25518002 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 5 | PA25518003 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 5 | PA25518004 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 5 | PA25518005 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 5 | PA25518006 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 5 | PA25518007 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 5 | PA25518008 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 5 | PA25518009 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 5 | PA25518010 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 5 | PA25518011 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 5 | PA25518012 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 5 | PA25518013 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 5 | PA25518014 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 5 | PA25518015 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 5 | PA25518016 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 5 | PA25518017 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 5 | PA25518018 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 5 | PA25518019 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 5 | PA25518020 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 5 | PA25518021 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 5 | PA25518022 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 5 | PA25518023 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 5 | PA25518024 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 5 | PA25518025 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 5 | PA25518026 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 5 | PA25518027 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 5 | PA25518028 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 5 | PA25518029 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 5 | PA25518030 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 5 | PA25518031 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 5 | PA25518032 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 5 | PA25518033 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 5 | PA25518034 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 5 | PA25518035 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 5 | PA25518036 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 5 | PA25518037 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 5 | PA25518038 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 5 | PA25518039 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 5 | PA25518040 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 5 | PA25518041 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 5 | PA25518042 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 5 | PA25518043 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 5 | PA25518044 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 5 | PA25518045 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 5 | PA25518046 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 5 | PA25518047 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 5 | PA25518048 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 5 | PA25518049 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 5 | PA25518050 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 5 | PA25518051 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 5 | PA25518052 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 5 | PA25518053 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 5 | PA25518054 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 5 | PA25518055 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 5 | PA25518056 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 5 | PA25518057 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 5 | PA25518058 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 5 | PA25518059 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición