## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517266 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 178.22 | 278.24 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.37849 | 21.1 | 1.391 | 1.391 | 283.43 | 293.48 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.37954 | 21.1 | 1.393 | 1.39 | 293.48 | 298.56 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 1.2356 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517267 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 177.58 | 277.2 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.77689 | 21.1 | 1.391 | 1.391 | 282.38 | 292.4 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.079904 | 21.1 | 1.393 | 1.39 | 292.4 | 297.48 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 1.2356 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517268 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 138.37 | 237.96 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.80677 | 21.1 | 1.391 | 1.391 | 243.14 | 253.16 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.079904 | 21.1 | 1.393 | 1.39 | 253.16 | 258.22 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 0.83699 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517269 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 176.19 | 276 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.58765 | 21.1 | 1.391 | 1.391 | 281.16 | 291.18 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.079904 | 21.1 | 1.393 | 1.39 | 291.18 | 296.25 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 1.0363 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517270 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 136.3 | 236.21 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.48805 | 21.1 | 1.391 | 1.391 | 241.39 | 251.48 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.77907 | 21.1 | 1.393 | 1.39 | 251.48 | 256.58 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 1.6341 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517271 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 138.16 | 237.68 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.87649 | 21.1 | 1.391 | 1.391 | 242.84 | 252.91 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.5793 | 21.1 | 1.393 | 1.39 | 252.91 | 258.02 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 1.8334 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517272 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 181.83 | 281.62 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | -0.60757 | 21.1 | 1.391 | 1.391 | 286.82 | 296.83 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | -0.019976 | 21.1 | 1.393 | 1.39 | 296.83 | 301.89 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 0.83699 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517273 | 25.6 | 0 | 21.1 | 8.315 | 0.552 | 137.29 | 237.69 | 100.4 | 2,30,990 | 2 | 30,990 | 0.041942 | 2393.8009 | 0 | 21.1 | 1.391 | 1.391 | 242.89 | 252.98 | 10.012 | 22,43,570 | 22 | 43,570 | 0.37877 | 26.433 | 0.77907 | 21.1 | 1.393 | 1.39 | 252.98 | 258.1 | 5.018 | 18,08,810 | 18 | 08,810 | 0.30245 | 16.5913 | 2.0327 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 0 | 1 | 1 | 011 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517274 | 25.6 | 0 | 21 | 8.313 | 0.539 | 91.6 | 190.98 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -1.0652 | 21.1 | 1.401 | 1.402 | 199.03 | 209.02 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.47818 | 21 | 1.4 | 1.399 | 209.02 | 214.05 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | -0.39604 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517275 | 25.6 | 0 | 21 | 8.313 | 0.539 | 132.35 | 232.28 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.51767 | 21.1 | 1.401 | 1.402 | 240.33 | 250.32 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.47818 | 21 | 1.4 | 1.399 | 250.32 | 255.39 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.39604 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517276 | 25.6 | 0 | 21 | 8.313 | 0.539 | 92.77 | 192.35 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.8661 | 21.1 | 1.401 | 1.402 | 200.4 | 210.37 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.67743 | 21 | 1.4 | 1.399 | 210.37 | 215.37 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | -0.9901 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517277 | 25.6 | 0 | 21 | 8.313 | 0.539 | 92.5 | 192.48 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.46789 | 21.1 | 1.401 | 1.402 | 200.52 | 210.59 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.31879 | 21 | 1.4 | 1.399 | 210.59 | 215.68 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.79208 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517278 | 25.6 | 0 | 21 | 8.313 | 0.539 | 93.58 | 193.89 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.13937 | 21.1 | 1.401 | 1.402 | 201.94 | 212 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.21917 | 21 | 1.4 | 1.399 | 212 | 217.06 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.19802 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517279 | 25.6 | 0 | 21 | 8.313 | 0.539 | 92.14 | 191.46 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -1.1249 | 21.1 | 1.401 | 1.402 | 199.48 | 209.48 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.37856 | 21 | 1.4 | 1.399 | 209.48 | 214.58 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.9901 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517280 | 25.6 | 0 | 21 | 8.313 | 0.539 | 92.48 | 191.88 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -1.0453 | 21.1 | 1.401 | 1.402 | 199.89 | 209.93 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.019924 | 21 | 1.4 | 1.399 | 209.93 | 215 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.39604 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517281 | 25.6 | 0 | 21 | 8.313 | 0.539 | 93 | 193.13 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.31857 | 21.1 | 1.401 | 1.402 | 201.18 | 211.28 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.61765 | 21 | 1.4 | 1.399 | 211.28 | 216.4 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 1.3861 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517282 | 25.6 | 0 | 21 | 8.313 | 0.539 | 93.3 | 192.9 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.84619 | 21.1 | 1.401 | 1.402 | 200.9 | 210.87 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.67743 | 21 | 1.4 | 1.399 | 210.87 | 215.92 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | -3.3417e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517283 | 25.6 | 0 | 21 | 8.313 | 0.539 | 132.06 | 232.24 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.26879 | 21.1 | 1.401 | 1.402 | 240.32 | 250.4 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.41841 | 21 | 1.4 | 1.399 | 250.4 | 255.49 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.79208 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517284 | 25.6 | 0 | 21 | 8.313 | 0.539 | 93.23 | 193.22 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.45794 | 21.1 | 1.401 | 1.402 | 201.24 | 211.02 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -2.5702 | 21 | 1.4 | 1.399 | 211.02 | 215.8 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | -5.3465 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
20 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517285 | 25.6 | 0 | 21 | 8.313 | 0.539 | 94.94 | 195.2 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.18915 | 21.1 | 1.401 | 1.402 | 203.29 | 213.35 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.21917 | 21 | 1.4 | 1.399 | 213.35 | 218.42 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 0.39604 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517286 | 25.6 | 0 | 21 | 8.313 | 0.539 | 95.83 | 195.82 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.45794 | 21.1 | 1.401 | 1.402 | 203.86 | 213.89 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.079697 | 21 | 1.4 | 1.399 | 213.89 | 219.02 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 1.5842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517287 | 25.6 | 0 | 21 | 8.313 | 0.539 | 92.5 | 192.14 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.80637 | 21.1 | 1.401 | 1.402 | 200.18 | 210.25 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.31879 | 21 | 1.4 | 1.399 | 210.25 | 215.38 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 1.5842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517288 | 25.6 | 0 | 21 | 8.313 | 0.539 | 96.54 | 196.69 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.29866 | 21.1 | 1.401 | 1.402 | 204.75 | 214.78 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | -0.079697 | 21 | 1.4 | 1.399 | 214.78 | 219.82 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | -0.19802 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517289 | 25.6 | 0 | 21 | 8.313 | 0.539 | 96.23 | 196.32 | 100.45 | 2,30,120 | 2 | 30,120 | 0.0417 | 2408.8729 | -0.35839 | 21.1 | 1.401 | 1.402 | 204.39 | 214.49 | 10.038 | 22,42,720 | 22 | 42,720 | 0.37853 | 26.5181 | 0.61765 | 21 | 1.4 | 1.399 | 214.49 | 219.63 | 5.05 | 18,10,810 | 18 | 10,810 | 0.303 | 16.6665 | 1.7822 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517290 | 25.6 | 0 | 21 | 8.314 | 0.553 | 165.58 | 265.96 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.1194 | 21.1 | 1.403 | 1.405 | 275.63 | 285.72 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.57815 | 21 | 1.399 | 1.401 | 285.72 | 290.84 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.7488 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517291 | 25.6 | 0 | 21 | 8.314 | 0.553 | 163.7 | 264.14 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.059701 | 21.1 | 1.403 | 1.405 | 273.8 | 283.94 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 1.0766 | 21 | 1.399 | 1.401 | 283.94 | 289.11 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 2.7424 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
27 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517292 | 25.6 | 0 | 21 | 8.314 | 0.553 | 163.42 | 264.27 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | 0.34826 | 21.1 | 1.403 | 1.405 | 273.96 | 284.09 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.97687 | 21 | 1.399 | 1.401 | 284.09 | 289.16 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 0.75517 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
28 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517293 | 25.6 | 0 | 21 | 8.314 | 0.553 | 164.47 | 264.85 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.1194 | 21.1 | 1.403 | 1.405 | 274.48 | 284.53 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.17943 | 21 | 1.399 | 1.401 | 284.53 | 289.64 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.5501 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517294 | 25.6 | 0 | 21 | 8.314 | 0.553 | 162.13 | 261.8 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.82587 | 21.1 | 1.403 | 1.405 | 271.38 | 281.45 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.37879 | 21 | 1.399 | 1.401 | 281.45 | 286.52 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 0.75517 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517295 | 25.6 | 0 | 21 | 8.314 | 0.553 | 164.38 | 264.48 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.39801 | 21.1 | 1.403 | 1.405 | 274.04 | 284.16 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.87719 | 21 | 1.399 | 1.401 | 284.16 | 289.26 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.3514 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517296 | 25.6 | 0 | 21 | 8.314 | 0.553 | 163.82 | 264.28 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.039801 | 21.1 | 1.403 | 1.405 | 273.88 | 283.9 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | -0.11962 | 21 | 1.399 | 1.401 | 283.9 | 288.96 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 0.55644 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517297 | 25.6 | 0 | 21 | 8.314 | 0.553 | 165.48 | 266.02 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | 0.039801 | 21.1 | 1.403 | 1.405 | 275.7 | 285.82 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.87719 | 21 | 1.399 | 1.401 | 285.82 | 290.92 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.3514 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
33 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517298 | 25.6 | 0 | 21 | 8.314 | 0.553 | 166.3 | 266.25 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.54726 | 21.1 | 1.403 | 1.405 | 275.84 | 285.95 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.77751 | 21 | 1.399 | 1.401 | 285.95 | 291.11 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 2.5437 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
34 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517299 | 25.6 | 0 | 21 | 8.314 | 0.553 | 164.63 | 264.94 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.18905 | 21.1 | 1.403 | 1.405 | 274.57 | 284.64 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.37879 | 21 | 1.399 | 1.401 | 284.64 | 289.75 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.5501 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
35 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517300 | 25.6 | 0 | 21 | 8.314 | 0.553 | 169.5 | 269.98 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.0199 | 21.1 | 1.403 | 1.405 | 279.52 | 289.65 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.97687 | 21 | 1.399 | 1.401 | 289.65 | 294.81 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 2.5437 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
36 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516931 | 25.6 | 0 | 21 | 8.314 | 0.553 | 167.59 | 267.8 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.28856 | 21.1 | 1.403 | 1.405 | 277.43 | 287.48 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.17943 | 21 | 1.399 | 1.401 | 287.48 | 292.54 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 0.55644 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516932 | 25.6 | 0 | 21 | 8.314 | 0.553 | 167.38 | 267.42 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.45771 | 21.1 | 1.403 | 1.405 | 277 | 287.04 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.079745 | 21 | 1.399 | 1.401 | 287.04 | 292.14 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.3514 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516933 | 25.6 | 0 | 21 | 8.314 | 0.553 | 167.65 | 268.2 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | 0.049751 | 21.1 | 1.403 | 1.405 | 277.82 | 287.9 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 0.47847 | 21 | 1.399 | 1.401 | 287.9 | 293 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.3514 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
39 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516934 | 25.6 | 1 | 21 | 8.314 | 0.553 |  |  | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -100 | 21.1 | 1.403 | 1.405 |  |  | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | -100 | 21 | 1.399 | 1.401 | 0 |  | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
40 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516935 | 25.6 | 0 | 21 | 8.314 | 0.553 | 167.28 | 267.4 | 100.5 | 2,30,400 | 2 | 30,400 | 0.041778 | 2405.5851 | -0.37811 | 21.1 | 1.403 | 1.405 | 277.02 | 287.18 | 10.032 | 22,47,990 | 22 | 47,990 | 0.38 | 26.4002 | 1.2759 | 21 | 1.399 | 1.401 | 287.18 | 292.29 | 5.032 | 18,06,970 | 18 | 06,970 | 0.30194 | 16.6658 | 1.5501 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
41 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516936 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.98 | 236.76 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.61753 | 21.1 | 1.399 | 1.395 | 245.11 | 255.09 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.4985 | 21 | 1.401 | 1.4 | 255.09 | 260.18 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.83201 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
42 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516937 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.18 | 236.8 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.77689 | 21.1 | 1.399 | 1.395 | 245.14 | 255.19 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | 0.1994 | 21 | 1.401 | 1.4 | 255.19 | 260.33 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 1.8225 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
43 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516938 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.58 | 237.33 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.64741 | 21.1 | 1.399 | 1.395 | 245.58 | 255.43 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -1.7946 | 21 | 1.401 | 1.4 | 255.43 | 260.39 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | -1.7433 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516939 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 135.98 | 235.36 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -1.0159 | 21.1 | 1.399 | 1.395 | 243.66 | 253.59 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.99701 | 21 | 1.401 | 1.4 | 253.59 | 258.62 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | -0.35658 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
45 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516940 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.8 | 237.64 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.55777 | 21.1 | 1.399 | 1.395 | 245.98 | 256.05 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | 0.3988 | 21 | 1.401 | 1.4 | 256.05 | 261.2 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 2.0206 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
46 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516941 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.24 | 236.89 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.74701 | 21.1 | 1.399 | 1.395 | 245.22 | 255.26 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | 0.099701 | 21 | 1.401 | 1.4 | 255.26 | 260.38 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 1.4263 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
47 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516942 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 138.29 | 239.04 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | 0.34861 | 21.1 | 1.399 | 1.395 | 247.39 | 257.42 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | 3.0108e-13 | 21 | 1.401 | 1.4 | 257.42 | 262.51 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.83201 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
48 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516943 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.83 | 236.4 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.82669 | 21.1 | 1.399 | 1.395 | 244.68 | 254.64 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.69791 | 21 | 1.401 | 1.4 | 254.64 | 259.69 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.03962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516944 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.82 | 237.58 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.63745 | 21.1 | 1.399 | 1.395 | 245.93 | 255.95 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.099701 | 21 | 1.401 | 1.4 | 255.95 | 261.02 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.43582 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
50 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516945 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.06 | 235.26 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -1.1952 | 21.1 | 1.399 | 1.395 | 243.53 | 253.49 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.69791 | 21 | 1.401 | 1.4 | 253.49 | 258.58 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.83201 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
51 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516946 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.59 | 236.33 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.65737 | 21.1 | 1.399 | 1.395 | 244.64 | 254.62 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.4985 | 21 | 1.401 | 1.4 | 254.62 | 259.68 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.23772 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
52 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516947 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.33 | 236.02 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.70717 | 21.1 | 1.399 | 1.395 | 244.34 | 254.43 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | 0.59821 | 21 | 1.401 | 1.4 | 254.43 | 259.54 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 1.2282 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
53 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516948 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.54 | 236.54 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.39841 | 21.1 | 1.399 | 1.395 | 244.86 | 254.83 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.59821 | 21 | 1.401 | 1.4 | 254.83 | 259.93 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 1.0301 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
54 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516949 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.58 | 237.23 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.74701 | 21.1 | 1.399 | 1.395 | 245.53 | 255.5 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.59821 | 21 | 1.401 | 1.4 | 255.5 | 260.54 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | -0.15848 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
55 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516950 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 136.71 | 236.15 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.95618 | 21.1 | 1.399 | 1.395 | 244.43 | 254.41 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | -0.4985 | 21 | 1.401 | 1.4 | 254.41 | 259.49 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.63391 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
56 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516951 | 25.6 | 0 | 21.1 | 8.311 | 0.525 | 137.34 | 237.1 | 100.4 | 2,30,700 | 2 | 30,700 | 0.041861 | 2398.4074 | -0.63745 | 21.1 | 1.399 | 1.395 | 245.4 | 255.44 | 10.03 | 22,30,770 | 22 | 30,770 | 0.37521 | 26.7314 | 0.099701 | 21 | 1.401 | 1.4 | 255.44 | 260.53 | 5.048 | 18,18,760 | 18 | 18,760 | 0.30521 | 16.5394 | 0.83201 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
57 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516952 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 78.96 | 178.49 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.78748 | 21.1 | 1.4 | 1.4 | 190.38 | 200.33 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.81738 | 21.1 | 1.402 | 1.402 | 200.33 | 205.36 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
58 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516953 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 79.07 | 178.43 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.95694 | 21.1 | 1.4 | 1.4 | 190.33 | 200.34 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.2193 | 21.1 | 1.402 | 1.402 | 200.34 | 205.41 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
59 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516954 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 78.91 | 178.74 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.48844 | 21.1 | 1.4 | 1.4 | 190.66 | 200.66 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.31898 | 21.1 | 1.402 | 1.402 | 200.66 | 205.74 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
60 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516955 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 82.47 | 181.96 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.82735 | 21.1 | 1.4 | 1.4 | 193.86 | 203.83 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.61802 | 21.1 | 1.402 | 1.402 | 203.83 | 208.9 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
61 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516956 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80 | 179.38 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.937 | 21.1 | 1.4 | 1.4 | 191.23 | 201.24 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.2193 | 21.1 | 1.402 | 1.402 | 201.24 | 206.33 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
62 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516957 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 79.84 | 179.7 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.45853 | 21.1 | 1.4 | 1.4 | 191.62 | 201.58 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.7177 | 21.1 | 1.402 | 1.402 | 201.58 | 206.54 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | -1.077 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
63 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516958 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.42 | 179.79 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.94697 | 21.1 | 1.4 | 1.4 | 191.63 | 201.63 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.31898 | 21.1 | 1.402 | 1.402 | 201.63 | 206.61 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | -0.6781 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
64 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516959 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.84 | 180.46 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.69777 | 21.1 | 1.4 | 1.4 | 192.36 | 202.39 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.019936 | 21.1 | 1.402 | 1.402 | 202.39 | 207.44 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
65 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516960 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.3 | 179.46 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -1.1563 | 21.1 | 1.4 | 1.4 | 191.37 | 201.36 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.41866 | 21.1 | 1.402 | 1.402 | 201.36 | 206.43 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
66 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516961 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.88 | 180.68 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.51834 | 21.1 | 1.4 | 1.4 | 192.6 | 202.65 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | 0.17943 | 21.1 | 1.402 | 1.402 | 202.65 | 207.76 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.9146 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
67 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516962 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.36 | 180.44 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.23923 | 21.1 | 1.4 | 1.4 | 192.39 | 202.46 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | 0.37879 | 21.1 | 1.402 | 1.402 | 202.46 | 207.52 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
68 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516963 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.09 | 179.63 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.77751 | 21.1 | 1.4 | 1.4 | 191.49 | 201.48 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.41866 | 21.1 | 1.402 | 1.402 | 201.48 | 206.49 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
69 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516964 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 78.9 | 178.33 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.88716 | 21.1 | 1.4 | 1.4 | 190.18 | 200.2 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.11962 | 21.1 | 1.402 | 1.402 | 200.2 | 205.28 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
70 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516965 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 81.23 | 181.64 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | 0.089713 | 21.1 | 1.4 | 1.4 | 193.63 | 203.66 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.019936 | 21.1 | 1.402 | 1.402 | 203.66 | 208.74 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
71 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516966 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.14 | 179.43 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -1.0267 | 21.1 | 1.4 | 1.4 | 191.3 | 201.28 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | -0.51834 | 21.1 | 1.402 | 1.402 | 201.28 | 206.32 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
72 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516967 | 25.6 | 0 | 21.1 | 8.333 | 0.534 | 80.2 | 179.82 | 100.32 | 2,30,130 | 2 | 30,130 | 0.041703 | 2405.5952 | -0.69777 | 21.1 | 1.4 | 1.4 | 191.72 | 201.78 | 10.032 | 22,50,450 | 22 | 50,450 | 0.38068 | 26.3528 | 0.27911 | 21.1 | 1.402 | 1.402 | 201.78 | 206.88 | 5.014 | 18,18,510 | 18 | 18,510 | 0.30514 | 16.4317 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
73 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516968 | 25.6 | 0 | 21 | 8.342 | 0.529 | 85.39 | 184.66 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.0269 | 21 | 1.4 | 1.4 | 189.06 | 199 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.89731 | 21 | 1.4 | 1.402 | 199 | 204 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
74 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516969 | 25.6 | 0 | 21 | 8.342 | 0.529 | 83.67 | 182.52 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.4457 | 21 | 1.4 | 1.4 | 186.89 | 196.86 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.59821 | 21 | 1.4 | 1.402 | 196.86 | 201.84 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
75 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516970 | 25.6 | 0 | 21 | 8.342 | 0.529 | 84.64 | 184.22 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.71785 | 21 | 1.4 | 1.4 | 188.6 | 198.62 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.099701 | 21 | 1.4 | 1.402 | 198.62 | 203.64 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -3.5386e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
76 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516971 | 25.6 | 0 | 21 | 8.342 | 0.529 | 84.58 | 184.12 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.75773 | 21 | 1.4 | 1.4 | 188.5 | 198.5 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.2991 | 21 | 1.4 | 1.402 | 198.5 | 203.56 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
77 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516972 | 25.6 | 0 | 21 | 8.342 | 0.529 | 83.61 | 182.84 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.0668 | 21 | 1.4 | 1.4 | 187.22 | 197.21 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.3988 | 21 | 1.4 | 1.402 | 197.21 | 202.21 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
78 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516973 | 25.6 | 0 | 21 | 8.342 | 0.529 | 85.29 | 184.53 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.0568 | 21 | 1.4 | 1.4 | 188.92 | 198.89 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.59821 | 21 | 1.4 | 1.402 | 198.89 | 203.9 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
79 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516974 | 25.6 | 0 | 21 | 8.342 | 0.529 | 86.34 | 186.2 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.43868 | 21 | 1.4 | 1.4 | 190.59 | 200.54 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.79761 | 21 | 1.4 | 1.402 | 200.54 | 205.62 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
80 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516975 | 25.6 | 0 | 21 | 8.342 | 0.529 | 85.94 | 185.18 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.0568 | 21 | 1.4 | 1.4 | 189.53 | 199.51 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.4985 | 21 | 1.4 | 1.402 | 199.51 | 204.53 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 2.1231e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
81 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516976 | 25.6 | 0 | 21 | 8.342 | 0.529 | 86.21 | 186.02 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.48853 | 21 | 1.4 | 1.4 | 190.42 | 200.48 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | 0.2991 | 21 | 1.4 | 1.402 | 200.48 | 205.58 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 1.5936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
82 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516977 | 25.6 | 0 | 21 | 8.342 | 0.529 | 87.04 | 186.41 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.92722 | 21 | 1.4 | 1.4 | 190.84 | 200.79 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.79761 | 21 | 1.4 | 1.402 | 200.79 | 205.89 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 1.5936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
83 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516978 | 25.6 | 0 | 21 | 8.342 | 0.529 | 87.7 | 186.52 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.4756 | 21 | 1.4 | 1.4 | 190.92 | 200.84 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -1.0967 | 21 | 1.4 | 1.402 | 200.84 | 205.74 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -2.3904 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
84 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516979 | 25.6 | 0 | 21 | 8.342 | 0.529 | 87.29 | 187.22 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.36889 | 21 | 1.4 | 1.4 | 191.6 | 201.6 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.2991 | 21 | 1.4 | 1.402 | 201.6 | 206.52 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -1.992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
85 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516980 | 25.6 | 0 | 21 | 8.342 | 0.529 | 87.61 | 187.59 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.31904 | 21 | 1.4 | 1.4 | 192 | 202.09 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | 0.59821 | 21 | 1.4 | 1.402 | 202.09 | 207.16 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
86 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516981 | 25.6 | 0 | 21 | 8.342 | 0.529 | 86.68 | 185.88 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -1.0967 | 21 | 1.4 | 1.4 | 190.28 | 200.29 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | -0.1994 | 21 | 1.4 | 1.402 | 200.29 | 205.28 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | -0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
87 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516982 | 25.6 | 0 | 21 | 8.342 | 0.529 | 87.23 | 186.68 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.84746 | 21 | 1.4 | 1.4 | 191.1 | 201.18 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | 0.4985 | 21 | 1.4 | 1.402 | 201.18 | 206.26 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
88 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516983 | 25.6 | 0 | 21 | 8.342 | 0.529 | 86.52 | 186.17 | 100.3 | 2,29,920 | 2 | 29,920 | 0.041644 | 2408.4845 | -0.64806 | 21 | 1.4 | 1.4 | 190.57 | 200.67 | 10.03 | 22,49,120 | 22 | 49,120 | 0.38031 | 26.3731 | 0.69791 | 21 | 1.4 | 1.402 | 200.67 | 205.74 | 5.02 | 18,15,640 | 18 | 15,640 | 0.30434 | 16.4945 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
89 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516984 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.73 | 210.69 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.3557 | 20.9 | 1.396 | 1.393 | 216.18 | 226.13 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.64903 | 20.8 | 1.396 | 1.393 | 226.13 | 231.28 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.78278 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
90 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516985 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.62 | 210.66 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.2759 | 20.9 | 1.396 | 1.393 | 216.16 | 226.11 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.64903 | 20.8 | 1.396 | 1.393 | 226.11 | 231.2 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | -0.39139 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
91 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516986 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 112.29 | 211.6 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.0068 | 20.9 | 1.396 | 1.393 | 217.09 | 227.13 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.24963 | 20.8 | 1.396 | 1.393 | 227.13 | 232.29 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.97847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
92 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516987 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.98 | 211.08 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.2161 | 20.9 | 1.396 | 1.393 | 216.53 | 226.54 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.049925 | 20.8 | 1.396 | 1.393 | 226.54 | 231.69 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.78278 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
93 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516988 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.88 | 210.63 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.565 | 20.9 | 1.396 | 1.393 | 216.09 | 226.05 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.54918 | 20.8 | 1.396 | 1.393 | 226.05 | 231.22 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 1.1742 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
94 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516989 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.93 | 211.29 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -0.95694 | 20.9 | 1.396 | 1.393 | 216.78 | 226.82 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.24963 | 20.8 | 1.396 | 1.393 | 226.82 | 231.98 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.97847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
95 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516990 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 112.52 | 212.28 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -0.55821 | 20.9 | 1.396 | 1.393 | 217.76 | 227.78 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.049925 | 20.8 | 1.396 | 1.393 | 227.78 | 232.93 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.78278 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
96 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516991 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.21 | 210.42 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.1065 | 20.9 | 1.396 | 1.393 | 215.89 | 225.86 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.44933 | 20.8 | 1.396 | 1.393 | 225.86 | 231 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.58708 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
97 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516992 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.75 | 210.86 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.2061 | 20.9 | 1.396 | 1.393 | 216.34 | 226.33 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.24963 | 20.8 | 1.396 | 1.393 | 226.33 | 231.44 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | -2.9548e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
98 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516993 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.24 | 210.49 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.0666 | 20.9 | 1.396 | 1.393 | 216 | 226 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.14978 | 20.8 | 1.396 | 1.393 | 226 | 231.16 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.97847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
99 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516994 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.94 | 211.5 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -0.75758 | 20.9 | 1.396 | 1.393 | 217 | 227.06 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.44933 | 20.8 | 1.396 | 1.393 | 227.06 | 232.25 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 1.5656 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
100 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516995 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 111.59 | 210.56 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.3457 | 20.9 | 1.396 | 1.393 | 216.03 | 226.02 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.24963 | 20.8 | 1.396 | 1.393 | 226.02 | 231.18 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.97847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
101 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516996 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 110.63 | 210.4 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -0.54825 | 20.9 | 1.396 | 1.393 | 215.9 | 225.9 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | -0.14978 | 20.8 | 1.396 | 1.393 | 225.9 | 231.05 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.78278 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
102 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516997 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 144.84 | 244.31 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -0.84729 | 20.9 | 1.396 | 1.393 | 249.78 | 259.82 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.24963 | 20.8 | 1.396 | 1.393 | 259.82 | 264.98 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | 0.97847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
103 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516998 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 108.33 | 207.59 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -1.0566 | 20.9 | 1.396 | 1.393 | 213.04 | 223.08 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.24963 | 20.8 | 1.396 | 1.393 | 223.08 | 228.1 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | -1.7613 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
104 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25516999 | 25.6 | 0 | 20.9 | 8.32 | 0.525 | 110.25 | 209.76 | 100.32 | 2,30,480 | 2 | 30,480 | 0.0418 | 2400 | -0.80742 | 20.9 | 1.396 | 1.393 | 215.28 | 225.33 | 10.015 | 22,49,160 | 22 | 49,160 | 0.38032 | 26.3329 | 0.34948 | 20.8 | 1.396 | 1.393 | 225.33 | 230.43 | 5.11 | 18,32,200 | 18 | 32,200 | 0.30894 | 16.5402 | -0.19569 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
105 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517000 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 78.34 | 177.36 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.2269 | 20.7 | 1.405 | 1.4 | 195.49 | 205.53 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.14918 | 20.6 | 1.401 | 1.397 | 205.53 | 210.59 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.75667 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
106 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517001 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 78.49 | 177.49 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.2469 | 20.7 | 1.405 | 1.4 | 195.52 | 205.53 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.44754 | 20.6 | 1.401 | 1.397 | 205.53 | 210.58 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.55755 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
107 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517002 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 80 | 179.78 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -0.46883 | 20.7 | 1.405 | 1.4 | 197.94 | 208.04 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | 0.44754 | 20.6 | 1.401 | 1.397 | 208.04 | 213.16 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 1.9514 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
108 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517003 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 84.13 | 183.15 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.2269 | 20.7 | 1.405 | 1.4 | 201.23 | 211.24 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.44754 | 20.6 | 1.401 | 1.397 | 211.24 | 216.3 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.75667 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
109 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517004 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 83.84 | 182.58 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.5062 | 20.7 | 1.405 | 1.4 | 200.6 | 210.62 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.34809 | 20.6 | 1.401 | 1.397 | 210.62 | 215.68 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.75667 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
110 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517005 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 84.52 | 183.6 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.1671 | 20.7 | 1.405 | 1.4 | 201.62 | 211.59 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.84535 | 20.6 | 1.401 | 1.397 | 211.59 | 216.61 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -0.039825 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
111 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517006 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 122.43 | 221.64 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.0374 | 20.7 | 1.405 | 1.4 | 239.74 | 249.76 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.34809 | 20.6 | 1.401 | 1.397 | 249.76 | 254.84 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 1.1549 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
112 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517007 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 83.38 | 182.23 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.3965 | 20.7 | 1.405 | 1.4 | 200.35 | 210.36 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.44754 | 20.6 | 1.401 | 1.397 | 210.36 | 215.4 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.35842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
113 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517008 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 84.21 | 183.61 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -0.84788 | 20.7 | 1.405 | 1.4 | 201.75 | 211.8 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.049727 | 20.6 | 1.401 | 1.397 | 211.8 | 216.88 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 1.1549 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
114 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517009 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 83.96 | 183.18 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.0274 | 20.7 | 1.405 | 1.4 | 201.26 | 211.3 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.14918 | 20.6 | 1.401 | 1.397 | 211.3 | 216.38 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 1.1549 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
115 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517010 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 84.43 | 183.89 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -0.78803 | 20.7 | 1.405 | 1.4 | 202.04 | 212.1 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | 0.049727 | 20.6 | 1.401 | 1.397 | 212.1 | 217.19 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 1.354 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
116 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517011 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 85.08 | 184.16 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.1671 | 20.7 | 1.405 | 1.4 | 202.24 | 212.24 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.54699 | 20.6 | 1.401 | 1.397 | 212.24 | 217.28 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.35842 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
117 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517012 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 84.23 | 183.94 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -0.53865 | 20.7 | 1.405 | 1.4 | 202.12 | 212.14 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.34809 | 20.6 | 1.401 | 1.397 | 212.14 | 217.23 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 1.354 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
118 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517013 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 155.03 | 254.24 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.0374 | 20.7 | 1.405 | 1.4 | 272.34 | 282.32 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.7459 | 20.6 | 1.401 | 1.397 | 282.32 | 287.34 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -0.039825 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
119 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517014 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 119.69 | 218.29 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.6459 | 20.7 | 1.405 | 1.4 | 236.33 | 246.3 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.84535 | 20.6 | 1.401 | 1.397 | 246.3 | 251.32 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -0.039825 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
120 | 2025-09-20 | 2025-09-20 | 4 | 1 |  | PA25517015 | 25.6 | 0 | 20.9 | 8.332 | 0.52 | 82.56 | 181.43 | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -1.3766 | 20.7 | 1.405 | 1.4 | 199.5 | 209.51 | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -0.44754 | 20.6 | 1.401 | 1.397 | 209.51 | 214.58 | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | 0.95579 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
121 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-20 | 2025-09-20 | 4 | 1 |  |  | 25.6 | 0 | 20.9 | 8.332 | 0.52 |  |  | 100.25 | 2,30,100 | 2 | 30,100 | 0.041694 | 2404.3971 | -100 | 20.7 | 1.405 | 1.4 |  |  | 10.055 | 22,52,040 | 22 | 52,040 | 0.38112 | 26.3826 | -100 | 20.6 | 1.401 | 1.397 | 0 |  | 5.022 | 18,16,060 | 18 | 16,060 | 0.30446 | 16.4947 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 4 | PA25517266 | 0 | -0.37849 | 0.37954 | 1.2356 | CONFORME | Medición
2 | 4 | PA25517267 | 0 | -0.77689 | 0.079904 | 1.2356 | CONFORME | Medición
3 | 4 | PA25517268 | 0 | -0.80677 | 0.079904 | 0.83699 | CONFORME | Medición
4 | 4 | PA25517269 | 0 | -0.58765 | 0.079904 | 1.0363 | CONFORME | Medición
5 | 4 | PA25517270 | 0 | -0.48805 | 0.77907 | 1.6341 | CONFORME | Medición
6 | 4 | PA25517271 | 0 | -0.87649 | 0.5793 | 1.8334 | CONFORME | Medición
7 | 4 | PA25517272 | 0 | -0.60757 | -0.019976 | 0.83699 | CONFORME | Medición
8 | 4 | PA25517273 | 0 | 0 | 0.77907 | 2.0327 | CONFORME | Medición
9 | 4 | PA25517274 | 0 | -1.0652 | -0.47818 | -0.39604 | CONFORME | Medición
10 | 4 | PA25517275 | 0 | -0.51767 | -0.47818 | 0.39604 | CONFORME | Medición
11 | 4 | PA25517276 | 0 | -0.8661 | -0.67743 | -0.9901 | CONFORME | Medición
12 | 4 | PA25517277 | 0 | -0.46789 | 0.31879 | 0.79208 | CONFORME | Medición
13 | 4 | PA25517278 | 0 | -0.13937 | 0.21917 | 0.19802 | CONFORME | Medición
14 | 4 | PA25517279 | 0 | -1.1249 | -0.37856 | 0.9901 | CONFORME | Medición
15 | 4 | PA25517280 | 0 | -1.0453 | 0.019924 | 0.39604 | CONFORME | Medición
16 | 4 | PA25517281 | 0 | -0.31857 | 0.61765 | 1.3861 | CONFORME | Medición
17 | 4 | PA25517282 | 0 | -0.84619 | -0.67743 | -3.3417e-13 | CONFORME | Medición
18 | 4 | PA25517283 | 0 | -0.26879 | 0.41841 | 0.79208 | CONFORME | Medición
19 | 4 | PA25517284 | 0 | -0.45794 | -2.5702 | -5.3465 | NO CONFORME | Medición
20 | 4 | PA25517285 | 0 | -0.18915 | 0.21917 | 0.39604 | CONFORME | Medición
21 | 4 | PA25517286 | 0 | -0.45794 | -0.079697 | 1.5842 | CONFORME | Medición
22 | 4 | PA25517287 | 0 | -0.80637 | 0.31879 | 1.5842 | CONFORME | Medición
23 | 4 | PA25517288 | 0 | -0.29866 | -0.079697 | -0.19802 | CONFORME | Medición
24 | 4 | PA25517289 | 0 | -0.35839 | 0.61765 | 1.7822 | CONFORME | Medición
25 | 4 | PA25517290 | 0 | -0.1194 | 0.57815 | 1.7488 | CONFORME | Medición
26 | 4 | PA25517291 | 0 | -0.059701 | 1.0766 | 2.7424 | CONFORME | Medición
27 | 4 | PA25517292 | 0 | 0.34826 | 0.97687 | 0.75517 | CONFORME | Medición
28 | 4 | PA25517293 | 0 | -0.1194 | 0.17943 | 1.5501 | CONFORME | Medición
29 | 4 | PA25517294 | 0 | -0.82587 | 0.37879 | 0.75517 | CONFORME | Medición
30 | 4 | PA25517295 | 0 | -0.39801 | 0.87719 | 1.3514 | CONFORME | Medición
31 | 4 | PA25517296 | 0 | -0.039801 | -0.11962 | 0.55644 | CONFORME | Medición
32 | 4 | PA25517297 | 0 | 0.039801 | 0.87719 | 1.3514 | CONFORME | Medición
33 | 4 | PA25517298 | 0 | -0.54726 | 0.77751 | 2.5437 | CONFORME | Medición
34 | 4 | PA25517299 | 0 | -0.18905 | 0.37879 | 1.5501 | CONFORME | Medición
35 | 4 | PA25517300 | 0 | -0.0199 | 0.97687 | 2.5437 | CONFORME | Medición
36 | 4 | PA25516931 | 0 | -0.28856 | 0.17943 | 0.55644 | CONFORME | Medición
37 | 4 | PA25516932 | 0 | -0.45771 | 0.079745 | 1.3514 | CONFORME | Medición
38 | 4 | PA25516933 | 0 | 0.049751 | 0.47847 | 1.3514 | CONFORME | Medición
39 | 4 | PA25516934 | 1 | -100 | -100 | -100 | NO CONFORME | Daño físico
40 | 4 | PA25516935 | 0 | -0.37811 | 1.2759 | 1.5501 | CONFORME | Medición
41 | 4 | PA25516936 | 0 | -0.61753 | -0.4985 | 0.83201 | CONFORME | Medición
42 | 4 | PA25516937 | 0 | -0.77689 | 0.1994 | 1.8225 | CONFORME | Medición
43 | 4 | PA25516938 | 0 | -0.64741 | -1.7946 | -1.7433 | CONFORME | Medición
44 | 4 | PA25516939 | 0 | -1.0159 | -0.99701 | -0.35658 | CONFORME | Medición
45 | 4 | PA25516940 | 0 | -0.55777 | 0.3988 | 2.0206 | CONFORME | Medición
46 | 4 | PA25516941 | 0 | -0.74701 | 0.099701 | 1.4263 | CONFORME | Medición
47 | 4 | PA25516942 | 0 | 0.34861 | 3.0108e-13 | 0.83201 | CONFORME | Medición
48 | 4 | PA25516943 | 0 | -0.82669 | -0.69791 | 0.03962 | CONFORME | Medición
49 | 4 | PA25516944 | 0 | -0.63745 | -0.099701 | 0.43582 | CONFORME | Medición
50 | 4 | PA25516945 | 0 | -1.1952 | -0.69791 | 0.83201 | CONFORME | Medición
51 | 4 | PA25516946 | 0 | -0.65737 | -0.4985 | 0.23772 | CONFORME | Medición
52 | 4 | PA25516947 | 0 | -0.70717 | 0.59821 | 1.2282 | CONFORME | Medición
53 | 4 | PA25516948 | 0 | -0.39841 | -0.59821 | 1.0301 | CONFORME | Medición
54 | 4 | PA25516949 | 0 | -0.74701 | -0.59821 | -0.15848 | CONFORME | Medición
55 | 4 | PA25516950 | 0 | -0.95618 | -0.4985 | 0.63391 | CONFORME | Medición
56 | 4 | PA25516951 | 0 | -0.63745 | 0.099701 | 0.83201 | CONFORME | Medición
57 | 4 | PA25516952 | 0 | -0.78748 | -0.81738 | 0.31911 | CONFORME | Medición
58 | 4 | PA25516953 | 0 | -0.95694 | -0.2193 | 1.1169 | CONFORME | Medición
59 | 4 | PA25516954 | 0 | -0.48844 | -0.31898 | 1.3163 | CONFORME | Medición
60 | 4 | PA25516955 | 0 | -0.82735 | -0.61802 | 1.1169 | CONFORME | Medición
61 | 4 | PA25516956 | 0 | -0.937 | -0.2193 | 1.5158 | CONFORME | Medición
62 | 4 | PA25516957 | 0 | -0.45853 | -0.7177 | -1.077 | CONFORME | Medición
63 | 4 | PA25516958 | 0 | -0.94697 | -0.31898 | -0.6781 | CONFORME | Medición
64 | 4 | PA25516959 | 0 | -0.69777 | -0.019936 | 0.71799 | CONFORME | Medición
65 | 4 | PA25516960 | 0 | -1.1563 | -0.41866 | 1.1169 | CONFORME | Medición
66 | 4 | PA25516961 | 0 | -0.51834 | 0.17943 | 1.9146 | CONFORME | Medición
67 | 4 | PA25516962 | 0 | -0.23923 | 0.37879 | 0.91743 | CONFORME | Medición
68 | 4 | PA25516963 | 0 | -0.77751 | -0.41866 | -0.079777 | CONFORME | Medición
69 | 4 | PA25516964 | 0 | -0.88716 | -0.11962 | 1.3163 | CONFORME | Medición
70 | 4 | PA25516965 | 0 | 0.089713 | -0.019936 | 1.3163 | CONFORME | Medición
71 | 4 | PA25516966 | 0 | -1.0267 | -0.51834 | 0.51855 | CONFORME | Medición
72 | 4 | PA25516967 | 0 | -0.69777 | 0.27911 | 1.7152 | CONFORME | Medición
73 | 4 | PA25516968 | 0 | -1.0269 | -0.89731 | -0.39841 | CONFORME | Medición
74 | 4 | PA25516969 | 0 | -1.4457 | -0.59821 | -0.79681 | CONFORME | Medición
75 | 4 | PA25516970 | 0 | -0.71785 | -0.099701 | -3.5386e-13 | CONFORME | Medición
76 | 4 | PA25516971 | 0 | -0.75773 | -0.2991 | 0.79681 | CONFORME | Medición
77 | 4 | PA25516972 | 0 | -1.0668 | -0.3988 | -0.39841 | CONFORME | Medición
78 | 4 | PA25516973 | 0 | -1.0568 | -0.59821 | -0.1992 | CONFORME | Medición
79 | 4 | PA25516974 | 0 | -0.43868 | -0.79761 | 1.1952 | CONFORME | Medición
80 | 4 | PA25516975 | 0 | -1.0568 | -0.4985 | 2.1231e-13 | CONFORME | Medición
81 | 4 | PA25516976 | 0 | -0.48853 | 0.2991 | 1.5936 | CONFORME | Medición
82 | 4 | PA25516977 | 0 | -0.92722 | -0.79761 | 1.5936 | CONFORME | Medición
83 | 4 | PA25516978 | 0 | -1.4756 | -1.0967 | -2.3904 | CONFORME | Medición
84 | 4 | PA25516979 | 0 | -0.36889 | -0.2991 | -1.992 | CONFORME | Medición
85 | 4 | PA25516980 | 0 | -0.31904 | 0.59821 | 0.99602 | CONFORME | Medición
86 | 4 | PA25516981 | 0 | -1.0967 | -0.1994 | -0.59761 | CONFORME | Medición
87 | 4 | PA25516982 | 0 | -0.84746 | 0.4985 | 1.1952 | CONFORME | Medición
88 | 4 | PA25516983 | 0 | -0.64806 | 0.69791 | 0.99602 | CONFORME | Medición
89 | 4 | PA25516984 | 0 | -1.3557 | -0.64903 | 0.78278 | CONFORME | Medición
90 | 4 | PA25516985 | 0 | -1.2759 | -0.64903 | -0.39139 | CONFORME | Medición
91 | 4 | PA25516986 | 0 | -1.0068 | 0.24963 | 0.97847 | CONFORME | Medición
92 | 4 | PA25516987 | 0 | -1.2161 | -0.049925 | 0.78278 | CONFORME | Medición
93 | 4 | PA25516988 | 0 | -1.565 | -0.54918 | 1.1742 | CONFORME | Medición
94 | 4 | PA25516989 | 0 | -0.95694 | 0.24963 | 0.97847 | CONFORME | Medición
95 | 4 | PA25516990 | 0 | -0.55821 | 0.049925 | 0.78278 | CONFORME | Medición
96 | 4 | PA25516991 | 0 | -1.1065 | -0.44933 | 0.58708 | CONFORME | Medición
97 | 4 | PA25516992 | 0 | -1.2061 | -0.24963 | -2.9548e-13 | CONFORME | Medición
98 | 4 | PA25516993 | 0 | -1.0666 | -0.14978 | 0.97847 | CONFORME | Medición
99 | 4 | PA25516994 | 0 | -0.75758 | 0.44933 | 1.5656 | CONFORME | Medición
100 | 4 | PA25516995 | 0 | -1.3457 | -0.24963 | 0.97847 | CONFORME | Medición
101 | 4 | PA25516996 | 0 | -0.54825 | -0.14978 | 0.78278 | CONFORME | Medición
102 | 4 | PA25516997 | 0 | -0.84729 | 0.24963 | 0.97847 | CONFORME | Medición
103 | 4 | PA25516998 | 0 | -1.0566 | 0.24963 | -1.7613 | CONFORME | Medición
104 | 4 | PA25516999 | 0 | -0.80742 | 0.34948 | -0.19569 | CONFORME | Medición
105 | 4 | PA25517000 | 0 | -1.2269 | -0.14918 | 0.75667 | CONFORME | Medición
106 | 4 | PA25517001 | 0 | -1.2469 | -0.44754 | 0.55755 | CONFORME | Medición
107 | 4 | PA25517002 | 0 | -0.46883 | 0.44754 | 1.9514 | CONFORME | Medición
108 | 4 | PA25517003 | 0 | -1.2269 | -0.44754 | 0.75667 | CONFORME | Medición
109 | 4 | PA25517004 | 0 | -1.5062 | -0.34809 | 0.75667 | CONFORME | Medición
110 | 4 | PA25517005 | 0 | -1.1671 | -0.84535 | -0.039825 | CONFORME | Medición
111 | 4 | PA25517006 | 0 | -1.0374 | -0.34809 | 1.1549 | CONFORME | Medición
112 | 4 | PA25517007 | 0 | -1.3965 | -0.44754 | 0.35842 | CONFORME | Medición
113 | 4 | PA25517008 | 0 | -0.84788 | -0.049727 | 1.1549 | CONFORME | Medición
114 | 4 | PA25517009 | 0 | -1.0274 | -0.14918 | 1.1549 | CONFORME | Medición
115 | 4 | PA25517010 | 0 | -0.78803 | 0.049727 | 1.354 | CONFORME | Medición
116 | 4 | PA25517011 | 0 | -1.1671 | -0.54699 | 0.35842 | CONFORME | Medición
117 | 4 | PA25517012 | 0 | -0.53865 | -0.34809 | 1.354 | CONFORME | Medición
118 | 4 | PA25517013 | 0 | -1.0374 | -0.7459 | -0.039825 | CONFORME | Medición
119 | 4 | PA25517014 | 0 | -1.6459 | -0.84535 | -0.039825 | CONFORME | Medición
120 | 4 | PA25517015 | 0 | -1.3766 | -0.44754 | 0.95579 | CONFORME | Medición
121 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición