## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517301 | 25.6 | 0 | 20 | 8.512 | 0.847 | 143.25 | 243.57 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.36746 | 20 | 1.107 | 1.103 | 247.29 | 257.29 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.089919 | 20.2 | 1.103 | 1.098 | 257.29 | 262.29 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.11986 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
2 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517302 | 25.6 | 0 | 20 | 8.512 | 0.847 | 144.08 | 243.97 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.79452 | 20 | 1.107 | 1.103 | 247.63 | 257.78 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | 1.4087 | 20.2 | 1.103 | 1.098 | 257.78 | 262.9 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 2.2773 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517303 | 25.6 | 0 | 20 | 8.512 | 0.847 | 144.38 | 245.15 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | 0.079452 | 20 | 1.107 | 1.103 | 248.8 | 258.91 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | 1.0091 | 20.2 | 1.103 | 1.098 | 258.91 | 264 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 1.678 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
4 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517304 | 25.6 | 0 | 20 | 8.512 | 0.847 | 146.3 | 246.1 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.8839 | 20 | 1.107 | 1.103 | 249.74 | 259.61 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -1.3888 | 20.2 | 1.103 | 1.098 | 259.61 | 264.6 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.31962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
5 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517305 | 25.6 | 0 | 20 | 8.512 | 0.847 | 148.19 | 248.3 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.57603 | 20 | 1.107 | 1.103 | 251.91 | 261.96 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | 0.40963 | 20.2 | 1.103 | 1.098 | 261.96 | 266.98 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.27966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517306 | 25.6 | 0 | 20 | 8.512 | 0.847 | 145.55 | 245.55 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.68527 | 20 | 1.107 | 1.103 | 249.25 | 259.25 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.089919 | 20.2 | 1.103 | 1.098 | 259.25 | 264.29 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.67918 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517307 | 25.6 | 0 | 20 | 8.512 | 0.847 | 145.82 | 246.4 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.10925 | 20 | 1.107 | 1.103 | 250.04 | 260.03 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.18983 | 20.2 | 1.103 | 1.098 | 260.03 | 265.08 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.87895 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517308 | 25.6 | 0 | 20 | 8.512 | 0.847 | 143.8 | 243.89 | 100.69 | 2,31,971 | 2 | 31,971 | 0.042214 | 2385.2182 | -0.59589 | 20 | 1.107 | 1.103 | 247.54 | 257.54 | 10.009 | 22,53,131 | 22 | 53,131 | 0.38143 | 26.2411 | -0.089919 | 20.2 | 1.103 | 1.098 | 257.54 | 262.58 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.67918 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517309 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 89.57 | 189.9 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.31793 | 20.4 | 1.098 | 1.088 | 195.56 | 205.55 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.17986 | 20.5 | 1.096 | 1.086 | 205.55 | 210.55 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517310 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 89.4 | 190.2 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | 0.14903 | 20.4 | 1.098 | 1.088 | 195.88 | 205.9 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | 0.1199 | 20.5 | 1.096 | 1.086 | 205.9 | 210.94 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | 0.71942 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
11 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517311 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 89.94 | 189.74 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.84451 | 20.4 | 1.098 | 1.088 | 195.43 | 205.4 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.3797 | 20.5 | 1.096 | 1.086 | 205.4 | 210.4 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
12 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517312 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 168.75 | 268.84 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.55638 | 20.4 | 1.098 | 1.088 | 274.51 | 284.49 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.27978 | 20.5 | 1.096 | 1.086 | 284.49 | 289.49 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
13 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517313 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 131.41 | 231.75 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.308 | 20.4 | 1.098 | 1.088 | 237.45 | 247.41 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.47962 | 20.5 | 1.096 | 1.086 | 247.41 | 252.4 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
14 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517314 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 126.03 | 225.53 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -1.1426 | 20.4 | 1.098 | 1.088 | 231.15 | 241.01 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -1.4788 | 20.5 | 1.096 | 1.086 | 241.01 | 245.95 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -1.279 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
15 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517315 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 127.75 | 227.55 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.84451 | 20.4 | 1.098 | 1.088 | 233.2 | 243.12 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.8793 | 20.5 | 1.096 | 1.086 | 243.12 | 248.11 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
16 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517316 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 128.38 | 228.7 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.32787 | 20.4 | 1.098 | 1.088 | 234.39 | 244.37 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.27978 | 20.5 | 1.096 | 1.086 | 244.37 | 249.37 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
17 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517317 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 126.75 | 226.64 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.75509 | 20.4 | 1.098 | 1.088 | 232.31 | 242.3 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.17986 | 20.5 | 1.096 | 1.086 | 242.3 | 247.35 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | 0.91926 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
18 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517318 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 130.4 | 230.79 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.25832 | 20.4 | 1.098 | 1.088 | 236.49 | 246.45 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.47962 | 20.5 | 1.096 | 1.086 | 246.45 | 251.44 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
19 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517319 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 129.43 | 229.04 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -1.0333 | 20.4 | 1.098 | 1.088 | 234.66 | 244.65 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.17986 | 20.5 | 1.096 | 1.086 | 244.65 | 249.66 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | 0.1199 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517320 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 93.82 | 193.87 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.59613 | 20.4 | 1.098 | 1.088 | 199.51 | 209.45 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.67946 | 20.5 | 1.096 | 1.086 | 209.45 | 214.45 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
21 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517321 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 92.82 | 193.06 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.40735 | 20.4 | 1.098 | 1.088 | 198.69 | 208.69 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.079936 | 20.5 | 1.096 | 1.086 | 208.69 | 213.69 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
22 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517322 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 94.51 | 195.36 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | 0.19871 | 20.4 | 1.098 | 1.088 | 201.01 | 211.02 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | 0.019984 | 20.5 | 1.096 | 1.086 | 211.02 | 216.05 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | 0.51958 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
23 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517323 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 95.89 | 195.95 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | -0.58619 | 20.4 | 1.098 | 1.088 | 201.58 | 211.55 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | -0.3797 | 20.5 | 1.096 | 1.086 | 211.55 | 216.55 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
24 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517324 | 25.6 | 0 | 20.2 | 8.509 | 0.849 | 94.18 | 195.24 | 100.65 | 2,32,441 | 2 | 32,441 | 0.042345 | 2376.9196 | 0.40735 | 20.4 | 1.098 | 1.088 | 200.89 | 210.92 | 10.008 | 22,55,141 | 22 | 55,141 | 0.38198 | 26.2001 | 0.21982 | 20.5 | 1.096 | 1.086 | 210.92 | 215.95 | 5.004 | 18,16,151 | 18 | 16,151 | 0.30449 | 16.4342 | 0.51958 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
25 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517325 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 89.05 | 188.68 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.97406 | 20.5 | 1.096 | 1.086 | 192.85 | 202.74 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -1.2383 | 20.6 | 1.097 | 1.092 | 202.74 | 207.73 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.33952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
26 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517326 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 87.55 | 187.5 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.656 | 20.5 | 1.096 | 1.086 | 191.66 | 201.64 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.33952 | 20.6 | 1.097 | 1.092 | 201.64 | 206.68 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 0.65908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
27 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517327 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 87.38 | 187.43 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.5566 | 20.5 | 1.096 | 1.086 | 191.59 | 201.56 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.43938 | 20.6 | 1.097 | 1.092 | 201.56 | 206.58 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 0.25964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517328 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 87.66 | 188.16 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.10933 | 20.5 | 1.096 | 1.086 | 192.35 | 202.28 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.83883 | 20.6 | 1.097 | 1.092 | 202.28 | 207.28 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.1398 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
29 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517329 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 87.16 | 186.84 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.92436 | 20.5 | 1.096 | 1.086 | 191.01 | 201.01 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.1398 | 20.6 | 1.097 | 1.092 | 201.01 | 206.07 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 1.0585 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517330 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 88.4 | 188.92 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.089454 | 20.5 | 1.096 | 1.086 | 193.13 | 203.1 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.43938 | 20.6 | 1.097 | 1.092 | 203.1 | 208.14 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 0.65908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517331 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 90.13 | 190.82 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | 0.079515 | 20.5 | 1.096 | 1.086 | 195.04 | 205 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.53925 | 20.6 | 1.097 | 1.092 | 205 | 210 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.1398 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517332 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 88.96 | 189.12 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.44727 | 20.5 | 1.096 | 1.086 | 193.3 | 203.28 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.33952 | 20.6 | 1.097 | 1.092 | 203.28 | 208.28 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.1398 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
33 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517333 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 87.26 | 187.69 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.17891 | 20.5 | 1.096 | 1.086 | 191.89 | 201.86 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.43938 | 20.6 | 1.097 | 1.092 | 201.86 | 206.93 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 1.2582 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
34 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517334 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 89.1 | 188.85 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.85479 | 20.5 | 1.096 | 1.086 | 193.03 | 202.95 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.93869 | 20.6 | 1.097 | 1.092 | 202.95 | 207.92 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.73897 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
35 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517335 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 92.02 | 192.56 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.069576 | 20.5 | 1.096 | 1.086 | 196.75 | 206.72 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.43938 | 20.6 | 1.097 | 1.092 | 206.72 | 211.73 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 0.059916 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
36 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517336 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 91.03 | 190.58 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -1.0536 | 20.5 | 1.096 | 1.086 | 194.72 | 204.65 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.83883 | 20.6 | 1.097 | 1.092 | 204.65 | 209.62 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.73897 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
37 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517337 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 90.86 | 191.35 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.11927 | 20.5 | 1.096 | 1.086 | 195.53 | 205.5 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.43938 | 20.6 | 1.097 | 1.092 | 205.5 | 210.53 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 0.45936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517338 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 89.75 | 190.28 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.079515 | 20.5 | 1.096 | 1.086 | 194.45 | 204.43 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.33952 | 20.6 | 1.097 | 1.092 | 204.43 | 209.44 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | 0.059916 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
39 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517339 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 88.7 | 188.55 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.75539 | 20.5 | 1.096 | 1.086 | 192.7 | 202.65 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.63911 | 20.6 | 1.097 | 1.092 | 202.65 | 207.62 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.73897 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
40 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517340 | 25.6 | 0 | 20.1 | 8.511 | 0.837 | 89.05 | 189.17 | 100.61 | 2,31,961 | 2 | 31,961 | 0.042211 | 2383.48 | -0.48703 | 20.5 | 1.096 | 1.086 | 193.35 | 203.31 | 10.014 | 22,51,911 | 22 | 51,911 | 0.38109 | 26.2775 | -0.53925 | 20.6 | 1.097 | 1.092 | 203.31 | 208.3 | 5.007 | 18,14,491 | 18 | 14,491 | 0.30403 | 16.469 | -0.33952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
41 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517341 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 237.36 | 336.77 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -1.281 | 20.8 | 1.099 | 1.09 | 340.54 | 350.43 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -1.1889 | 20.8 | 1.095 | 1.086 | 350.43 | 355.36 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -1.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
42 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517342 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 239.1 | 339.29 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.50645 | 20.8 | 1.099 | 1.09 | 343.05 | 352.93 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -1.2888 | 20.8 | 1.095 | 1.086 | 352.93 | 357.89 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.8991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
43 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517343 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 239.49 | 339.66 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.52632 | 20.8 | 1.099 | 1.09 | 343.48 | 353.43 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.58947 | 20.8 | 1.095 | 1.086 | 353.43 | 358.43 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517344 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 242.4 | 343.05 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.049652 | 20.8 | 1.099 | 1.09 | 346.85 | 356.86 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | 0.009991 | 20.8 | 1.095 | 1.086 | 356.86 | 361.86 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
45 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517345 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 240.92 | 341.44 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.17875 | 20.8 | 1.099 | 1.09 | 345.2 | 355.2 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.089919 | 20.8 | 1.095 | 1.086 | 355.2 | 360.25 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | 0.8991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
46 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517346 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 239.65 | 339.5 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.84409 | 20.8 | 1.099 | 1.09 | 343.26 | 353.17 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.98911 | 20.8 | 1.095 | 1.086 | 353.17 | 358.17 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
47 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517347 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 240.75 | 340.76 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.6852 | 20.8 | 1.099 | 1.09 | 344.55 | 354.54 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.18983 | 20.8 | 1.095 | 1.086 | 354.54 | 359.55 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | 0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
48 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517348 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 240.87 | 340.8 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.76465 | 20.8 | 1.099 | 1.09 | 344.59 | 354.58 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.18983 | 20.8 | 1.095 | 1.086 | 354.58 | 359.6 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | 0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517349 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 238.51 | 339 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.20854 | 20.8 | 1.099 | 1.09 | 342.75 | 352.68 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.78929 | 20.8 | 1.095 | 1.086 | 352.68 | 357.69 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | 0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
50 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517350 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 238.95 | 338.4 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -1.2413 | 20.8 | 1.099 | 1.09 | 342.11 | 351.97 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -1.4887 | 20.8 | 1.095 | 1.086 | 351.97 | 356.9 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -1.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
51 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517351 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 239.1 | 339.69 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.10924 | 20.8 | 1.099 | 1.09 | 343.48 | 353.36 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -1.2888 | 20.8 | 1.095 | 1.086 | 353.36 | 358.2 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -3.2967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
52 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517352 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 236.44 | 336.02 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -1.1122 | 20.8 | 1.099 | 1.09 | 339.76 | 349.74 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.28974 | 20.8 | 1.095 | 1.086 | 349.74 | 354.74 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
53 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517353 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 239.1 | 338.83 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.96326 | 20.8 | 1.099 | 1.09 | 342.6 | 352.53 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.78929 | 20.8 | 1.095 | 1.086 | 352.53 | 357.52 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
54 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517354 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 239.85 | 339.6 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.9434 | 20.8 | 1.099 | 1.09 | 343.38 | 353.36 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.28974 | 20.8 | 1.095 | 1.086 | 353.36 | 358.4 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | 0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517355 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 235.36 | 334.38 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -1.6683 | 20.8 | 1.099 | 1.09 | 338.1 | 348.04 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.68938 | 20.8 | 1.095 | 1.086 | 348.04 | 353.04 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
56 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517356 | 25.6 | 0 | 20.3 | 8.513 | 0.847 | 283.29 | 383.69 | 100.7 | 2,32,201 | 2 | 32,201 | 0.042278 | 2381.8503 | -0.29791 | 20.8 | 1.099 | 1.09 | 387.49 | 397.45 | 10.009 | 22,53,121 | 22 | 53,121 | 0.38142 | 26.2412 | -0.48956 | 20.8 | 1.095 | 1.086 | 397.45 | 402.45 | 5.005 | 18,16,161 | 18 | 16,161 | 0.30449 | 16.4374 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
57 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517357 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 127.39 | 227.25 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.90305 | 20.4 | 1.097 | 1.088 | 232.8 | 242.7 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -1.0989 | 20.6 | 1.098 | 1.089 | 242.7 | 247.68 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.51938 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
58 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517358 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 130.55 | 229.88 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -1.429 | 20.4 | 1.097 | 1.088 | 235.45 | 245.35 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -1.0989 | 20.6 | 1.098 | 1.089 | 245.35 | 250.26 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -1.9177 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
59 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517359 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 129.24 | 229.35 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.65496 | 20.4 | 1.097 | 1.088 | 234.93 | 244.8 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -1.3986 | 20.6 | 1.098 | 1.089 | 244.8 | 249.79 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.31962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
60 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517360 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 131.34 | 231.28 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.82366 | 20.4 | 1.097 | 1.088 | 236.85 | 246.87 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | 0.0999 | 20.6 | 1.098 | 1.089 | 246.87 | 251.95 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 1.4782 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
61 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517361 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 133.15 | 233.4 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.51603 | 20.4 | 1.097 | 1.088 | 239 | 248.93 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.7992 | 20.6 | 1.098 | 1.089 | 248.93 | 253.94 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.079904 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
62 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517362 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 86.55 | 186.5 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.81373 | 20.4 | 1.097 | 1.088 | 192.1 | 202 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -1.0989 | 20.6 | 1.098 | 1.089 | 202 | 206.98 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.51938 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
63 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517363 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 87.59 | 187.94 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.41679 | 20.4 | 1.097 | 1.088 | 193.56 | 203.55 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.1998 | 20.6 | 1.098 | 1.089 | 203.55 | 208.6 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.87895 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
64 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517364 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 86.24 | 186.65 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.35725 | 20.4 | 1.097 | 1.088 | 192.3 | 202.25 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.5994 | 20.6 | 1.098 | 1.089 | 202.25 | 207.25 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.11986 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
65 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517365 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 86.4 | 186.65 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.51603 | 20.4 | 1.097 | 1.088 | 192.29 | 202.23 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.6993 | 20.6 | 1.098 | 1.089 | 202.23 | 207.24 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.079904 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
66 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517366 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 127.44 | 227.89 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.31755 | 20.4 | 1.097 | 1.088 | 233.51 | 243.49 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.2997 | 20.6 | 1.098 | 1.089 | 243.49 | 248.49 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.11986 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
67 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517367 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 89.63 | 190.05 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.34733 | 20.4 | 1.097 | 1.088 | 195.65 | 205.58 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.7992 | 20.6 | 1.098 | 1.089 | 205.58 | 210.57 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.31962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
68 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517368 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 86.49 | 186.05 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -1.2008 | 20.4 | 1.097 | 1.088 | 191.6 | 201.52 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.8991 | 20.6 | 1.098 | 1.089 | 201.52 | 206.55 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.47942 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
69 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517369 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 87.38 | 187.97 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.17862 | 20.4 | 1.097 | 1.088 | 193.6 | 203.55 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.5994 | 20.6 | 1.098 | 1.089 | 203.55 | 208.55 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.11986 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
70 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517370 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 86.05 | 185.41 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -1.3992 | 20.4 | 1.097 | 1.088 | 190.95 | 200.85 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -1.0989 | 20.6 | 1.098 | 1.089 | 200.85 | 205.81 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.9189 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
71 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517371 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 87.45 | 187.58 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.63511 | 20.4 | 1.097 | 1.088 | 193.2 | 203.15 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.5994 | 20.6 | 1.098 | 1.089 | 203.15 | 208.15 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | -0.11986 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
72 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517372 | 25.6 | 0 | 20.3 | 8.51 | 0.851 | 86.9 | 187.24 | 100.77 | 2,32,017 | 2 | 32,017 | 0.042227 | 2386.391 | -0.42671 | 20.4 | 1.097 | 1.088 | 192.84 | 202.77 | 10.01 | 22,50,961 | 22 | 50,961 | 0.38082 | 26.2852 | -0.7992 | 20.6 | 1.098 | 1.089 | 202.77 | 207.79 | 5.006 | 18,15,171 | 18 | 15,171 | 0.30421 | 16.4555 | 0.27966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
73 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517373 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 82.1 | 181.99 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.74523 | 20.7 | 1.095 | 1.086 | 186.02 | 195.97 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.58947 | 20.7 | 1.096 | 1.087 | 195.97 | 200.96 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
74 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517374 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 84.03 | 184.39 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.27822 | 20.7 | 1.095 | 1.086 | 188.44 | 198.37 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.78929 | 20.7 | 1.096 | 1.087 | 198.37 | 203.37 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
75 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517375 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.84 | 183.84 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.63593 | 20.7 | 1.095 | 1.086 | 187.85 | 197.76 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.98911 | 20.7 | 1.096 | 1.087 | 197.76 | 202.73 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.67946 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
76 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517376 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.5 | 183.08 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -1.0533 | 20.7 | 1.095 | 1.086 | 187.09 | 197.01 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.8892 | 20.7 | 1.096 | 1.087 | 197.01 | 202 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
77 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517377 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.29 | 182.96 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.96383 | 20.7 | 1.095 | 1.086 | 187 | 196.94 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.68938 | 20.7 | 1.096 | 1.087 | 196.94 | 201.92 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.47962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
78 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517378 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.35 | 183.35 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.63593 | 20.7 | 1.095 | 1.086 | 187.38 | 197.34 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.48956 | 20.7 | 1.096 | 1.087 | 197.34 | 202.32 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.47962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
79 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517379 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.76 | 183.45 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.94396 | 20.7 | 1.095 | 1.086 | 187.46 | 197.37 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.98911 | 20.7 | 1.096 | 1.087 | 197.37 | 202.34 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.67946 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
80 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517380 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 84.31 | 184.1 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.84459 | 20.7 | 1.095 | 1.086 | 188.1 | 198.01 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.98911 | 20.7 | 1.096 | 1.087 | 198.01 | 202.98 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.67946 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
81 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517381 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.71 | 183.97 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.37758 | 20.7 | 1.095 | 1.086 | 188.01 | 197.94 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.78929 | 20.7 | 1.096 | 1.087 | 197.94 | 202.95 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | 0.1199 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
82 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517382 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.55 | 183.55 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.63593 | 20.7 | 1.095 | 1.086 | 187.55 | 197.49 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.68938 | 20.7 | 1.096 | 1.087 | 197.49 | 202.48 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
83 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517383 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 84.8 | 184.98 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.45707 | 20.7 | 1.095 | 1.086 | 189.02 | 198.94 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.8892 | 20.7 | 1.096 | 1.087 | 198.94 | 203.92 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.47962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
84 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517384 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 84.74 | 185.24 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.13911 | 20.7 | 1.095 | 1.086 | 189.26 | 199.23 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.38965 | 20.7 | 1.096 | 1.087 | 199.23 | 204.24 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | 0.1199 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
85 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517385 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.2 | 183.05 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.78498 | 20.7 | 1.095 | 1.086 | 187.07 | 196.97 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -1.089 | 20.7 | 1.096 | 1.087 | 196.97 | 201.91 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -1.279 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
86 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517386 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 85.19 | 185.74 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.089428 | 20.7 | 1.095 | 1.086 | 189.79 | 199.75 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.48956 | 20.7 | 1.096 | 1.087 | 199.75 | 204.74 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
87 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517387 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 85.54 | 186 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.17886 | 20.7 | 1.095 | 1.086 | 190.05 | 200.01 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.48956 | 20.7 | 1.096 | 1.087 | 200.01 | 205 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.27978 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
88 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517388 | 25.6 | 0 | 20.4 | 8.512 | 0.841 | 83.63 | 183.64 | 100.64 | 2,32,371 | 2 | 32,371 | 0.042325 | 2377.7753 | -0.62599 | 20.7 | 1.095 | 1.086 | 187.67 | 197.64 | 10.009 | 22,54,120 | 22 | 54,120 | 0.3817 | 26.2222 | -0.38965 | 20.7 | 1.096 | 1.087 | 197.64 | 202.64 | 5.004 | 18,15,991 | 18 | 15,991 | 0.30444 | 16.4366 | -0.079936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
89 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517389 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 98.41 | 198.2 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.82489 | 20.4 | 1.098 | 1.089 | 202.65 | 212.55 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -1.0791 | 20.5 | 1.099 | 1.09 | 212.55 | 217.52 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
90 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517390 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 98.05 | 198.1 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.56649 | 20.4 | 1.098 | 1.089 | 202.56 | 212.55 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.17986 | 20.5 | 1.099 | 1.09 | 212.55 | 217.58 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | 0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
91 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517391 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 97.55 | 197.51 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.65593 | 20.4 | 1.098 | 1.089 | 202 | 211.95 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.57954 | 20.5 | 1.099 | 1.09 | 211.95 | 216.98 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | 0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
92 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517392 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 98.44 | 198.57 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.48698 | 20.4 | 1.098 | 1.089 | 203.06 | 213 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.67946 | 20.5 | 1.099 | 1.09 | 213 | 217.96 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.8991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
93 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517393 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 97.49 | 197 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -1.1032 | 20.4 | 1.098 | 1.089 | 201.46 | 211.38 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.8793 | 20.5 | 1.099 | 1.09 | 211.38 | 216.36 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
94 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517394 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 99.39 | 199.08 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.92427 | 20.4 | 1.098 | 1.089 | 203.55 | 213.48 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.77938 | 20.5 | 1.099 | 1.09 | 213.48 | 218.48 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
95 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517395 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 99.7 | 200.44 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | 0.11926 | 20.4 | 1.098 | 1.089 | 204.88 | 214.84 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.47962 | 20.5 | 1.099 | 1.09 | 214.84 | 219.88 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | 0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
96 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517396 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 97.85 | 198.03 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.43729 | 20.4 | 1.098 | 1.089 | 202.5 | 212.45 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.57954 | 20.5 | 1.099 | 1.09 | 212.45 | 217.44 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
97 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517397 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 100.31 | 199.94 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.9839 | 20.4 | 1.098 | 1.089 | 204.42 | 214.35 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.77938 | 20.5 | 1.099 | 1.09 | 214.35 | 219.34 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
98 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517398 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 100.5 | 200.89 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.22858 | 20.4 | 1.098 | 1.089 | 205.37 | 215.32 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.57954 | 20.5 | 1.099 | 1.09 | 215.32 | 220.33 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | 0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
99 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517399 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 99.64 | 200.14 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.11926 | 20.4 | 1.098 | 1.089 | 204.6 | 214.55 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.57954 | 20.5 | 1.099 | 1.09 | 214.55 | 219.55 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
100 | 2025-09-22 | 2025-09-22 | 3 | 8 |  | PA25517400 | 25.6 | 0 | 20.4 | 8.515 | 0.846 | 100.38 | 200.59 | 100.62 | 2,31,891 | 2 | 31,891 | 0.042192 | 2384.8154 | -0.40747 | 20.4 | 1.098 | 1.089 | 205.1 | 215.02 | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -0.8793 | 20.5 | 1.099 | 1.09 | 215.02 | 220.02 | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
101 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
102 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
103 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
104 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
105 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
106 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
107 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
108 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
109 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
110 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
111 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
112 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
113 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
114 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
115 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
116 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
117 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
118 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
119 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
120 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
121 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
122 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
123 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
124 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
125 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
126 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
127 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
128 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
129 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
130 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
131 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
132 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
133 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
134 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
135 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
136 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
137 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
138 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
139 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
140 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
141 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
142 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
143 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
144 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
145 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
146 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
147 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
148 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
149 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
150 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
151 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
152 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
153 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
154 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
155 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
156 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
157 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
158 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
159 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
160 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
161 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
162 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
163 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
164 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
165 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
166 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
167 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
168 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
169 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
170 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
171 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
172 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
173 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
174 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
175 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
176 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
177 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
178 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
179 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
180 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
181 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
182 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
183 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
184 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
185 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
186 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
187 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
188 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
189 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
190 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
191 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
192 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
193 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
194 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
195 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
196 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
197 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
198 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
199 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
200 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
201 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
202 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
203 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
204 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
205 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
206 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
207 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
208 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
209 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
210 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
211 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
212 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
213 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
214 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
215 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
216 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
217 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
218 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
219 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
220 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
221 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
222 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
223 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
224 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
225 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
226 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
227 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
228 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
229 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
230 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
231 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
232 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
233 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
234 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
235 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
236 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
237 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
238 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
239 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
240 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
241 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
242 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
243 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
244 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
245 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
246 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
247 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
248 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
249 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
250 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
251 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
252 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
253 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
254 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
255 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
256 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
257 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
258 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
259 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
260 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
261 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
262 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
263 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
264 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
265 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
266 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
267 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
268 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
269 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
270 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
271 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
272 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
273 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
274 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
275 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
276 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
277 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
278 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
279 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
280 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
281 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
282 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
283 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
284 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
285 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
286 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
287 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
288 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
289 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
290 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
291 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
292 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
293 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
294 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
295 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
296 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
297 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
298 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
299 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
300 | [17 empty] | #VALUE! | #VALUE! | #DIV/0! | 20.4 | 1.098 | 1.089 |  |  | 10.008 | 22,51,021 | 22 | 51,021 | 0.38084 | 26.2788 | -100 | 20.5 | 1.099 | 1.09 | 0 |  | 5.005 | 18,14,801 | 18 | 14,801 | 0.30411 | 16.4578 | -100 | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | aceptable | aceptable | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | -1 | -1 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 3 | PA25517301 | 0 | -0.36746 | -0.089919 | -0.11986 | CONFORME | Medición
2 | 3 | PA25517302 | 0 | -0.79452 | 1.4087 | 2.2773 | CONFORME | Medición
3 | 3 | PA25517303 | 0 | 0.079452 | 1.0091 | 1.678 | CONFORME | Medición
4 | 3 | PA25517304 | 0 | -0.8839 | -1.3888 | -0.31962 | CONFORME | Medición
5 | 3 | PA25517305 | 0 | -0.57603 | 0.40963 | 0.27966 | CONFORME | Medición
6 | 3 | PA25517306 | 0 | -0.68527 | -0.089919 | 0.67918 | CONFORME | Medición
7 | 3 | PA25517307 | 0 | -0.10925 | -0.18983 | 0.87895 | CONFORME | Medición
8 | 3 | PA25517308 | 0 | -0.59589 | -0.089919 | 0.67918 | CONFORME | Medición
9 | 3 | PA25517309 | 0 | -0.31793 | -0.17986 | -0.079936 | CONFORME | Medición
10 | 3 | PA25517310 | 0 | 0.14903 | 0.1199 | 0.71942 | CONFORME | Medición
11 | 3 | PA25517311 | 0 | -0.84451 | -0.3797 | -0.079936 | CONFORME | Medición
12 | 3 | PA25517312 | 0 | -0.55638 | -0.27978 | -0.079936 | CONFORME | Medición
13 | 3 | PA25517313 | 0 | -0.308 | -0.47962 | -0.27978 | CONFORME | Medición
14 | 3 | PA25517314 | 0 | -1.1426 | -1.4788 | -1.279 | CONFORME | Medición
15 | 3 | PA25517315 | 0 | -0.84451 | -0.8793 | -0.27978 | CONFORME | Medición
16 | 3 | PA25517316 | 0 | -0.32787 | -0.27978 | -0.079936 | CONFORME | Medición
17 | 3 | PA25517317 | 0 | -0.75509 | -0.17986 | 0.91926 | CONFORME | Medición
18 | 3 | PA25517318 | 0 | -0.25832 | -0.47962 | -0.27978 | CONFORME | Medición
19 | 3 | PA25517319 | 0 | -1.0333 | -0.17986 | 0.1199 | CONFORME | Medición
20 | 3 | PA25517320 | 0 | -0.59613 | -0.67946 | -0.079936 | CONFORME | Medición
21 | 3 | PA25517321 | 0 | -0.40735 | -0.079936 | -0.079936 | CONFORME | Medición
22 | 3 | PA25517322 | 0 | 0.19871 | 0.019984 | 0.51958 | CONFORME | Medición
23 | 3 | PA25517323 | 0 | -0.58619 | -0.3797 | -0.079936 | CONFORME | Medición
24 | 3 | PA25517324 | 0 | 0.40735 | 0.21982 | 0.51958 | CONFORME | Medición
25 | 3 | PA25517325 | 0 | -0.97406 | -1.2383 | -0.33952 | CONFORME | Medición
26 | 3 | PA25517326 | 0 | -0.656 | -0.33952 | 0.65908 | CONFORME | Medición
27 | 3 | PA25517327 | 0 | -0.5566 | -0.43938 | 0.25964 | CONFORME | Medición
28 | 3 | PA25517328 | 0 | -0.10933 | -0.83883 | -0.1398 | CONFORME | Medición
29 | 3 | PA25517329 | 0 | -0.92436 | -0.1398 | 1.0585 | CONFORME | Medición
30 | 3 | PA25517330 | 0 | -0.089454 | -0.43938 | 0.65908 | CONFORME | Medición
31 | 3 | PA25517331 | 0 | 0.079515 | -0.53925 | -0.1398 | CONFORME | Medición
32 | 3 | PA25517332 | 0 | -0.44727 | -0.33952 | -0.1398 | CONFORME | Medición
33 | 3 | PA25517333 | 0 | -0.17891 | -0.43938 | 1.2582 | CONFORME | Medición
34 | 3 | PA25517334 | 0 | -0.85479 | -0.93869 | -0.73897 | CONFORME | Medición
35 | 3 | PA25517335 | 0 | -0.069576 | -0.43938 | 0.059916 | CONFORME | Medición
36 | 3 | PA25517336 | 0 | -1.0536 | -0.83883 | -0.73897 | CONFORME | Medición
37 | 3 | PA25517337 | 0 | -0.11927 | -0.43938 | 0.45936 | CONFORME | Medición
38 | 3 | PA25517338 | 0 | -0.079515 | -0.33952 | 0.059916 | CONFORME | Medición
39 | 3 | PA25517339 | 0 | -0.75539 | -0.63911 | -0.73897 | CONFORME | Medición
40 | 3 | PA25517340 | 0 | -0.48703 | -0.53925 | -0.33952 | CONFORME | Medición
41 | 3 | PA25517341 | 0 | -1.281 | -1.1889 | -1.4985 | CONFORME | Medición
42 | 3 | PA25517342 | 0 | -0.50645 | -1.2888 | -0.8991 | CONFORME | Medición
43 | 3 | PA25517343 | 0 | -0.52632 | -0.58947 | -0.0999 | CONFORME | Medición
44 | 3 | PA25517344 | 0 | -0.049652 | 0.009991 | -0.0999 | CONFORME | Medición
45 | 3 | PA25517345 | 0 | -0.17875 | -0.089919 | 0.8991 | CONFORME | Medición
46 | 3 | PA25517346 | 0 | -0.84409 | -0.98911 | -0.0999 | CONFORME | Medición
47 | 3 | PA25517347 | 0 | -0.6852 | -0.18983 | 0.0999 | CONFORME | Medición
48 | 3 | PA25517348 | 0 | -0.76465 | -0.18983 | 0.2997 | CONFORME | Medición
49 | 3 | PA25517349 | 0 | -0.20854 | -0.78929 | 0.0999 | CONFORME | Medición
50 | 3 | PA25517350 | 0 | -1.2413 | -1.4887 | -1.4985 | CONFORME | Medición
51 | 3 | PA25517351 | 0 | -0.10924 | -1.2888 | -3.2967 | CONFORME | Medición
52 | 3 | PA25517352 | 0 | -1.1122 | -0.28974 | -0.0999 | CONFORME | Medición
53 | 3 | PA25517353 | 0 | -0.96326 | -0.78929 | -0.2997 | CONFORME | Medición
54 | 3 | PA25517354 | 0 | -0.9434 | -0.28974 | 0.6993 | CONFORME | Medición
55 | 3 | PA25517355 | 0 | -1.6683 | -0.68938 | -0.0999 | CONFORME | Medición
56 | 3 | PA25517356 | 0 | -0.29791 | -0.48956 | -0.0999 | CONFORME | Medición
57 | 3 | PA25517357 | 0 | -0.90305 | -1.0989 | -0.51938 | CONFORME | Medición
58 | 3 | PA25517358 | 0 | -1.429 | -1.0989 | -1.9177 | CONFORME | Medición
59 | 3 | PA25517359 | 0 | -0.65496 | -1.3986 | -0.31962 | CONFORME | Medición
60 | 3 | PA25517360 | 0 | -0.82366 | 0.0999 | 1.4782 | CONFORME | Medición
61 | 3 | PA25517361 | 0 | -0.51603 | -0.7992 | 0.079904 | CONFORME | Medición
62 | 3 | PA25517362 | 0 | -0.81373 | -1.0989 | -0.51938 | CONFORME | Medición
63 | 3 | PA25517363 | 0 | -0.41679 | -0.1998 | 0.87895 | CONFORME | Medición
64 | 3 | PA25517364 | 0 | -0.35725 | -0.5994 | -0.11986 | CONFORME | Medición
65 | 3 | PA25517365 | 0 | -0.51603 | -0.6993 | 0.079904 | CONFORME | Medición
66 | 3 | PA25517366 | 0 | -0.31755 | -0.2997 | -0.11986 | CONFORME | Medición
67 | 3 | PA25517367 | 0 | -0.34733 | -0.7992 | -0.31962 | CONFORME | Medición
68 | 3 | PA25517368 | 0 | -1.2008 | -0.8991 | 0.47942 | CONFORME | Medición
69 | 3 | PA25517369 | 0 | -0.17862 | -0.5994 | -0.11986 | CONFORME | Medición
70 | 3 | PA25517370 | 0 | -1.3992 | -1.0989 | -0.9189 | CONFORME | Medición
71 | 3 | PA25517371 | 0 | -0.63511 | -0.5994 | -0.11986 | CONFORME | Medición
72 | 3 | PA25517372 | 0 | -0.42671 | -0.7992 | 0.27966 | CONFORME | Medición
73 | 3 | PA25517373 | 0 | -0.74523 | -0.58947 | -0.27978 | CONFORME | Medición
74 | 3 | PA25517374 | 0 | -0.27822 | -0.78929 | -0.079936 | CONFORME | Medición
75 | 3 | PA25517375 | 0 | -0.63593 | -0.98911 | -0.67946 | CONFORME | Medición
76 | 3 | PA25517376 | 0 | -1.0533 | -0.8892 | -0.27978 | CONFORME | Medición
77 | 3 | PA25517377 | 0 | -0.96383 | -0.68938 | -0.47962 | CONFORME | Medición
78 | 3 | PA25517378 | 0 | -0.63593 | -0.48956 | -0.47962 | CONFORME | Medición
79 | 3 | PA25517379 | 0 | -0.94396 | -0.98911 | -0.67946 | CONFORME | Medición
80 | 3 | PA25517380 | 0 | -0.84459 | -0.98911 | -0.67946 | CONFORME | Medición
81 | 3 | PA25517381 | 0 | -0.37758 | -0.78929 | 0.1199 | CONFORME | Medición
82 | 3 | PA25517382 | 0 | -0.63593 | -0.68938 | -0.27978 | CONFORME | Medición
83 | 3 | PA25517383 | 0 | -0.45707 | -0.8892 | -0.47962 | CONFORME | Medición
84 | 3 | PA25517384 | 0 | -0.13911 | -0.38965 | 0.1199 | CONFORME | Medición
85 | 3 | PA25517385 | 0 | -0.78498 | -1.089 | -1.279 | CONFORME | Medición
86 | 3 | PA25517386 | 0 | -0.089428 | -0.48956 | -0.27978 | CONFORME | Medición
87 | 3 | PA25517387 | 0 | -0.17886 | -0.48956 | -0.27978 | CONFORME | Medición
88 | 3 | PA25517388 | 0 | -0.62599 | -0.38965 | -0.079936 | CONFORME | Medición
89 | 3 | PA25517389 | 0 | -0.82489 | -1.0791 | -0.6993 | CONFORME | Medición
90 | 3 | PA25517390 | 0 | -0.56649 | -0.17986 | 0.4995 | CONFORME | Medición
91 | 3 | PA25517391 | 0 | -0.65593 | -0.57954 | 0.4995 | CONFORME | Medición
92 | 3 | PA25517392 | 0 | -0.48698 | -0.67946 | -0.8991 | CONFORME | Medición
93 | 3 | PA25517393 | 0 | -1.1032 | -0.8793 | -0.4995 | CONFORME | Medición
94 | 3 | PA25517394 | 0 | -0.92427 | -0.77938 | -0.0999 | CONFORME | Medición
95 | 3 | PA25517395 | 0 | 0.11926 | -0.47962 | 0.6993 | CONFORME | Medición
96 | 3 | PA25517396 | 0 | -0.43729 | -0.57954 | -0.2997 | CONFORME | Medición
97 | 3 | PA25517397 | 0 | -0.9839 | -0.77938 | -0.2997 | CONFORME | Medición
98 | 3 | PA25517398 | 0 | -0.22858 | -0.57954 | 0.0999 | CONFORME | Medición
99 | 3 | PA25517399 | 0 | -0.11926 | -0.57954 | -0.0999 | CONFORME | Medición
100 | 3 | PA25517400 | 0 | -0.40747 | -0.8793 | -0.0999 | CONFORME | Medición
101 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
102 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
103 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
104 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
105 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
106 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
107 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
108 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
109 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
110 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
111 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
112 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
113 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
114 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
115 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
116 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
117 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
118 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
119 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
120 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
121 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
122 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
123 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
124 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
125 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
126 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
127 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
128 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
129 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
130 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
131 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
132 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
133 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
134 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
135 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
136 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
137 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
138 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
139 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
140 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
141 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
142 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
143 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
144 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
145 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
146 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
147 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
148 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
149 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
150 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
151 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
152 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
153 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
154 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
155 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
156 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
157 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
158 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
159 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
160 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
161 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
162 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
163 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
164 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
165 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
166 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
167 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
168 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
169 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
170 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
171 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
172 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
173 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
174 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
175 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
176 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
177 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
178 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
179 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
180 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
181 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
182 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
183 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
184 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
185 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
186 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
187 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
188 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
189 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
190 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
191 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
192 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
193 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
194 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
195 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
196 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
197 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
198 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
199 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
200 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
201 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
202 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
203 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
204 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
205 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
206 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
207 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
208 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
209 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
210 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
211 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
212 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
213 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
214 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
215 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
216 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
217 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
218 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
219 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
220 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
221 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
222 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
223 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
224 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
225 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
226 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
227 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
228 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
229 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
230 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
231 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
232 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
233 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
234 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
235 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
236 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
237 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
238 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
239 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
240 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
241 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
242 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
243 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
244 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
245 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
246 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
247 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
248 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
249 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
250 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
251 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
252 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
253 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
254 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
255 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
256 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
257 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
258 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
259 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
260 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
261 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
262 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
263 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
264 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
265 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
266 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
267 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
268 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
269 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
270 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
271 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
272 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
273 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
274 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
275 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
276 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
277 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
278 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
279 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
280 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
281 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
282 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
283 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
284 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
285 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
286 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
287 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
288 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
289 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
290 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
291 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
292 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
293 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
294 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
295 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
296 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
297 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
298 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
299 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
300 | 0 | 0 | 0 | #DIV/0! | -100 | -100 | #DIV/0! | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición