## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517421 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 110.79 | 210.78 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.68534 | 20.7 | 1.393 | 1.392 | 213.39 | 223.4 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 223.4 | 228.48 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517422 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 111.14 | 211.49 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.32777 | 20.7 | 1.393 | 1.392 | 214.05 | 224.06 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 224.06 | 229.15 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517423 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 111.71 | 211.75 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.63568 | 20.7 | 1.393 | 1.392 | 214.36 | 224.39 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | 0.0998 | 20.7 | 1.399 | 1.392 | 224.39 | 229.48 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517424 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 111.62 | 212.24 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.059595 | 20.7 | 1.393 | 1.392 | 214.81 | 224.82 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 224.82 | 229.89 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517425 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 110.74 | 210.44 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.97338 | 20.7 | 1.393 | 1.392 | 212.99 | 222.97 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.3992 | 20.7 | 1.399 | 1.392 | 222.97 | 228.04 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517426 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 110.92 | 210.94 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.65554 | 20.7 | 1.393 | 1.392 | 213.54 | 223.55 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 223.55 | 228.63 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517427 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 112.09 | 212.74 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.029797 | 20.7 | 1.393 | 1.392 | 215.35 | 225.36 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 225.36 | 230.44 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517428 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 110.69 | 210.71 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.65554 | 20.7 | 1.393 | 1.392 | 213.31 | 223.32 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 223.32 | 228.38 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517429 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 111.86 | 212.05 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.48669 | 20.7 | 1.393 | 1.392 | 214.64 | 224.64 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.1996 | 20.7 | 1.399 | 1.392 | 224.64 | 229.7 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517430 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 112.1 | 211.65 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -1.1224 | 20.7 | 1.393 | 1.392 | 214.21 | 224.14 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.8982 | 20.7 | 1.399 | 1.392 | 224.14 | 229.08 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | -1.4759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517431 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 112.96 | 213.24 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.3973 | 20.7 | 1.393 | 1.392 | 215.79 | 225.8 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 225.8 | 230.85 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517432 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 112.36 | 212.64 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.3973 | 20.7 | 1.393 | 1.392 | 215.25 | 225.28 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | 0.0998 | 20.7 | 1.399 | 1.392 | 225.28 | 230.32 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517433 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 113.25 | 213.7 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.22845 | 20.7 | 1.393 | 1.392 | 216.3 | 226.34 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | 0.1996 | 20.7 | 1.399 | 1.392 | 226.34 | 231.41 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517434 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 112.31 | 212.19 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.7946 | 20.7 | 1.393 | 1.392 | 214.75 | 224.8 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | 0.2994 | 20.7 | 1.399 | 1.392 | 224.8 | 229.9 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517435 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 113.71 | 214.26 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.12912 | 20.7 | 1.393 | 1.392 | 216.84 | 226.85 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | -0.0998 | 20.7 | 1.399 | 1.392 | 226.85 | 231.9 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517436 | 25.6 | 0 | 20.7 | 8.325 | 0.525 | 112.91 | 212.9 | 100.68 | 2,31,120 | 2 | 31,120 | 0.041978 | 2398.4119 | -0.68534 | 20.7 | 1.393 | 1.392 | 215.5 | 225.58 | 10.02 | 22,39,500 | 22 | 39,500 | 0.37764 | 26.5333 | 0.5988 | 20.7 | 1.399 | 1.392 | 225.58 | 230.68 | 5.014 | 18,13,850 | 18 | 13,850 | 0.30385 | 16.5017 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517437 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.42 | 211.79 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.31781 | 20.6 | 1.392 | 1.386 | 214.31 | 224.35 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | 0.1896 | 20.8 | 1.394 | 1.398 | 224.35 | 229.45 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517438 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.89 | 212.35 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.22842 | 20.6 | 1.392 | 1.386 | 214.81 | 224.81 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.20956 | 20.8 | 1.394 | 1.398 | 224.81 | 229.85 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517439 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.78 | 211.86 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.60582 | 20.6 | 1.392 | 1.386 | 214.39 | 224.38 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.30935 | 20.8 | 1.394 | 1.398 | 224.38 | 229.44 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517440 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.81 | 211.98 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.51644 | 20.6 | 1.392 | 1.386 | 214.48 | 224.48 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.20956 | 20.8 | 1.394 | 1.398 | 224.48 | 229.54 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517441 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110.85 | 211.04 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.49657 | 20.6 | 1.392 | 1.386 | 213.55 | 223.57 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.009979 | 20.8 | 1.394 | 1.398 | 223.57 | 228.64 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517442 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110 | 209.82 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.86404 | 20.6 | 1.392 | 1.386 | 212.31 | 222.29 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.40914 | 20.8 | 1.394 | 1.398 | 222.29 | 227.3 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517443 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110.11 | 210.01 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.78459 | 20.6 | 1.392 | 1.386 | 212.5 | 222.5 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.20956 | 20.8 | 1.394 | 1.398 | 222.5 | 227.55 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517444 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.18 | 211.51 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.35753 | 20.6 | 1.392 | 1.386 | 214 | 224.07 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | 0.48897 | 20.8 | 1.394 | 1.398 | 224.07 | 229.18 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 1.9146 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517445 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 109.64 | 209.6 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.725 | 20.6 | 1.392 | 1.386 | 212.09 | 222.1 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.10977 | 20.8 | 1.394 | 1.398 | 222.1 | 227.16 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517446 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 108.4 | 207.62 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -1.4599 | 20.6 | 1.392 | 1.386 | 210.1 | 220 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -1.2075 | 20.8 | 1.394 | 1.398 | 220 | 224.98 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | -0.6781 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
27 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517447 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110.61 | 211 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.29794 | 20.6 | 1.392 | 1.386 | 213.5 | 223.52 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.009979 | 20.8 | 1.394 | 1.398 | 223.52 | 228.58 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517448 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110.87 | 211.25 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.30788 | 20.6 | 1.392 | 1.386 | 213.74 | 223.75 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.10977 | 20.8 | 1.394 | 1.398 | 223.75 | 228.81 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517449 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.23 | 211.78 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.13904 | 20.6 | 1.392 | 1.386 | 214.3 | 224.35 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | 0.28939 | 20.8 | 1.394 | 1.398 | 224.35 | 229.44 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517450 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110.9 | 211.39 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.19863 | 20.6 | 1.392 | 1.386 | 213.86 | 223.86 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | -0.20956 | 20.8 | 1.394 | 1.398 | 223.86 | 228.93 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517451 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 110.39 | 210.37 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.70513 | 20.6 | 1.392 | 1.386 | 212.82 | 222.85 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | 0.089811 | 20.8 | 1.394 | 1.398 | 222.85 | 227.91 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517452 | 25.6 | 0 | 20.7 | 8.329 | 0.517 | 111.17 | 211.55 | 100.69 | 2,30,690 | 2 | 30,690 | 0.041858 | 2405.4947 | -0.30788 | 20.6 | 1.392 | 1.386 | 214.05 | 224.15 | 10.021 | 22,40,600 | 22 | 40,600 | 0.37794 | 26.5145 | 0.78834 | 20.8 | 1.394 | 1.398 | 224.15 | 229.22 | 5.014 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4827 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
33 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517453 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.51 | 203.34 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | 0.23859 | 20.6 | 1.391 | 1.388 | 205.91 | 216.02 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | 0.8982 | 20.8 | 1.399 | 1.394 | 216.02 | 221.15 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 2.3135 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
34 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517454 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 103.55 | 203.54 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.59648 | 20.6 | 1.391 | 1.388 | 206.11 | 216.15 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | 0.1996 | 20.8 | 1.399 | 1.394 | 216.15 | 221.25 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
35 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517455 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 101.48 | 201.24 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.82513 | 20.6 | 1.391 | 1.388 | 203.8 | 213.81 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.0998 | 20.8 | 1.399 | 1.394 | 213.81 | 218.89 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
36 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517456 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 101.26 | 199.96 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -1.8789 | 20.6 | 1.391 | 1.388 | 202.05 | 212.04 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.2994 | 20.8 | 1.399 | 1.394 | 212.04 | 217.09 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517457 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 103.41 | 203.44 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.55672 | 20.6 | 1.391 | 1.388 | 206 | 216.03 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | 0.0998 | 20.8 | 1.399 | 1.394 | 216.03 | 221.09 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517458 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 103 | 203.16 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.42748 | 20.6 | 1.391 | 1.388 | 205.72 | 215.72 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.1996 | 20.8 | 1.399 | 1.394 | 215.72 | 220.74 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
39 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517459 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.6 | 202.78 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.4076 | 20.6 | 1.391 | 1.388 | 205.4 | 215.4 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.1996 | 20.8 | 1.399 | 1.394 | 215.4 | 220.44 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
40 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517460 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.28 | 201.65 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -1.2128 | 20.6 | 1.391 | 1.388 | 204.25 | 214.2 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.6986 | 20.8 | 1.399 | 1.394 | 214.2 | 219.24 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
41 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517461 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 103.03 | 203.09 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.52689 | 20.6 | 1.391 | 1.388 | 205.65 | 215.67 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -1.7728e-13 | 20.8 | 1.399 | 1.394 | 215.67 | 220.72 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
42 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517462 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.79 | 202.95 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.42748 | 20.6 | 1.391 | 1.388 | 205.55 | 215.58 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | 0.0998 | 20.8 | 1.399 | 1.394 | 215.58 | 220.65 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
43 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517463 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.61 | 202.5 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.69589 | 20.6 | 1.391 | 1.388 | 205.09 | 215.08 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.2994 | 20.8 | 1.399 | 1.394 | 215.08 | 220.11 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
44 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517464 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.69 | 202.95 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.32806 | 20.6 | 1.391 | 1.388 | 205.55 | 215.54 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.2994 | 20.8 | 1.399 | 1.394 | 215.54 | 220.55 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
45 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517465 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102.82 | 203.2 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.20877 | 20.6 | 1.391 | 1.388 | 205.77 | 215.82 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | 0.2994 | 20.8 | 1.399 | 1.394 | 215.82 | 220.89 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
46 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517466 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 104.64 | 205 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -0.22865 | 20.6 | 1.391 | 1.388 | 207.56 | 217.55 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.2994 | 20.8 | 1.399 | 1.394 | 217.55 | 222.57 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
47 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517467 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 102 | 201.49 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | -1.0935 | 20.6 | 1.391 | 1.388 | 203.55 | 213.51 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | -0.5988 | 20.8 | 1.399 | 1.394 | 213.51 | 218.52 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | -0.079777 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
48 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517468 | 25.6 | 0 | 20.6 | 8.33 | 0.535 | 103.75 | 204.35 | 100.59 | 2,30,740 | 2 | 30,740 | 0.041872 | 2402.3086 | 0.0099413 | 20.6 | 1.391 | 1.388 | 206.92 | 217 | 10.02 | 22,41,280 | 22 | 41,280 | 0.37813 | 26.4986 | 0.5988 | 20.8 | 1.399 | 1.394 | 217 | 222.06 | 5.014 | 18,13,400 | 18 | 13,400 | 0.30372 | 16.5085 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
49 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517469 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.68 | 202.24 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -1.1124 | 20.6 | 1.392 | 1.389 | 204.26 | 214.24 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.40914 | 20.9 | 1.397 | 1.396 | 214.24 | 219.28 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
50 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517470 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.92 | 203.05 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.54629 | 20.6 | 1.392 | 1.389 | 205.31 | 215.32 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.10977 | 20.9 | 1.397 | 1.396 | 215.32 | 220.39 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
51 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517471 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 101.4 | 201.38 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.69527 | 20.6 | 1.392 | 1.389 | 203.59 | 213.59 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.20956 | 20.9 | 1.397 | 1.396 | 213.59 | 218.64 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
52 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517472 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 100.79 | 200.4 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -1.0628 | 20.6 | 1.392 | 1.389 | 202.6 | 212.59 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.30935 | 20.9 | 1.397 | 1.396 | 212.59 | 217.64 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
53 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517473 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.55 | 202.61 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.61581 | 20.6 | 1.392 | 1.389 | 204.84 | 214.89 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | 0.28939 | 20.9 | 1.397 | 1.396 | 214.89 | 220 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 1.874 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
54 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517474 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 103.95 | 204.49 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.13905 | 20.6 | 1.392 | 1.389 | 206.71 | 216.72 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.10977 | 20.9 | 1.397 | 1.396 | 216.72 | 221.76 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517475 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 103.26 | 203.29 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.64561 | 20.6 | 1.392 | 1.389 | 205.5 | 215.54 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | 0.1896 | 20.9 | 1.397 | 1.396 | 215.54 | 220.6 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.87719 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
56 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517476 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.25 | 201.6 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -1.321 | 20.6 | 1.392 | 1.389 | 203.8 | 213.74 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.8083 | 20.9 | 1.397 | 1.396 | 213.74 | 218.74 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
57 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517477 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 103.1 | 203.34 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.43703 | 20.6 | 1.392 | 1.389 | 205.55 | 215.55 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.20956 | 20.9 | 1.397 | 1.396 | 215.55 | 220.59 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
58 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517478 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.15 | 202.25 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.57608 | 20.6 | 1.392 | 1.389 | 204.46 | 214.46 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.20956 | 20.9 | 1.397 | 1.396 | 214.46 | 219.48 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.079745 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
59 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517479 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.06 | 202.38 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.35757 | 20.6 | 1.392 | 1.389 | 204.59 | 214.53 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.8083 | 20.9 | 1.397 | 1.396 | 214.53 | 219.5 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | -0.91707 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
60 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517480 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 105.03 | 205.15 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.55622 | 20.6 | 1.392 | 1.389 | 207.38 | 217.41 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | 0.089811 | 20.9 | 1.397 | 1.396 | 217.41 | 222.45 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
61 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517481 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 105.4 | 205.79 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.28804 | 20.6 | 1.392 | 1.389 | 208.05 | 218.08 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | 0.089811 | 20.9 | 1.397 | 1.396 | 218.08 | 223.15 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
62 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517482 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 105.23 | 205.6 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.30791 | 20.6 | 1.392 | 1.389 | 207.84 | 217.82 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | -0.40914 | 20.9 | 1.397 | 1.396 | 217.82 | 222.82 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | -0.31898 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
63 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517483 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 105.48 | 205.91 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.24831 | 20.6 | 1.392 | 1.389 | 208.18 | 218.25 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | 0.48897 | 20.9 | 1.397 | 1.396 | 218.25 | 223.35 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 1.6746 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
64 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517484 | 25.6 | 0 | 20.6 | 8.346 | 0.541 | 102.46 | 202.3 | 100.68 | 2,30,840 | 2 | 30,840 | 0.0419 | 2402.864 | -0.83433 | 20.6 | 1.392 | 1.389 | 204.51 | 214.58 | 10.021 | 22,39,300 | 22 | 39,300 | 0.37758 | 26.5398 | 0.48897 | 20.9 | 1.397 | 1.396 | 214.58 | 219.65 | 5.016 | 18,14,110 | 18 | 14,110 | 0.30392 | 16.5044 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
65 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517485 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 93.94 | 194.02 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.5169 | 20.7 | 1.395 | 1.386 | 196.44 | 206.51 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.47895 | 21.1 | 1.395 | 1.396 | 206.51 | 211.62 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.8537 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
66 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517486 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 95.2 | 195.15 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.64612 | 20.7 | 1.395 | 1.386 | 197.56 | 207.56 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | -0.21952 | 21.1 | 1.395 | 1.396 | 207.56 | 212.64 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.2557 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
67 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517487 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 94.39 | 194.67 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.31809 | 20.7 | 1.395 | 1.386 | 197.1 | 207.09 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | -0.3193 | 21.1 | 1.395 | 1.396 | 207.09 | 212.12 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 0.25912 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
68 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517488 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 90.05 | 190 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.64612 | 20.7 | 1.395 | 1.386 | 192.44 | 202.44 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | -0.21952 | 21.1 | 1.395 | 1.396 | 202.44 | 207.45 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | -0.13953 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
69 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517489 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 91.11 | 191.49 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.21869 | 20.7 | 1.395 | 1.386 | 193.89 | 203.94 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.27939 | 21.1 | 1.395 | 1.396 | 203.94 | 209.02 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.2557 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
70 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517490 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 89.79 | 189.7 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.68588 | 20.7 | 1.395 | 1.386 | 192.14 | 202.15 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | -0.11974 | 21.1 | 1.395 | 1.396 | 202.15 | 207.21 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 0.85709 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
71 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517491 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 91.17 | 191.23 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.53678 | 20.7 | 1.395 | 1.386 | 193.64 | 203.69 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.27939 | 21.1 | 1.395 | 1.396 | 203.69 | 208.75 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 0.85709 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
72 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517492 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 91.43 | 192.15 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | 0.11928 | 20.7 | 1.395 | 1.386 | 194.59 | 204.6 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | -0.11974 | 21.1 | 1.395 | 1.396 | 204.6 | 209.64 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 0.45844 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
73 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517493 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 95.92 | 196.15 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.36779 | 20.7 | 1.395 | 1.386 | 198.55 | 208.59 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.1796 | 21.1 | 1.395 | 1.396 | 208.59 | 213.64 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 0.65776 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
74 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517494 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 96.34 | 196.58 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.35785 | 20.7 | 1.395 | 1.386 | 199 | 209.09 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.67851 | 21.1 | 1.395 | 1.396 | 209.09 | 214.21 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 2.053 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
75 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517495 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 96.69 | 196.81 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.47714 | 20.7 | 1.395 | 1.386 | 199.25 | 209.34 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.67851 | 21.1 | 1.395 | 1.396 | 209.34 | 214.45 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.8537 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
76 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517496 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 97.1 | 197.12 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.57654 | 20.7 | 1.395 | 1.386 | 199.55 | 209.59 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.1796 | 21.1 | 1.395 | 1.396 | 209.59 | 214.67 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.2557 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
77 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517497 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 133.7 | 234.04 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.25845 | 20.7 | 1.395 | 1.386 | 236.49 | 246.59 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.77829 | 21.1 | 1.395 | 1.396 | 246.59 | 251.7 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.8537 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
78 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517498 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 91.82 | 192.1 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.31809 | 20.7 | 1.395 | 1.386 | 194.54 | 204.58 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.1796 | 21.1 | 1.395 | 1.396 | 204.58 | 209.65 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.0564 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
79 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517499 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 91.21 | 190.78 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -1.0239 | 20.7 | 1.395 | 1.386 | 193.22 | 203.27 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.27939 | 21.1 | 1.395 | 1.396 | 203.27 | 208.38 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.8537 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
80 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517500 | 25.6 | 0 | 20.7 | 8.352 | 0.549 | 91.95 | 191.95 | 100.6 | 2,31,120 | 2 | 31,120 | 0.041978 | 2396.5061 | -0.59642 | 20.7 | 1.395 | 1.386 | 194.4 | 204.5 | 10.022 | 22,42,010 | 22 | 42,010 | 0.37834 | 26.4897 | 0.77829 | 21.1 | 1.395 | 1.396 | 204.5 | 209.61 | 5.017 | 18,15,110 | 18 | 15,110 | 0.3042 | 16.4926 | 1.8537 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
81 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517501 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 95.7 | 195.69 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.68534 | 20.8 | 1.392 | 1.387 | 198.01 | 208 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | -0.30935 | 21.1 | 1.397 | 1.395 | 208 | 213.11 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.8943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
82 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517502 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 134.94 | 235.01 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.60588 | 20.8 | 1.392 | 1.387 | 237.35 | 247.35 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | -0.20956 | 21.1 | 1.397 | 1.395 | 247.35 | 252.39 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
83 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517503 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 95.59 | 195.56 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.7052 | 20.8 | 1.392 | 1.387 | 197.88 | 207.92 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.1896 | 21.1 | 1.397 | 1.395 | 207.92 | 213.01 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
84 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517504 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 91.56 | 191.25 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.98331 | 20.8 | 1.392 | 1.387 | 193.55 | 203.6 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.28939 | 21.1 | 1.397 | 1.395 | 203.6 | 208.67 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
85 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517505 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 131.95 | 231.79 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.83433 | 20.8 | 1.392 | 1.387 | 234.14 | 244.19 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.28939 | 21.1 | 1.397 | 1.395 | 244.19 | 249.28 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
86 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517506 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 92.36 | 192.31 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.72507 | 20.8 | 1.392 | 1.387 | 194.62 | 204.65 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.089811 | 21.1 | 1.397 | 1.395 | 204.65 | 209.74 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
87 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517507 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 91.74 | 191.95 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.46683 | 20.8 | 1.392 | 1.387 | 194.31 | 204.34 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.089811 | 21.1 | 1.397 | 1.395 | 204.34 | 209.41 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
88 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517508 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 135.19 | 235.11 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.75487 | 20.8 | 1.392 | 1.387 | 237.44 | 247.5 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.38918 | 21.1 | 1.397 | 1.395 | 247.5 | 252.59 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
89 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517509 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 96.12 | 195.96 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.83433 | 20.8 | 1.392 | 1.387 | 198.31 | 208.38 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.48897 | 21.1 | 1.397 | 1.395 | 208.38 | 213.43 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
90 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517510 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 92.45 | 192.49 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.63568 | 20.8 | 1.392 | 1.387 | 194.8 | 204.86 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.38918 | 21.1 | 1.397 | 1.395 | 204.86 | 209.98 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 2.0937 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
91 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517511 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 92.34 | 191.55 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -1.4601 | 20.8 | 1.392 | 1.387 | 193.84 | 203.8 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | -0.60872 | 21.1 | 1.397 | 1.395 | 203.8 | 208.88 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
92 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517512 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 92.95 | 193.44 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.18872 | 20.8 | 1.392 | 1.387 | 195.74 | 205.79 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.28939 | 21.1 | 1.397 | 1.395 | 205.79 | 210.85 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
93 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517513 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 92.35 | 192.31 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.71514 | 20.8 | 1.392 | 1.387 | 194.62 | 204.65 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.089811 | 21.1 | 1.397 | 1.395 | 204.65 | 209.75 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
94 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517514 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 130.21 | 229.55 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -1.3309 | 20.8 | 1.392 | 1.387 | 231.85 | 241.89 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.1896 | 21.1 | 1.397 | 1.395 | 241.89 | 246.94 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
95 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517515 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 93.59 | 194.45 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | 0.17878 | 20.8 | 1.392 | 1.387 | 196.76 | 206.82 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.38918 | 21.1 | 1.397 | 1.395 | 206.82 | 211.89 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
96 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517516 | 25.6 | 0 | 20.8 | 8.339 | 0.548 | 92.54 | 192.31 | 100.68 | 2,30,960 | 2 | 30,960 | 0.041933 | 2400.9539 | -0.90385 | 20.8 | 1.392 | 1.387 | 194.62 | 204.71 | 10.021 | 22,40,090 | 22 | 40,090 | 0.3778 | 26.5244 | 0.68855 | 21.1 | 1.397 | 1.395 | 204.71 | 209.83 | 5.015 | 18,18,340 | 18 | 18,340 | 0.30509 | 16.4375 | 2.0937 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
97 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517517 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 181.3 | 281.2 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.79444 | 20.9 | 1.399 | 1.394 | 283.71 | 293.74 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.10979 | 21.5 | 1.402 | 1.396 | 293.74 | 298.77 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
98 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517518 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 102.34 | 202.3 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.73486 | 20.9 | 1.399 | 1.394 | 204.83 | 214.86 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.10979 | 21.5 | 1.402 | 1.396 | 214.86 | 219.95 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 1.5158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
99 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517519 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 101.35 | 201.31 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.73486 | 20.9 | 1.399 | 1.394 | 203.84 | 213.86 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.009981 | 21.5 | 1.402 | 1.396 | 213.86 | 218.91 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
100 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517520 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 103.34 | 204.62 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | 0.57597 | 20.9 | 1.399 | 1.394 | 207.21 | 217.3 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.70865 | 21.5 | 1.402 | 1.396 | 217.3 | 222.4 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
101 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517521 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 103.05 | 203.57 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.17875 | 20.9 | 1.399 | 1.394 | 206.14 | 216.28 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 1.2077 | 21.5 | 1.402 | 1.396 | 216.28 | 221.38 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
102 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517522 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 104.41 | 204.78 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.32771 | 20.9 | 1.399 | 1.394 | 207.35 | 217.39 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.2096 | 21.5 | 1.402 | 1.396 | 217.39 | 222.44 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
103 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517523 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 139.94 | 239.71 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.92354 | 20.9 | 1.399 | 1.394 | 242.29 | 252.29 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | -0.18964 | 21.5 | 1.402 | 1.396 | 252.29 | 257.31 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
104 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517524 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 103.06 | 204.1 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | 0.33764 | 20.9 | 1.399 | 1.394 | 206.65 | 216.71 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.40922 | 21.5 | 1.402 | 1.396 | 216.71 | 221.79 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
105 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517525 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 137.34 | 237.86 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.17875 | 20.9 | 1.399 | 1.394 | 240.44 | 250.54 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.80846 | 21.5 | 1.402 | 1.396 | 250.54 | 255.67 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 2.3135 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
106 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517526 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 103.15 | 203.76 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.089374 | 20.9 | 1.399 | 1.394 | 206.35 | 216.4 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.30941 | 21.5 | 1.402 | 1.396 | 216.4 | 221.46 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
107 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517527 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 101.68 | 202.55 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | 0.16882 | 20.9 | 1.399 | 1.394 | 205.11 | 215.13 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.009981 | 21.5 | 1.402 | 1.396 | 215.13 | 220.15 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
108 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517528 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 144.94 | 244.99 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.64548 | 20.9 | 1.399 | 1.394 | 247.53 | 257.56 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.10979 | 21.5 | 1.402 | 1.396 | 257.56 | 262.58 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.11966 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
109 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517529 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 103.75 | 203.89 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.55611 | 20.9 | 1.399 | 1.394 | 206.45 | 216.51 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.40922 | 21.5 | 1.402 | 1.396 | 216.51 | 221.56 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
110 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517530 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 145.84 | 246.15 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.38729 | 20.9 | 1.399 | 1.394 | 248.68 | 258.67 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | -0.28945 | 21.5 | 1.402 | 1.396 | 258.67 | 263.7 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
111 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517531 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 143.61 | 243.9 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.40715 | 20.9 | 1.399 | 1.394 | 246.45 | 256.5 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.30941 | 21.5 | 1.402 | 1.396 | 256.5 | 261.55 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
112 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517532 | 25.6 | 0 | 20.9 | 8.308 | 0.52 | 141.19 | 241.02 | 100.7 | 2,31,150 | 2 | 31,150 | 0.041986 | 2398.4122 | -0.86395 | 20.9 | 1.399 | 1.394 | 243.56 | 253.58 | 10.019 | 22,40,450 | 22 | 40,450 | 0.3779 | 26.5121 | 0.009981 | 21.5 | 1.402 | 1.396 | 253.58 | 258.64 | 5.014 | 18,17,910 | 18 | 17,910 | 0.30497 | 16.4407 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
113 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517533 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 140.09 | 239.84 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.92372 | 21.1 | 1.39 | 1.383 | 243.29 | 253.31 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -0.049875 | 21.3 | 1.403 | 1.396 | 253.31 | 258.34 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
114 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517534 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 99.56 | 198.77 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -1.4601 | 21.1 | 1.39 | 1.383 | 202.19 | 212.09 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -1.2469 | 21.3 | 1.403 | 1.396 | 212.09 | 216.99 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | -2.2931 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
115 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517535 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 99.76 | 199.84 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.59595 | 21.1 | 1.39 | 1.383 | 203.3 | 213.31 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -0.14963 | 21.3 | 1.403 | 1.396 | 213.31 | 218.35 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
116 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517536 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 97.01 | 197.09 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.59595 | 21.1 | 1.39 | 1.383 | 200.54 | 210.58 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.14963 | 21.3 | 1.403 | 1.396 | 210.58 | 215.6 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
117 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517537 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 101.14 | 201.23 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.58602 | 21.1 | 1.39 | 1.383 | 204.64 | 214.66 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -0.049875 | 21.3 | 1.403 | 1.396 | 214.66 | 219.65 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
118 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517538 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 97.39 | 198.24 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | 0.16885 | 21.1 | 1.39 | 1.383 | 201.68 | 211.74 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.34913 | 21.3 | 1.403 | 1.396 | 211.74 | 216.81 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
119 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517539 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 101.54 | 201.41 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.80453 | 21.1 | 1.39 | 1.383 | 204.82 | 214.88 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.34913 | 21.3 | 1.403 | 1.396 | 214.88 | 219.95 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
120 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517540 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 98.92 | 198.55 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -1.0429 | 21.1 | 1.39 | 1.383 | 201.97 | 211.94 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -0.54863 | 21.3 | 1.403 | 1.396 | 211.94 | 216.96 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
121 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517541 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 100.7 | 201.21 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.16885 | 21.1 | 1.39 | 1.383 | 204.65 | 214.71 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.34913 | 21.3 | 1.403 | 1.396 | 214.71 | 219.8 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
122 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517542 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 98.49 | 198.51 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.65554 | 21.1 | 1.39 | 1.383 | 201.92 | 211.93 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -0.14963 | 21.3 | 1.403 | 1.396 | 211.93 | 216.94 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
123 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517543 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 100.75 | 200.8 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.62574 | 21.1 | 1.39 | 1.383 | 204.27 | 214.31 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.14963 | 21.3 | 1.403 | 1.396 | 214.31 | 219.39 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
124 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517544 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 102.56 | 202.49 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.74493 | 21.1 | 1.39 | 1.383 | 205.89 | 215.91 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | -0.049875 | 21.3 | 1.403 | 1.396 | 215.91 | 220.95 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
125 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517545 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 101.91 | 201.49 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -1.0926 | 21.1 | 1.39 | 1.383 | 204.88 | 214.93 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.24938 | 21.3 | 1.403 | 1.396 | 214.93 | 219.99 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
126 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517546 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 100.7 | 200.71 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.66547 | 21.1 | 1.39 | 1.383 | 204.16 | 214.21 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.24938 | 21.3 | 1.403 | 1.396 | 214.21 | 219.29 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
127 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517547 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 101.45 | 201.59 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.53635 | 21.1 | 1.39 | 1.383 | 205.01 | 215.05 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.14963 | 21.3 | 1.403 | 1.396 | 215.05 | 220.09 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
128 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517548 | 25.6 | 0 | 21.1 | 8.323 | 0.532 | 103.62 | 204.25 | 100.68 | 2,30,850 | 2 | 30,850 | 0.041903 | 2402.7047 | -0.049662 | 21.1 | 1.39 | 1.383 | 207.7 | 217.82 | 10.025 | 22,36,150 | 22 | 36,150 | 0.37671 | 26.6121 | 0.94763 | 21.3 | 1.403 | 1.396 | 217.82 | 222.93 | 5.015 | 18,15,050 | 18 | 15,050 | 0.30418 | 16.4869 | 1.8943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
129 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517549 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 97.84 | 197.87 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.5666 | 21.3 | 1.392 | 1.388 | 200.12 | 210.19 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.48897 | 21.3 | 1.392 | 1.388 | 210.19 | 215.25 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
130 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517550 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 96.26 | 195.7 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -1.1531 | 21.3 | 1.392 | 1.388 | 197.91 | 207.94 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.089811 | 21.3 | 1.392 | 1.388 | 207.94 | 212.95 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
131 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517551 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 97.16 | 198.24 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | 0.47714 | 21.3 | 1.392 | 1.388 | 200.49 | 210.55 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.38918 | 21.3 | 1.392 | 1.388 | 210.55 | 215.6 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
132 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517552 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 94.29 | 194.58 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.30815 | 21.3 | 1.392 | 1.388 | 196.8 | 206.76 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | -0.60872 | 21.3 | 1.392 | 1.388 | 206.76 | 211.76 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | -0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
133 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517553 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 96.07 | 196.53 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.13917 | 21.3 | 1.392 | 1.388 | 198.76 | 208.82 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.38918 | 21.3 | 1.392 | 1.388 | 208.82 | 213.87 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
134 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517554 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 97.4 | 198.14 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | 0.13917 | 21.3 | 1.392 | 1.388 | 200.39 | 210.39 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | -0.20956 | 21.3 | 1.392 | 1.388 | 210.39 | 215.4 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
135 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517555 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 135.24 | 235.85 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | 0.0099404 | 21.3 | 1.392 | 1.388 | 238.11 | 248.19 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.58876 | 21.3 | 1.392 | 1.388 | 248.19 | 253.24 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
136 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517556 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 97.65 | 197.7 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.54672 | 21.3 | 1.392 | 1.388 | 199.91 | 210 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.68855 | 21.3 | 1.392 | 1.388 | 210 | 215.05 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
137 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517557 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 95.36 | 195.01 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.94433 | 21.3 | 1.392 | 1.388 | 197.25 | 207.29 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.1896 | 21.3 | 1.392 | 1.388 | 207.29 | 212.31 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 2.1231e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
138 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517558 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 134.03 | 234.62 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.0099404 | 21.3 | 1.392 | 1.388 | 236.85 | 246.89 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.1896 | 21.3 | 1.392 | 1.388 | 246.89 | 251.93 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
139 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517559 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 96.45 | 196.19 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.85487 | 21.3 | 1.392 | 1.388 | 198.41 | 208.49 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.58876 | 21.3 | 1.392 | 1.388 | 208.49 | 213.53 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
140 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517560 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 96.55 | 196.76 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.38767 | 21.3 | 1.392 | 1.388 | 199 | 209.05 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.28939 | 21.3 | 1.392 | 1.388 | 209.05 | 214.1 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
141 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517561 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 98.85 | 198.76 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.68588 | 21.3 | 1.392 | 1.388 | 201.01 | 211.06 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.28939 | 21.3 | 1.392 | 1.388 | 211.06 | 216.11 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
142 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517562 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 94.76 | 194.51 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.84493 | 21.3 | 1.392 | 1.388 | 196.74 | 206.79 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.28939 | 21.3 | 1.392 | 1.388 | 206.79 | 211.83 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
143 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517563 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 98.26 | 198.54 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.31809 | 21.3 | 1.392 | 1.388 | 200.76 | 210.85 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.68855 | 21.3 | 1.392 | 1.388 | 210.85 | 215.92 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
144 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517564 | 25.6 | 0 | 21.4 | 8.338 | 0.548 | 96.21 | 196.21 | 100.6 | 2,31,100 | 2 | 31,100 | 0.041972 | 2396.8233 | -0.59642 | 21.3 | 1.392 | 1.388 | 198.45 | 208.54 | 10.021 | 22,39,720 | 22 | 39,720 | 0.3777 | 26.5316 | 0.68855 | 21.3 | 1.392 | 1.388 | 208.54 | 213.61 | 5.02 | 18,13,550 | 18 | 13,550 | 0.30376 | 16.526 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
145 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517565 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.05 | 195.46 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -1.2516 | 21.5 | 1.389 | 1.374 | 198 | 208 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | -0.1996 | 21.5 | 1.398 | 1.392 | 208 | 213.06 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
146 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517566 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.29 | 196.26 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.69534 | 21.5 | 1.389 | 1.374 | 198.8 | 208.89 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.6986 | 21.5 | 1.398 | 1.392 | 208.89 | 214.04 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 2.6919 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
147 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517567 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 95.33 | 195.19 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.80461 | 21.5 | 1.389 | 1.374 | 197.75 | 207.81 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.3992 | 21.5 | 1.398 | 1.392 | 207.81 | 212.9 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
148 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517568 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 95.19 | 195.1 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.75494 | 21.5 | 1.389 | 1.374 | 197.65 | 207.71 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.3992 | 21.5 | 1.398 | 1.392 | 207.71 | 212.8 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
149 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517569 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.01 | 195.72 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.95361 | 21.5 | 1.389 | 1.374 | 198.31 | 208.4 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.6986 | 21.5 | 1.398 | 1.392 | 208.4 | 213.5 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
150 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517570 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.24 | 196.06 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.84434 | 21.5 | 1.389 | 1.374 | 198.64 | 208.7 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.3992 | 21.5 | 1.398 | 1.392 | 208.7 | 213.8 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
151 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517571 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 93.84 | 193.45 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -1.0529 | 21.5 | 1.389 | 1.374 | 195.97 | 205.88 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | -1.0978 | 21.5 | 1.398 | 1.392 | 205.88 | 210.8 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | -1.8943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
152 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517572 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.39 | 196.89 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.16887 | 21.5 | 1.389 | 1.374 | 199.49 | 209.55 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.3992 | 21.5 | 1.398 | 1.392 | 209.55 | 214.63 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
153 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517573 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.74 | 196.78 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.62581 | 21.5 | 1.389 | 1.374 | 199.36 | 209.46 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.7984 | 21.5 | 1.398 | 1.392 | 209.46 | 214.55 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
154 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517574 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 95.01 | 195.06 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.61587 | 21.5 | 1.389 | 1.374 | 197.64 | 207.7 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.3992 | 21.5 | 1.398 | 1.392 | 207.7 | 212.79 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
155 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517575 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 133.98 | 233.66 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.98341 | 21.5 | 1.389 | 1.374 | 236.24 | 246.25 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | -0.0998 | 21.5 | 1.398 | 1.392 | 246.25 | 251.3 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
156 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517576 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 94.41 | 193.76 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -1.3112 | 21.5 | 1.389 | 1.374 | 196.36 | 206.42 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.3992 | 21.5 | 1.398 | 1.392 | 206.42 | 211.53 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.8943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
157 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517577 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 95.8 | 195.16 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -1.3013 | 21.5 | 1.389 | 1.374 | 197.69 | 207.69 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | -0.1996 | 21.5 | 1.398 | 1.392 | 207.69 | 212.71 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
158 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517578 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.77 | 196.74 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.69534 | 21.5 | 1.389 | 1.374 | 199.31 | 209.41 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.7984 | 21.5 | 1.398 | 1.392 | 209.41 | 214.51 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
159 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517579 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 94.25 | 194.31 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.60594 | 21.5 | 1.389 | 1.374 | 196.85 | 206.95 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.7984 | 21.5 | 1.398 | 1.392 | 206.95 | 212.05 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
160 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517580 | 25.6 | 0 | 21.5 | 8.325 | 0.537 | 96.1 | 196.2 | 100.67 | 2,30,860 | 2 | 30,860 | 0.041906 | 2402.3068 | -0.56621 | 21.5 | 1.389 | 1.374 | 198.75 | 208.87 | 10.02 | 22,40,360 | 22 | 40,360 | 0.37788 | 26.5165 | 0.998 | 21.5 | 1.398 | 1.392 | 208.87 | 214 | 5.015 | 18,14,130 | 18 | 14,130 | 0.30393 | 16.5008 | 2.2931 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
161 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517581 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 94.86 | 194.96 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.9009 | 21.5 | 1.392 | 1.385 | 197.8 | 207.84 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.1996 | 21.8 | 1.39 | 1.385 | 207.84 | 212.86 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
162 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517582 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 94.36 | 194.05 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -1.3068 | 21.5 | 1.392 | 1.385 | 196.9 | 206.98 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.5988 | 21.8 | 1.39 | 1.385 | 206.98 | 212.01 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
163 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517583 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 132.19 | 232.35 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.8415 | 21.5 | 1.392 | 1.385 | 235.18 | 245.26 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.5988 | 21.8 | 1.39 | 1.385 | 245.26 | 250.3 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
164 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517584 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 93.21 | 193.44 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.7722 | 21.5 | 1.392 | 1.385 | 196.29 | 206.31 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 1.0637e-13 | 21.8 | 1.39 | 1.385 | 206.31 | 211.32 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
165 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517585 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 130 | 230.15 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.8514 | 21.5 | 1.392 | 1.385 | 232.96 | 242.98 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | -1.7728e-13 | 21.8 | 1.39 | 1.385 | 242.98 | 247.98 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
166 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517586 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 95.05 | 195.09 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.9603 | 21.5 | 1.392 | 1.385 | 197.91 | 207.93 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 1.0637e-13 | 21.8 | 1.39 | 1.385 | 207.93 | 212.94 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
167 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517587 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 95.9 | 196.7 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.2079 | 21.5 | 1.392 | 1.385 | 199.59 | 209.69 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.7984 | 21.8 | 1.39 | 1.385 | 209.69 | 214.75 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
168 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517588 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 131.21 | 231.56 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.6534 | 21.5 | 1.392 | 1.385 | 234.41 | 244.41 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | -0.1996 | 21.8 | 1.39 | 1.385 | 244.41 | 249.43 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
169 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517589 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 94.65 | 195.51 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.1485 | 21.5 | 1.392 | 1.385 | 198.4 | 208.46 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.3992 | 21.8 | 1.39 | 1.385 | 208.46 | 213.52 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
170 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517590 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 95.89 | 195.92 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.9702 | 21.5 | 1.392 | 1.385 | 198.76 | 208.74 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | -0.3992 | 21.8 | 1.39 | 1.385 | 208.74 | 213.72 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
171 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517591 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 95.44 | 196.84 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | 0.3861 | 21.5 | 1.392 | 1.385 | 199.71 | 209.82 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.8982 | 21.8 | 1.39 | 1.385 | 209.82 | 214.88 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
172 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517592 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 128.66 | 229.21 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.4554 | 21.5 | 1.392 | 1.385 | 232.05 | 242.15 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.7984 | 21.8 | 1.39 | 1.385 | 242.15 | 247.21 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
173 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517593 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 93.85 | 194 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.8514 | 21.5 | 1.392 | 1.385 | 196.81 | 206.83 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 1.0637e-13 | 21.8 | 1.39 | 1.385 | 206.83 | 211.85 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
174 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517594 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 96 | 196.49 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.5148 | 21.5 | 1.392 | 1.385 | 199.34 | 209.34 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | -0.1996 | 21.8 | 1.39 | 1.385 | 209.34 | 214.34 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
175 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517595 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 94.95 | 195.17 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | -0.7821 | 21.5 | 1.392 | 1.385 | 198.01 | 208.05 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.1996 | 21.8 | 1.39 | 1.385 | 208.05 | 213.09 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
176 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517596 | 25.6 | 0 | 21.6 | 8.34 | 0.528 | 95.19 | 196.49 | 101.01 | 2,31,490 | 2 | 31,490 | 0.042081 | 2400.3961 | 0.2871 | 21.5 | 1.392 | 1.385 | 199.35 | 209.4 | 10.02 | 22,40,010 | 22 | 40,010 | 0.37778 | 26.5233 | 0.2994 | 21.8 | 1.39 | 1.385 | 209.4 | 214.51 | 5.01 | 18,17,090 | 18 | 17,090 | 0.30475 | 16.4399 | 1.996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
177 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517597 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 92.9 | 192.45 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -1.1322 | 21.6 | 1.395 | 1.386 | 194.65 | 204.59 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.7984 | 21.7 | 1.401 | 1.388 | 204.59 | 209.55 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -1.0178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
178 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517598 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 92.51 | 192.46 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.73493 | 21.6 | 1.395 | 1.386 | 194.69 | 204.65 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.5988 | 21.7 | 1.401 | 1.388 | 204.65 | 209.64 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
179 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517599 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 92.9 | 192.64 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.94349 | 21.6 | 1.395 | 1.386 | 194.86 | 204.85 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.2994 | 21.7 | 1.401 | 1.388 | 204.85 | 209.86 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
180 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517600 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 92.63 | 192.55 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.76472 | 21.6 | 1.395 | 1.386 | 194.79 | 204.8 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.0998 | 21.7 | 1.401 | 1.388 | 204.8 | 209.81 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
181 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517401 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94.05 | 194.09 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.64555 | 21.6 | 1.395 | 1.386 | 196.32 | 206.3 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.3992 | 21.7 | 1.401 | 1.388 | 206.3 | 211.3 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
182 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517402 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 93.21 | 193.1 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.79452 | 21.6 | 1.395 | 1.386 | 195.35 | 205.32 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.499 | 21.7 | 1.401 | 1.388 | 205.32 | 210.34 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | 0.1796 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
183 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517403 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94.54 | 194.95 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.27808 | 21.6 | 1.395 | 1.386 | 197.21 | 207.2 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.2994 | 21.7 | 1.401 | 1.388 | 207.2 | 212.2 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
184 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517404 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94.8 | 194.91 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.57603 | 21.6 | 1.395 | 1.386 | 197.15 | 207.13 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.3992 | 21.7 | 1.401 | 1.388 | 207.13 | 212.06 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -1.6164 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
185 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517405 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94.91 | 194.45 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -1.1421 | 21.6 | 1.395 | 1.386 | 196.66 | 206.6 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.7984 | 21.7 | 1.401 | 1.388 | 206.6 | 211.55 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -1.2173 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
186 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517406 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94 | 194.05 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.63561 | 21.6 | 1.395 | 1.386 | 196.29 | 206.3 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.0998 | 21.7 | 1.401 | 1.388 | 206.3 | 211.31 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
187 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517407 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 93.58 | 193.75 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.51644 | 21.6 | 1.395 | 1.386 | 196 | 206.01 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.0998 | 21.7 | 1.401 | 1.388 | 206.01 | 211.05 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | 0.57873 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
188 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517408 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94.5 | 195 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.1887 | 21.6 | 1.395 | 1.386 | 197.25 | 207.27 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | 1.0637e-13 | 21.7 | 1.401 | 1.388 | 207.27 | 212.29 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | 0.1796 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
189 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517409 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 93.48 | 193.25 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.9137 | 21.6 | 1.395 | 1.386 | 195.49 | 205.44 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.6986 | 21.7 | 1.401 | 1.388 | 205.44 | 210.41 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.8182 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
190 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517410 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 94.95 | 195.59 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -0.049657 | 21.6 | 1.395 | 1.386 | 197.83 | 207.86 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | 0.0998 | 21.7 | 1.401 | 1.388 | 207.86 | 212.89 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | 0.37917 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
191 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517411 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 93.85 | 194.77 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | 0.22842 | 21.6 | 1.395 | 1.386 | 197.03 | 207.12 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | 0.6986 | 21.7 | 1.401 | 1.388 | 207.12 | 212.18 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | 0.97785 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
192 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517412 | 25.6 | 0 | 21.6 | 8.341 | 0.522 | 93.4 | 192.94 | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -1.1421 | 21.6 | 1.395 | 1.386 | 195.17 | 205.14 | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -0.499 | 21.7 | 1.401 | 1.388 | 205.14 | 210.13 | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
193 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517413 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517414 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517415 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517416 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517417 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517418 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517419 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517420 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517421 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517422 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517423 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517424 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517425 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517426 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517427 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517428 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517429 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517430 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517431 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517432 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517433 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517434 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517435 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517436 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517437 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517438 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517439 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517440 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517441 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517442 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517443 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517444 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517445 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517446 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517447 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517448 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517449 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517450 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517451 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517452 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517453 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517454 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517455 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517456 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517457 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517458 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517459 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517460 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517461 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517462 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517463 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517464 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517465 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517466 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517467 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517468 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517469 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517470 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517471 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517472 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517473 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517474 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517475 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517476 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517477 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517478 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517479 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517480 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517481 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517482 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517483 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517484 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517485 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517486 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517487 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517488 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517489 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517490 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517491 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517492 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517493 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517494 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517495 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517496 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517497 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517498 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517499 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517500 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517501 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517502 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517503 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517504 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517505 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517506 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517507 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517508 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517509 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517510 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517511 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517512 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517513 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517514 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517715 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-20 | 2025-09-20 | 4 | 9 |  | PA25517716 | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-20 | 2025-09-20 | 4 | 9 |  |  | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-20 | 2025-09-20 | 4 | 9 |  |  | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-20 | 2025-09-20 | 4 | 9 |  |  | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-20 | 2025-09-20 | 4 | 9 |  |  | 25.6 | 0 | 21.6 | 8.341 | 0.522 |  |  | 100.69 | 2,30,820 | 2 | 30,820 | 0.041894 | 2403.4213 | -100 | 21.6 | 1.395 | 1.386 |  |  | 10.02 | 22,39,360 | 22 | 39,360 | 0.3776 | 26.536 | -100 | 21.7 | 1.401 | 1.388 | 0 |  | 5.011 | 18,12,010 | 18 | 12,010 | 0.30334 | 16.5196 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
108
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 4 | PA25517421 | 0 | -0.68534 | -0.0998 | 1.3163 | CONFORME | Medición
2 | 4 | PA25517422 | 0 | -0.32777 | -0.0998 | 1.5158 | CONFORME | Medición
3 | 4 | PA25517423 | 0 | -0.63568 | 0.0998 | 1.5158 | CONFORME | Medición
4 | 4 | PA25517424 | 0 | -0.059595 | -0.0998 | 1.1169 | CONFORME | Medición
5 | 4 | PA25517425 | 0 | -0.97338 | -0.3992 | 1.1169 | CONFORME | Medición
6 | 4 | PA25517426 | 0 | -0.65554 | -0.0998 | 1.3163 | CONFORME | Medición
7 | 4 | PA25517427 | 0 | -0.029797 | -0.0998 | 1.3163 | CONFORME | Medición
8 | 4 | PA25517428 | 0 | -0.65554 | -0.0998 | 0.91743 | CONFORME | Medición
9 | 4 | PA25517429 | 0 | -0.48669 | -0.1996 | 0.91743 | CONFORME | Medición
10 | 4 | PA25517430 | 0 | -1.1224 | -0.8982 | -1.4759 | CONFORME | Medición
11 | 4 | PA25517431 | 0 | -0.3973 | -0.0998 | 0.71799 | CONFORME | Medición
12 | 4 | PA25517432 | 0 | -0.3973 | 0.0998 | 0.51855 | CONFORME | Medición
13 | 4 | PA25517433 | 0 | -0.22845 | 0.1996 | 1.1169 | CONFORME | Medición
14 | 4 | PA25517434 | 0 | -0.7946 | 0.2994 | 1.7152 | CONFORME | Medición
15 | 4 | PA25517435 | 0 | -0.12912 | -0.0998 | 0.71799 | CONFORME | Medición
16 | 4 | PA25517436 | 0 | -0.68534 | 0.5988 | 1.7152 | CONFORME | Medición
17 | 4 | PA25517437 | 0 | -0.31781 | 0.1896 | 1.7152 | CONFORME | Medición
18 | 4 | PA25517438 | 0 | -0.22842 | -0.20956 | 0.51855 | CONFORME | Medición
19 | 4 | PA25517439 | 0 | -0.60582 | -0.30935 | 0.91743 | CONFORME | Medición
20 | 4 | PA25517440 | 0 | -0.51644 | -0.20956 | 0.91743 | CONFORME | Medición
21 | 4 | PA25517441 | 0 | -0.49657 | -0.009979 | 1.1169 | CONFORME | Medición
22 | 4 | PA25517442 | 0 | -0.86404 | -0.40914 | -0.079777 | CONFORME | Medición
23 | 4 | PA25517443 | 0 | -0.78459 | -0.20956 | 0.71799 | CONFORME | Medición
24 | 4 | PA25517444 | 0 | -0.35753 | 0.48897 | 1.9146 | CONFORME | Medición
25 | 4 | PA25517445 | 0 | -0.725 | -0.10977 | 0.91743 | CONFORME | Medición
26 | 4 | PA25517446 | 0 | -1.4599 | -1.2075 | -0.6781 | CONFORME | Medición
27 | 4 | PA25517447 | 0 | -0.29794 | -0.009979 | 0.91743 | CONFORME | Medición
28 | 4 | PA25517448 | 0 | -0.30788 | -0.10977 | 0.91743 | CONFORME | Medición
29 | 4 | PA25517449 | 0 | -0.13904 | 0.28939 | 1.5158 | CONFORME | Medición
30 | 4 | PA25517450 | 0 | -0.19863 | -0.20956 | 1.1169 | CONFORME | Medición
31 | 4 | PA25517451 | 0 | -0.70513 | 0.089811 | 0.91743 | CONFORME | Medición
32 | 4 | PA25517452 | 0 | -0.30788 | 0.78834 | 1.1169 | CONFORME | Medición
33 | 4 | PA25517453 | 0 | 0.23859 | 0.8982 | 2.3135 | CONFORME | Medición
34 | 4 | PA25517454 | 0 | -0.59648 | 0.1996 | 1.7152 | CONFORME | Medición
35 | 4 | PA25517455 | 0 | -0.82513 | -0.0998 | 1.3163 | CONFORME | Medición
36 | 4 | PA25517456 | 0 | -1.8789 | -0.2994 | 0.71799 | CONFORME | Medición
37 | 4 | PA25517457 | 0 | -0.55672 | 0.0998 | 0.91743 | CONFORME | Medición
38 | 4 | PA25517458 | 0 | -0.42748 | -0.1996 | 0.11966 | CONFORME | Medición
39 | 4 | PA25517459 | 0 | -0.4076 | -0.1996 | 0.51855 | CONFORME | Medición
40 | 4 | PA25517460 | 0 | -1.2128 | -0.6986 | 0.51855 | CONFORME | Medición
41 | 4 | PA25517461 | 0 | -0.52689 | -1.7728e-13 | 0.71799 | CONFORME | Medición
42 | 4 | PA25517462 | 0 | -0.42748 | 0.0998 | 1.1169 | CONFORME | Medición
43 | 4 | PA25517463 | 0 | -0.69589 | -0.2994 | 0.31911 | CONFORME | Medición
44 | 4 | PA25517464 | 0 | -0.32806 | -0.2994 | -0.079777 | CONFORME | Medición
45 | 4 | PA25517465 | 0 | -0.20877 | 0.2994 | 1.1169 | CONFORME | Medición
46 | 4 | PA25517466 | 0 | -0.22865 | -0.2994 | 0.11966 | CONFORME | Medición
47 | 4 | PA25517467 | 0 | -1.0935 | -0.5988 | -0.079777 | CONFORME | Medición
48 | 4 | PA25517468 | 0 | 0.0099413 | 0.5988 | 0.91743 | CONFORME | Medición
49 | 4 | PA25517469 | 0 | -1.1124 | -0.40914 | 0.47847 | CONFORME | Medición
50 | 4 | PA25517470 | 0 | -0.54629 | -0.10977 | 1.0766 | CONFORME | Medición
51 | 4 | PA25517471 | 0 | -0.69527 | -0.20956 | 0.67783 | CONFORME | Medición
52 | 4 | PA25517472 | 0 | -1.0628 | -0.30935 | 0.67783 | CONFORME | Medición
53 | 4 | PA25517473 | 0 | -0.61581 | 0.28939 | 1.874 | CONFORME | Medición
54 | 4 | PA25517474 | 0 | -0.13905 | -0.10977 | 0.47847 | CONFORME | Medición
55 | 4 | PA25517475 | 0 | -0.64561 | 0.1896 | 0.87719 | CONFORME | Medición
56 | 4 | PA25517476 | 0 | -1.321 | -0.8083 | -0.31898 | CONFORME | Medición
57 | 4 | PA25517477 | 0 | -0.43703 | -0.20956 | 0.47847 | CONFORME | Medición
58 | 4 | PA25517478 | 0 | -0.57608 | -0.20956 | 0.079745 | CONFORME | Medición
59 | 4 | PA25517479 | 0 | -0.35757 | -0.8083 | -0.91707 | CONFORME | Medición
60 | 4 | PA25517480 | 0 | -0.55622 | 0.089811 | 0.47847 | CONFORME | Medición
61 | 4 | PA25517481 | 0 | -0.28804 | 0.089811 | 1.0766 | CONFORME | Medición
62 | 4 | PA25517482 | 0 | -0.30791 | -0.40914 | -0.31898 | CONFORME | Medición
63 | 4 | PA25517483 | 0 | -0.24831 | 0.48897 | 1.6746 | CONFORME | Medición
64 | 4 | PA25517484 | 0 | -0.83433 | 0.48897 | 1.0766 | CONFORME | Medición
65 | 4 | PA25517485 | 0 | -0.5169 | 0.47895 | 1.8537 | CONFORME | Medición
66 | 4 | PA25517486 | 0 | -0.64612 | -0.21952 | 1.2557 | CONFORME | Medición
67 | 4 | PA25517487 | 0 | -0.31809 | -0.3193 | 0.25912 | CONFORME | Medición
68 | 4 | PA25517488 | 0 | -0.64612 | -0.21952 | -0.13953 | CONFORME | Medición
69 | 4 | PA25517489 | 0 | -0.21869 | 0.27939 | 1.2557 | CONFORME | Medición
70 | 4 | PA25517490 | 0 | -0.68588 | -0.11974 | 0.85709 | CONFORME | Medición
71 | 4 | PA25517491 | 0 | -0.53678 | 0.27939 | 0.85709 | CONFORME | Medición
72 | 4 | PA25517492 | 0 | 0.11928 | -0.11974 | 0.45844 | CONFORME | Medición
73 | 4 | PA25517493 | 0 | -0.36779 | 0.1796 | 0.65776 | CONFORME | Medición
74 | 4 | PA25517494 | 0 | -0.35785 | 0.67851 | 2.053 | CONFORME | Medición
75 | 4 | PA25517495 | 0 | -0.47714 | 0.67851 | 1.8537 | CONFORME | Medición
76 | 4 | PA25517496 | 0 | -0.57654 | 0.1796 | 1.2557 | CONFORME | Medición
77 | 4 | PA25517497 | 0 | -0.25845 | 0.77829 | 1.8537 | CONFORME | Medición
78 | 4 | PA25517498 | 0 | -0.31809 | 0.1796 | 1.0564 | CONFORME | Medición
79 | 4 | PA25517499 | 0 | -1.0239 | 0.27939 | 1.8537 | CONFORME | Medición
80 | 4 | PA25517500 | 0 | -0.59642 | 0.77829 | 1.8537 | CONFORME | Medición
81 | 4 | PA25517501 | 0 | -0.68534 | -0.30935 | 1.8943 | CONFORME | Medición
82 | 4 | PA25517502 | 0 | -0.60588 | -0.20956 | 0.4985 | CONFORME | Medición
83 | 4 | PA25517503 | 0 | -0.7052 | 0.1896 | 1.4955 | CONFORME | Medición
84 | 4 | PA25517504 | 0 | -0.98331 | 0.28939 | 1.0967 | CONFORME | Medición
85 | 4 | PA25517505 | 0 | -0.83433 | 0.28939 | 1.4955 | CONFORME | Medición
86 | 4 | PA25517506 | 0 | -0.72507 | 0.089811 | 1.4955 | CONFORME | Medición
87 | 4 | PA25517507 | 0 | -0.46683 | 0.089811 | 1.0967 | CONFORME | Medición
88 | 4 | PA25517508 | 0 | -0.75487 | 0.38918 | 1.4955 | CONFORME | Medición
89 | 4 | PA25517509 | 0 | -0.83433 | 0.48897 | 0.69791 | CONFORME | Medición
90 | 4 | PA25517510 | 0 | -0.63568 | 0.38918 | 2.0937 | CONFORME | Medición
91 | 4 | PA25517511 | 0 | -1.4601 | -0.60872 | 1.2961 | CONFORME | Medición
92 | 4 | PA25517512 | 0 | -0.18872 | 0.28939 | 0.89731 | CONFORME | Medición
93 | 4 | PA25517513 | 0 | -0.71514 | 0.089811 | 1.6949 | CONFORME | Medición
94 | 4 | PA25517514 | 0 | -1.3309 | 0.1896 | 0.69791 | CONFORME | Medición
95 | 4 | PA25517515 | 0 | 0.17878 | 0.38918 | 1.0967 | CONFORME | Medición
96 | 4 | PA25517516 | 0 | -0.90385 | 0.68855 | 2.0937 | CONFORME | Medición
97 | 4 | PA25517517 | 0 | -0.79444 | 0.10979 | 0.31911 | CONFORME | Medición
98 | 4 | PA25517518 | 0 | -0.73486 | 0.10979 | 1.5158 | CONFORME | Medición
99 | 4 | PA25517519 | 0 | -0.73486 | 0.009981 | 0.71799 | CONFORME | Medición
100 | 4 | PA25517520 | 0 | 0.57597 | 0.70865 | 1.7152 | CONFORME | Medición
101 | 4 | PA25517521 | 0 | -0.17875 | 1.2077 | 1.7152 | CONFORME | Medición
102 | 4 | PA25517522 | 0 | -0.32771 | 0.2096 | 0.71799 | CONFORME | Medición
103 | 4 | PA25517523 | 0 | -0.92354 | -0.18964 | 0.11966 | CONFORME | Medición
104 | 4 | PA25517524 | 0 | 0.33764 | 0.40922 | 1.3163 | CONFORME | Medición
105 | 4 | PA25517525 | 0 | -0.17875 | 0.80846 | 2.3135 | CONFORME | Medición
106 | 4 | PA25517526 | 0 | -0.089374 | 0.30941 | 0.91743 | CONFORME | Medición
107 | 4 | PA25517527 | 0 | 0.16882 | 0.009981 | 0.11966 | CONFORME | Medición
108 | 4 | PA25517528 | 0 | -0.64548 | 0.10979 | 0.11966 | CONFORME | Medición
109 | 4 | PA25517529 | 0 | -0.55611 | 0.40922 | 0.71799 | CONFORME | Medición
110 | 4 | PA25517530 | 0 | -0.38729 | -0.28945 | 0.31911 | CONFORME | Medición
111 | 4 | PA25517531 | 0 | -0.40715 | 0.30941 | 0.71799 | CONFORME | Medición
112 | 4 | PA25517532 | 0 | -0.86395 | 0.009981 | 0.91743 | CONFORME | Medición
113 | 4 | PA25517533 | 0 | -0.92372 | -0.049875 | 0.2991 | CONFORME | Medición
114 | 4 | PA25517534 | 0 | -1.4601 | -1.2469 | -2.2931 | CONFORME | Medición
115 | 4 | PA25517535 | 0 | -0.59595 | -0.14963 | 0.4985 | CONFORME | Medición
116 | 4 | PA25517536 | 0 | -0.59595 | 0.14963 | 0.099701 | CONFORME | Medición
117 | 4 | PA25517537 | 0 | -0.58602 | -0.049875 | -0.4985 | CONFORME | Medición
118 | 4 | PA25517538 | 0 | 0.16885 | 0.34913 | 1.0967 | CONFORME | Medición
119 | 4 | PA25517539 | 0 | -0.80453 | 0.34913 | 1.0967 | CONFORME | Medición
120 | 4 | PA25517540 | 0 | -1.0429 | -0.54863 | 0.099701 | CONFORME | Medición
121 | 4 | PA25517541 | 0 | -0.16885 | 0.34913 | 1.4955 | CONFORME | Medición
122 | 4 | PA25517542 | 0 | -0.65554 | -0.14963 | -0.099701 | CONFORME | Medición
123 | 4 | PA25517543 | 0 | -0.62574 | 0.14963 | 1.2961 | CONFORME | Medición
124 | 4 | PA25517544 | 0 | -0.74493 | -0.049875 | 0.4985 | CONFORME | Medición
125 | 4 | PA25517545 | 0 | -1.0926 | 0.24938 | 0.89731 | CONFORME | Medición
126 | 4 | PA25517546 | 0 | -0.66547 | 0.24938 | 1.2961 | CONFORME | Medición
127 | 4 | PA25517547 | 0 | -0.53635 | 0.14963 | 0.4985 | CONFORME | Medición
128 | 4 | PA25517548 | 0 | -0.049662 | 0.94763 | 1.8943 | CONFORME | Medición
129 | 4 | PA25517549 | 0 | -0.5666 | 0.48897 | 0.79681 | CONFORME | Medición
130 | 4 | PA25517550 | 0 | -1.1531 | 0.089811 | -0.1992 | CONFORME | Medición
131 | 4 | PA25517551 | 0 | 0.47714 | 0.38918 | 0.59761 | CONFORME | Medición
132 | 4 | PA25517552 | 0 | -0.30815 | -0.60872 | -0.39841 | CONFORME | Medición
133 | 4 | PA25517553 | 0 | -0.13917 | 0.38918 | 0.59761 | CONFORME | Medición
134 | 4 | PA25517554 | 0 | 0.13917 | -0.20956 | -0.1992 | CONFORME | Medición
135 | 4 | PA25517555 | 0 | 0.0099404 | 0.58876 | 0.59761 | CONFORME | Medición
136 | 4 | PA25517556 | 0 | -0.54672 | 0.68855 | 0.59761 | CONFORME | Medición
137 | 4 | PA25517557 | 0 | -0.94433 | 0.1896 | 2.1231e-13 | CONFORME | Medición
138 | 4 | PA25517558 | 0 | -0.0099404 | 0.1896 | 0.39841 | CONFORME | Medición
139 | 4 | PA25517559 | 0 | -0.85487 | 0.58876 | 0.39841 | CONFORME | Medición
140 | 4 | PA25517560 | 0 | -0.38767 | 0.28939 | 0.59761 | CONFORME | Medición
141 | 4 | PA25517561 | 0 | -0.68588 | 0.28939 | 0.59761 | CONFORME | Medición
142 | 4 | PA25517562 | 0 | -0.84493 | 0.28939 | 0.39841 | CONFORME | Medición
143 | 4 | PA25517563 | 0 | -0.31809 | 0.68855 | 0.99602 | CONFORME | Medición
144 | 4 | PA25517564 | 0 | -0.59642 | 0.68855 | 0.99602 | CONFORME | Medición
145 | 4 | PA25517565 | 0 | -1.2516 | -0.1996 | 0.89731 | CONFORME | Medición
146 | 4 | PA25517566 | 0 | -0.69534 | 0.6986 | 2.6919 | CONFORME | Medición
147 | 4 | PA25517567 | 0 | -0.80461 | 0.3992 | 1.4955 | CONFORME | Medición
148 | 4 | PA25517568 | 0 | -0.75494 | 0.3992 | 1.4955 | CONFORME | Medición
149 | 4 | PA25517569 | 0 | -0.95361 | 0.6986 | 1.6949 | CONFORME | Medición
150 | 4 | PA25517570 | 0 | -0.84434 | 0.3992 | 1.6949 | CONFORME | Medición
151 | 4 | PA25517571 | 0 | -1.0529 | -1.0978 | -1.8943 | CONFORME | Medición
152 | 4 | PA25517572 | 0 | -0.16887 | 0.3992 | 1.2961 | CONFORME | Medición
153 | 4 | PA25517573 | 0 | -0.62581 | 0.7984 | 1.4955 | CONFORME | Medición
154 | 4 | PA25517574 | 0 | -0.61587 | 0.3992 | 1.4955 | CONFORME | Medición
155 | 4 | PA25517575 | 0 | -0.98341 | -0.0998 | 0.69791 | CONFORME | Medición
156 | 4 | PA25517576 | 0 | -1.3112 | 0.3992 | 1.8943 | CONFORME | Medición
157 | 4 | PA25517577 | 0 | -1.3013 | -0.1996 | 0.099701 | CONFORME | Medición
158 | 4 | PA25517578 | 0 | -0.69534 | 0.7984 | 1.6949 | CONFORME | Medición
159 | 4 | PA25517579 | 0 | -0.60594 | 0.7984 | 1.6949 | CONFORME | Medición
160 | 4 | PA25517580 | 0 | -0.56621 | 0.998 | 2.2931 | CONFORME | Medición
161 | 4 | PA25517581 | 0 | -0.9009 | 0.1996 | 0.1996 | CONFORME | Medición
162 | 4 | PA25517582 | 0 | -1.3068 | 0.5988 | 0.3992 | CONFORME | Medición
163 | 4 | PA25517583 | 0 | -0.8415 | 0.5988 | 0.5988 | CONFORME | Medición
164 | 4 | PA25517584 | 0 | -0.7722 | 1.0637e-13 | -1.7728e-13 | CONFORME | Medición
165 | 4 | PA25517585 | 0 | -0.8514 | -1.7728e-13 | -0.1996 | CONFORME | Medición
166 | 4 | PA25517586 | 0 | -0.9603 | 1.0637e-13 | -1.7728e-13 | CONFORME | Medición
167 | 4 | PA25517587 | 0 | -0.2079 | 0.7984 | 0.998 | CONFORME | Medición
168 | 4 | PA25517588 | 0 | -0.6534 | -0.1996 | 0.1996 | CONFORME | Medición
169 | 4 | PA25517589 | 0 | -0.1485 | 0.3992 | 0.998 | CONFORME | Medición
170 | 4 | PA25517590 | 0 | -0.9702 | -0.3992 | -0.5988 | CONFORME | Medición
171 | 4 | PA25517591 | 0 | 0.3861 | 0.8982 | 0.998 | CONFORME | Medición
172 | 4 | PA25517592 | 0 | -0.4554 | 0.7984 | 0.998 | CONFORME | Medición
173 | 4 | PA25517593 | 0 | -0.8514 | 1.0637e-13 | 0.1996 | CONFORME | Medición
174 | 4 | PA25517594 | 0 | -0.5148 | -0.1996 | -0.1996 | CONFORME | Medición
175 | 4 | PA25517595 | 0 | -0.7821 | 0.1996 | 0.5988 | CONFORME | Medición
176 | 4 | PA25517596 | 0 | 0.2871 | 0.2994 | 1.996 | CONFORME | Medición
177 | 4 | PA25517597 | 0 | -1.1322 | -0.7984 | -1.0178 | CONFORME | Medición
178 | 4 | PA25517598 | 0 | -0.73493 | -0.5988 | -0.41908 | CONFORME | Medición
179 | 4 | PA25517599 | 0 | -0.94349 | -0.2994 | -0.019956 | CONFORME | Medición
180 | 4 | PA25517600 | 0 | -0.76472 | -0.0998 | -0.019956 | CONFORME | Medición
181 | 4 | PA25517401 | 0 | -0.64555 | -0.3992 | -0.21952 | CONFORME | Medición
182 | 4 | PA25517402 | 0 | -0.79452 | -0.499 | 0.1796 | CONFORME | Medición
183 | 4 | PA25517403 | 0 | -0.27808 | -0.2994 | -0.21952 | CONFORME | Medición
184 | 4 | PA25517404 | 0 | -0.57603 | -0.3992 | -1.6164 | CONFORME | Medición
185 | 4 | PA25517405 | 0 | -1.1421 | -0.7984 | -1.2173 | CONFORME | Medición
186 | 4 | PA25517406 | 0 | -0.63561 | -0.0998 | -0.019956 | CONFORME | Medición
187 | 4 | PA25517407 | 0 | -0.51644 | -0.0998 | 0.57873 | CONFORME | Medición
188 | 4 | PA25517408 | 0 | -0.1887 | 1.0637e-13 | 0.1796 | CONFORME | Medición
189 | 4 | PA25517409 | 0 | -0.9137 | -0.6986 | -0.8182 | CONFORME | Medición
190 | 4 | PA25517410 | 0 | -0.049657 | 0.0998 | 0.37917 | CONFORME | Medición
191 | 4 | PA25517411 | 0 | 0.22842 | 0.6986 | 0.97785 | CONFORME | Medición
192 | 4 | PA25517412 | 0 | -1.1421 | -0.499 | -0.41908 | CONFORME | Medición
193 | 4 | PA25517413 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 4 | PA25517414 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 4 | PA25517415 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 4 | PA25517416 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 4 | PA25517417 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 4 | PA25517418 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 4 | PA25517419 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 4 | PA25517420 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 4 | PA25517421 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 4 | PA25517422 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 4 | PA25517423 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 4 | PA25517424 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 4 | PA25517425 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 4 | PA25517426 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 4 | PA25517427 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 4 | PA25517428 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 4 | PA25517429 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 4 | PA25517430 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 4 | PA25517431 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 4 | PA25517432 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 4 | PA25517433 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 4 | PA25517434 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 4 | PA25517435 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 4 | PA25517436 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 4 | PA25517437 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 4 | PA25517438 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 4 | PA25517439 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 4 | PA25517440 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 4 | PA25517441 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 4 | PA25517442 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 4 | PA25517443 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 4 | PA25517444 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 4 | PA25517445 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 4 | PA25517446 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 4 | PA25517447 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 4 | PA25517448 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 4 | PA25517449 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 4 | PA25517450 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 4 | PA25517451 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 4 | PA25517452 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 4 | PA25517453 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 4 | PA25517454 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 4 | PA25517455 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 4 | PA25517456 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 4 | PA25517457 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 4 | PA25517458 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 4 | PA25517459 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 4 | PA25517460 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 4 | PA25517461 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 4 | PA25517462 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 4 | PA25517463 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 4 | PA25517464 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 4 | PA25517465 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 4 | PA25517466 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 4 | PA25517467 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 4 | PA25517468 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 4 | PA25517469 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 4 | PA25517470 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 4 | PA25517471 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 4 | PA25517472 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 4 | PA25517473 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 4 | PA25517474 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 4 | PA25517475 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 4 | PA25517476 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 4 | PA25517477 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 4 | PA25517478 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 4 | PA25517479 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 4 | PA25517480 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 4 | PA25517481 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 4 | PA25517482 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 4 | PA25517483 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 4 | PA25517484 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 4 | PA25517485 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 4 | PA25517486 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 4 | PA25517487 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 4 | PA25517488 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 4 | PA25517489 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 4 | PA25517490 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 4 | PA25517491 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 4 | PA25517492 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 4 | PA25517493 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 4 | PA25517494 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 4 | PA25517495 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 4 | PA25517496 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 4 | PA25517497 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 4 | PA25517498 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 4 | PA25517499 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 4 | PA25517500 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 4 | PA25517501 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 4 | PA25517502 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 4 | PA25517503 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 4 | PA25517504 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 4 | PA25517505 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 4 | PA25517506 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 4 | PA25517507 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 4 | PA25517508 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 4 | PA25517509 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 4 | PA25517510 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 4 | PA25517511 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 4 | PA25517512 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 4 | PA25517513 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 4 | PA25517514 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 4 | PA25517715 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 4 | PA25517716 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 4 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición