## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516921 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 116.98 | 217.6 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.089366 | 21 | 1.508 | 1.505 | 229.54 | 239.5 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.688 | 21.5 | 1.508 | 1.503 | 239.5 | 244.55 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 0.49751 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516922 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 115.2 | 215.04 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.86387 | 21 | 1.508 | 1.505 | 226.84 | 236.65 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -2.1837 | 21.5 | 1.508 | 1.503 | 236.65 | 241.58 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | -1.8905 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
3 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516923 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 115.3 | 215.5 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.5064 | 21 | 1.508 | 1.505 | 227.4 | 237.35 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.78772 | 21.5 | 1.508 | 1.503 | 237.35 | 242.35 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | -0.49751 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
4 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516924 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 114.47 | 214.7 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.47662 | 21 | 1.508 | 1.505 | 226.6 | 236.55 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.78772 | 21.5 | 1.508 | 1.503 | 236.55 | 241.6 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 0.49751 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516925 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 115.69 | 216 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.39718 | 21 | 1.508 | 1.505 | 227.9 | 237.86 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.688 | 21.5 | 1.508 | 1.503 | 237.86 | 242.89 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 0.099502 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516926 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 115.55 | 216.52 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | 0.25817 | 21 | 1.508 | 1.505 | 228.45 | 238.41 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.688 | 21.5 | 1.508 | 1.503 | 238.41 | 243.52 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 1.6915 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516927 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 116.11 | 216.65 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.1688 | 21 | 1.508 | 1.505 | 228.55 | 238.5 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.78772 | 21.5 | 1.508 | 1.503 | 238.5 | 243.54 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 0.29851 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516928 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 113.8 | 214.49 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.019859 | 21 | 1.508 | 1.505 | 226.5 | 236.47 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.58829 | 21.5 | 1.508 | 1.503 | 236.47 | 241.53 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 0.69652 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516929 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 114.14 | 214.3 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | -0.54612 | 21 | 1.508 | 1.505 | 226.18 | 236.1 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -1.0868 | 21.5 | 1.508 | 1.503 | 236.1 | 241.07 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | -1.0945 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516930 | 25.6 | 0 | 21.2 | 8.08 | 0.573 | 114.32 | 215.16 | 100.71 | 2,28,370 | 2 | 28,370 | 0.041214 | 2443.5937 | 0.12908 | 21 | 1.508 | 1.505 | 227.17 | 237.19 | 10.029 | 23,11,780 | 23 | 11,780 | 0.38661 | 25.9412 | -0.08974 | 21.5 | 1.508 | 1.503 | 237.19 | 242.29 | 5.025 | 18,43,300 | 18 | 43,300 | 0.31203 | 16.1043 | 1.4925 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516931 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
12 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516932 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
13 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516933 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
14 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516934 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
15 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516935 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
16 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516936 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
17 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516937 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
18 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516938 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
19 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516939 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
20 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516940 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
21 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516941 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
22 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516942 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
23 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516943 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
24 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516944 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
25 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516945 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
26 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516946 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
27 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516947 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
28 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516948 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
29 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516949 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
30 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516950 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
31 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516951 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
32 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516952 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
33 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516953 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
34 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516954 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
35 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516955 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
36 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516956 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
37 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516957 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
38 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516958 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
39 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516959 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
40 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516960 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
41 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516961 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
42 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516962 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
43 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516963 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
44 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516964 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
45 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516965 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
46 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516966 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
47 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516967 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
48 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516968 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
49 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516969 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
50 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516970 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
51 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516971 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
52 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516972 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
53 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516973 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
54 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516974 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
55 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516975 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
56 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516976 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
57 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516977 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
58 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516978 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
59 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516979 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
60 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516980 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
61 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516981 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
62 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516982 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
63 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516983 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
64 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516984 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
65 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516985 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
66 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516986 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
67 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516987 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
68 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516988 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
69 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516989 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
70 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516990 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
71 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516991 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
72 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516992 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
73 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516993 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
74 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516994 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
75 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516995 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
76 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516996 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
77 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516997 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
78 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516998 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
79 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25516999 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
80 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517000 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
81 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517001 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
82 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517002 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
83 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517003 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
84 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517004 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
85 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517005 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
86 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517006 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
87 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517007 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
88 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517008 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
89 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517009 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
90 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517010 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
91 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517011 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
92 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517012 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
93 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517013 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
94 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517014 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
95 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517015 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
96 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517016 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
97 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517017 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
98 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517018 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
99 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517019 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
100 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517020 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
101 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517021 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
102 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517022 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
103 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517023 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
104 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517024 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
105 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517025 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
106 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517026 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
107 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517027 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
108 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517028 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
109 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517029 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
110 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517030 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
111 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517031 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
112 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517032 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
113 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517033 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
114 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517034 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
115 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517035 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
116 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517036 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
117 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517037 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
118 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517038 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
119 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517039 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
120 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517040 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
121 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517041 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
122 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517042 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
123 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517043 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
124 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517044 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
125 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517045 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
126 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517046 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
127 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517047 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
128 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517048 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
129 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517049 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
130 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517050 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
131 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517051 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
132 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517052 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
133 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517053 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
134 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517054 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
135 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517055 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
136 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517056 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
137 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517057 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
138 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517058 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
139 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517059 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
140 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517060 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
141 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517061 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
142 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517062 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
143 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517063 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
144 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517064 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
145 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517065 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
146 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517066 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
147 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517067 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
148 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517068 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
149 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517069 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
150 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517070 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
151 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517071 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
152 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517072 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
153 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517073 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
154 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517074 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
155 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517075 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
156 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517076 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
157 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517077 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
158 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517078 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
159 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517079 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
160 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517080 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
161 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517081 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
162 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517082 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
163 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517083 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
164 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517084 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
165 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517085 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
166 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517086 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
167 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517087 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
168 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517088 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
169 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517089 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
170 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517090 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
171 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517091 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
172 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517092 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
173 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517093 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
174 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517094 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
175 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517095 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
176 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517096 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
177 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517097 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
178 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517098 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
179 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517099 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
180 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517100 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
181 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517101 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
182 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517102 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
183 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517103 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
184 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517104 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
185 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517105 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
186 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517106 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
187 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517107 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
188 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517108 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
189 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517109 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
190 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517110 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
191 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517111 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
192 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517112 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
193 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517113 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
194 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517114 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
195 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517115 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
196 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517116 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
197 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517117 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
198 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517118 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
199 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517119 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
200 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517120 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
201 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517121 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
202 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517122 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
203 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517123 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
204 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517124 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
205 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517125 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
206 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517126 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
207 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517127 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
208 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517128 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
209 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517129 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
210 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517130 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
211 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517131 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
212 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517132 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
213 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517133 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
214 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517134 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
215 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517135 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
216 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517136 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
217 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517137 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
218 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517138 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
219 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517139 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
220 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517140 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
221 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517141 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
222 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517142 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
223 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517143 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
224 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517144 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
225 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517145 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
226 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517146 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
227 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517147 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
228 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517148 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
229 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517149 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
230 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517150 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
231 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517151 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
232 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517152 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
233 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517153 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
234 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517154 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
235 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517155 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
236 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517156 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
237 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517157 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
238 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517158 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
239 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517159 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
240 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517160 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
241 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517161 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
242 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517162 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
243 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517163 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
244 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517164 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
245 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517165 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
246 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517166 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
247 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517167 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
248 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517168 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
249 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517169 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
250 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517170 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
251 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517171 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
252 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517172 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
253 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517173 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
254 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517174 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
255 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517175 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
256 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517176 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
257 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517177 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
258 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517178 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
259 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517179 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
260 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517180 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
261 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517181 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
262 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517182 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
263 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517183 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
264 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517184 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
265 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517185 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
266 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517186 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
267 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517187 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
268 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517188 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
269 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517189 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
270 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517190 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
271 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517191 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
272 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517192 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
273 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517193 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
274 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517194 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
275 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517195 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
276 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517196 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
277 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517197 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
278 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517198 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
279 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517199 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
280 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517200 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
281 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517201 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
282 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517202 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
283 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517203 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
284 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517204 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
285 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517205 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
286 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517206 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
287 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517207 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
288 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517208 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
289 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517209 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
290 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517210 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
291 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517211 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
292 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517212 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
293 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517213 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
294 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517214 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
295 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517215 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
296 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517216 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
297 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517217 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
298 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517218 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
299 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517219 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
300 | 2025-09-19 | 2025-09-19 | 5 | 10 |  | PA25517220 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
1
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 5 | PA25516921 | 0 | -0.089366 | -0.688 | 0.49751 | CONFORME | Medición
2 | 5 | PA25516922 | 0 | -0.86387 | -2.1837 | -1.8905 | NO CONFORME | Medición
3 | 5 | PA25516923 | 0 | -0.5064 | -0.78772 | -0.49751 | CONFORME | Medición
4 | 5 | PA25516924 | 0 | -0.47662 | -0.78772 | 0.49751 | CONFORME | Medición
5 | 5 | PA25516925 | 0 | -0.39718 | -0.688 | 0.099502 | CONFORME | Medición
6 | 5 | PA25516926 | 0 | 0.25817 | -0.688 | 1.6915 | CONFORME | Medición
7 | 5 | PA25516927 | 0 | -0.1688 | -0.78772 | 0.29851 | CONFORME | Medición
8 | 5 | PA25516928 | 0 | -0.019859 | -0.58829 | 0.69652 | CONFORME | Medición
9 | 5 | PA25516929 | 0 | -0.54612 | -1.0868 | -1.0945 | CONFORME | Medición
10 | 5 | PA25516930 | 0 | 0.12908 | -0.08974 | 1.4925 | CONFORME | Medición
11 | 5 | PA25516931 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
12 | 5 | PA25516932 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
13 | 5 | PA25516933 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
14 | 5 | PA25516934 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
15 | 5 | PA25516935 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
16 | 5 | PA25516936 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
17 | 5 | PA25516937 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
18 | 5 | PA25516938 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
19 | 5 | PA25516939 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
20 | 5 | PA25516940 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
21 | 5 | PA25516941 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
22 | 5 | PA25516942 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
23 | 5 | PA25516943 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
24 | 5 | PA25516944 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
25 | 5 | PA25516945 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
26 | 5 | PA25516946 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
27 | 5 | PA25516947 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
28 | 5 | PA25516948 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
29 | 5 | PA25516949 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
30 | 5 | PA25516950 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
31 | 5 | PA25516951 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
32 | 5 | PA25516952 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
33 | 5 | PA25516953 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
34 | 5 | PA25516954 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
35 | 5 | PA25516955 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
36 | 5 | PA25516956 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
37 | 5 | PA25516957 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
38 | 5 | PA25516958 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
39 | 5 | PA25516959 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
40 | 5 | PA25516960 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
41 | 5 | PA25516961 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
42 | 5 | PA25516962 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
43 | 5 | PA25516963 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
44 | 5 | PA25516964 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
45 | 5 | PA25516965 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
46 | 5 | PA25516966 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
47 | 5 | PA25516967 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
48 | 5 | PA25516968 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
49 | 5 | PA25516969 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
50 | 5 | PA25516970 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
51 | 5 | PA25516971 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
52 | 5 | PA25516972 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
53 | 5 | PA25516973 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
54 | 5 | PA25516974 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
55 | 5 | PA25516975 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
56 | 5 | PA25516976 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
57 | 5 | PA25516977 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
58 | 5 | PA25516978 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
59 | 5 | PA25516979 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
60 | 5 | PA25516980 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
61 | 5 | PA25516981 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
62 | 5 | PA25516982 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
63 | 5 | PA25516983 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
64 | 5 | PA25516984 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
65 | 5 | PA25516985 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
66 | 5 | PA25516986 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
67 | 5 | PA25516987 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
68 | 5 | PA25516988 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
69 | 5 | PA25516989 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
70 | 5 | PA25516990 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
71 | 5 | PA25516991 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
72 | 5 | PA25516992 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
73 | 5 | PA25516993 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
74 | 5 | PA25516994 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
75 | 5 | PA25516995 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
76 | 5 | PA25516996 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
77 | 5 | PA25516997 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
78 | 5 | PA25516998 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
79 | 5 | PA25516999 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
80 | 5 | PA25517000 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
81 | 5 | PA25517001 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
82 | 5 | PA25517002 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
83 | 5 | PA25517003 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
84 | 5 | PA25517004 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
85 | 5 | PA25517005 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
86 | 5 | PA25517006 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
87 | 5 | PA25517007 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
88 | 5 | PA25517008 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
89 | 5 | PA25517009 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
90 | 5 | PA25517010 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
91 | 5 | PA25517011 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
92 | 5 | PA25517012 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
93 | 5 | PA25517013 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
94 | 5 | PA25517014 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
95 | 5 | PA25517015 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
96 | 5 | PA25517016 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
97 | 5 | PA25517017 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
98 | 5 | PA25517018 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
99 | 5 | PA25517019 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
100 | 5 | PA25517020 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
101 | 5 | PA25517021 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
102 | 5 | PA25517022 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
103 | 5 | PA25517023 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
104 | 5 | PA25517024 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
105 | 5 | PA25517025 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
106 | 5 | PA25517026 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
107 | 5 | PA25517027 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
108 | 5 | PA25517028 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
109 | 5 | PA25517029 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
110 | 5 | PA25517030 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
111 | 5 | PA25517031 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
112 | 5 | PA25517032 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
113 | 5 | PA25517033 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
114 | 5 | PA25517034 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
115 | 5 | PA25517035 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
116 | 5 | PA25517036 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
117 | 5 | PA25517037 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
118 | 5 | PA25517038 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
119 | 5 | PA25517039 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
120 | 5 | PA25517040 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
121 | 5 | PA25517041 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
122 | 5 | PA25517042 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
123 | 5 | PA25517043 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
124 | 5 | PA25517044 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
125 | 5 | PA25517045 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
126 | 5 | PA25517046 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
127 | 5 | PA25517047 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
128 | 5 | PA25517048 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
129 | 5 | PA25517049 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
130 | 5 | PA25517050 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
131 | 5 | PA25517051 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
132 | 5 | PA25517052 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
133 | 5 | PA25517053 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
134 | 5 | PA25517054 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
135 | 5 | PA25517055 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
136 | 5 | PA25517056 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
137 | 5 | PA25517057 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
138 | 5 | PA25517058 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
139 | 5 | PA25517059 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
140 | 5 | PA25517060 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
141 | 5 | PA25517061 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
142 | 5 | PA25517062 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
143 | 5 | PA25517063 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
144 | 5 | PA25517064 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
145 | 5 | PA25517065 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
146 | 5 | PA25517066 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
147 | 5 | PA25517067 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
148 | 5 | PA25517068 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
149 | 5 | PA25517069 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
150 | 5 | PA25517070 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
151 | 5 | PA25517071 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
152 | 5 | PA25517072 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
153 | 5 | PA25517073 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
154 | 5 | PA25517074 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
155 | 5 | PA25517075 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
156 | 5 | PA25517076 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
157 | 5 | PA25517077 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
158 | 5 | PA25517078 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
159 | 5 | PA25517079 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
160 | 5 | PA25517080 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
161 | 5 | PA25517081 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
162 | 5 | PA25517082 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
163 | 5 | PA25517083 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
164 | 5 | PA25517084 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
165 | 5 | PA25517085 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
166 | 5 | PA25517086 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
167 | 5 | PA25517087 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
168 | 5 | PA25517088 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
169 | 5 | PA25517089 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
170 | 5 | PA25517090 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
171 | 5 | PA25517091 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
172 | 5 | PA25517092 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
173 | 5 | PA25517093 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
174 | 5 | PA25517094 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
175 | 5 | PA25517095 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
176 | 5 | PA25517096 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
177 | 5 | PA25517097 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
178 | 5 | PA25517098 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
179 | 5 | PA25517099 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
180 | 5 | PA25517100 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
181 | 5 | PA25517101 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
182 | 5 | PA25517102 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
183 | 5 | PA25517103 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
184 | 5 | PA25517104 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
185 | 5 | PA25517105 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
186 | 5 | PA25517106 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
187 | 5 | PA25517107 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
188 | 5 | PA25517108 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
189 | 5 | PA25517109 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
190 | 5 | PA25517110 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
191 | 5 | PA25517111 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
192 | 5 | PA25517112 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
193 | 5 | PA25517113 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
194 | 5 | PA25517114 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
195 | 5 | PA25517115 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
196 | 5 | PA25517116 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
197 | 5 | PA25517117 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
198 | 5 | PA25517118 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
199 | 5 | PA25517119 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
200 | 5 | PA25517120 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
201 | 5 | PA25517121 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
202 | 5 | PA25517122 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
203 | 5 | PA25517123 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
204 | 5 | PA25517124 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
205 | 5 | PA25517125 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
206 | 5 | PA25517126 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
207 | 5 | PA25517127 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
208 | 5 | PA25517128 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
209 | 5 | PA25517129 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
210 | 5 | PA25517130 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
211 | 5 | PA25517131 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
212 | 5 | PA25517132 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
213 | 5 | PA25517133 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
214 | 5 | PA25517134 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
215 | 5 | PA25517135 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
216 | 5 | PA25517136 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
217 | 5 | PA25517137 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
218 | 5 | PA25517138 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
219 | 5 | PA25517139 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
220 | 5 | PA25517140 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
221 | 5 | PA25517141 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
222 | 5 | PA25517142 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
223 | 5 | PA25517143 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
224 | 5 | PA25517144 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
225 | 5 | PA25517145 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
226 | 5 | PA25517146 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
227 | 5 | PA25517147 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
228 | 5 | PA25517148 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
229 | 5 | PA25517149 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
230 | 5 | PA25517150 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
231 | 5 | PA25517151 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
232 | 5 | PA25517152 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
233 | 5 | PA25517153 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
234 | 5 | PA25517154 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
235 | 5 | PA25517155 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
236 | 5 | PA25517156 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
237 | 5 | PA25517157 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
238 | 5 | PA25517158 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
239 | 5 | PA25517159 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
240 | 5 | PA25517160 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
241 | 5 | PA25517161 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
242 | 5 | PA25517162 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
243 | 5 | PA25517163 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
244 | 5 | PA25517164 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
245 | 5 | PA25517165 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
246 | 5 | PA25517166 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
247 | 5 | PA25517167 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
248 | 5 | PA25517168 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
249 | 5 | PA25517169 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
250 | 5 | PA25517170 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
251 | 5 | PA25517171 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
252 | 5 | PA25517172 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
253 | 5 | PA25517173 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
254 | 5 | PA25517174 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
255 | 5 | PA25517175 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
256 | 5 | PA25517176 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
257 | 5 | PA25517177 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
258 | 5 | PA25517178 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
259 | 5 | PA25517179 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
260 | 5 | PA25517180 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
261 | 5 | PA25517181 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
262 | 5 | PA25517182 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
263 | 5 | PA25517183 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
264 | 5 | PA25517184 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
265 | 5 | PA25517185 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
266 | 5 | PA25517186 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
267 | 5 | PA25517187 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
268 | 5 | PA25517188 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
269 | 5 | PA25517189 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
270 | 5 | PA25517190 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
271 | 5 | PA25517191 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
272 | 5 | PA25517192 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
273 | 5 | PA25517193 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
274 | 5 | PA25517194 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
275 | 5 | PA25517195 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
276 | 5 | PA25517196 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
277 | 5 | PA25517197 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
278 | 5 | PA25517198 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
279 | 5 | PA25517199 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
280 | 5 | PA25517200 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
281 | 5 | PA25517201 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
282 | 5 | PA25517202 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
283 | 5 | PA25517203 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
284 | 5 | PA25517204 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
285 | 5 | PA25517205 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
286 | 5 | PA25517206 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
287 | 5 | PA25517207 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
288 | 5 | PA25517208 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
289 | 5 | PA25517209 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
290 | 5 | PA25517210 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
291 | 5 | PA25517211 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
292 | 5 | PA25517212 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
293 | 5 | PA25517213 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
294 | 5 | PA25517214 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
295 | 5 | PA25517215 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
296 | 5 | PA25517216 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
297 | 5 | PA25517217 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
298 | 5 | PA25517218 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
299 | 5 | PA25517219 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
300 | 5 | PA25517220 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición