## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517736 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 528.15 | 628.64 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -0.40634 | 20.83 | 1.513 | 1.513 | 669.38 | 679.38 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.18964 | 21.58 | 1.516 | 1.515 | 679.38 | 684.4 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.13964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517737 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 523.89 | 623.5 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.2785 | 20.83 | 1.513 | 1.513 | 663.94 | 673.9 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.58888 | 21.58 | 1.516 | 1.515 | 673.9 | 678.95 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.73808 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517738 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 521.74 | 621.17 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.4569 | 20.83 | 1.513 | 1.513 | 661.61 | 671.65 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | 0.2096 | 21.58 | 1.516 | 1.515 | 671.65 | 676.76 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 1.935 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517739 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 526.74 | 626.91 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -0.72349 | 20.83 | 1.513 | 1.513 | 667.52 | 677.45 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.88831 | 21.58 | 1.516 | 1.515 | 677.45 | 682.45 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | -0.25933 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517740 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 524.54 | 624.35 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.0803 | 20.83 | 1.513 | 1.513 | 664.76 | 674.7 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.7885 | 21.58 | 1.516 | 1.515 | 674.7 | 679.74 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.5386 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517741 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 525.1 | 625.2 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -0.79286 | 20.83 | 1.513 | 1.513 | 665.69 | 675.61 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.98812 | 21.58 | 1.516 | 1.515 | 675.61 | 680.64 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.33912 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517742 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 524.66 | 624.16 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.3875 | 20.83 | 1.513 | 1.513 | 664.51 | 674.46 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.68869 | 21.58 | 1.516 | 1.515 | 674.46 | 679.5 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.5386 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517743 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 526.44 | 626.47 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -0.86224 | 20.83 | 1.513 | 1.513 | 667.01 | 677.02 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.089829 | 21.58 | 1.516 | 1.515 | 677.02 | 682.1 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 1.3365 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517744 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 528.1 | 628.6 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -0.39643 | 20.83 | 1.513 | 1.513 | 669.24 | 679.25 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.089829 | 21.58 | 1.516 | 1.515 | 679.25 | 684.3 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.73808 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517745 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 523.86 | 623.53 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.219 | 20.83 | 1.513 | 1.513 | 663.9 | 673.9 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.18964 | 21.58 | 1.516 | 1.515 | 673.9 | 678.95 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.73808 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517746 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 525.15 | 624.98 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.0605 | 20.83 | 1.513 | 1.513 | 665.4 | 675.37 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.48907 | 21.58 | 1.516 | 1.515 | 675.37 | 680.4 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.33912 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517747 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 526.1 | 625.8 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.1893 | 20.83 | 1.513 | 1.513 | 666.34 | 676.32 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.38926 | 21.58 | 1.516 | 1.515 | 676.32 | 681.36 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 0.5386 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517748 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 532.27 | 633.04 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -0.12884 | 20.83 | 1.513 | 1.513 | 673.9 | 683.95 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | 0.30941 | 21.58 | 1.516 | 1.515 | 683.95 | 689.02 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 1.137 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517749 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 525.9 | 625.66 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | -1.1298 | 20.83 | 1.513 | 1.513 | 666.03 | 675.96 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | -0.88831 | 21.58 | 1.516 | 1.515 | 675.96 | 680.95 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | -0.45881 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517750 | 25.6 | 0 | 22.23 | 7.648 | 0.539 | 527.61 | 628.57 | 100.9 | 2,32,832 | 2 | 32,832 | 0.042453 | 2376.7274 | 0.059465 | 20.83 | 1.513 | 1.513 | 669.54 | 679.64 | 10.019 | 22,45,010 | 22 | 45,010 | 0.37917 | 26.4235 | 0.80846 | 21.58 | 1.516 | 1.515 | 679.64 | 684.71 | 5.013 | 18,30,200 | 18 | 30,200 | 0.30839 | 16.2554 | 1.137 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517751 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 298.05 | 398.6 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.1494 | 21.85 | 1.512 | 1.507 | 435.91 | 445.95 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | 0.1996 | 22.43 | 1.514 | 1.511 | 445.95 | 451.02 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517752 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 295.73 | 396.06 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.069721 | 21.85 | 1.512 | 1.507 | 433.3 | 443.3 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.1996 | 22.43 | 1.514 | 1.511 | 443.3 | 448.35 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517753 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 294.21 | 394.1 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.50797 | 21.85 | 1.512 | 1.507 | 431.44 | 441.45 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.0998 | 22.43 | 1.514 | 1.511 | 441.45 | 446.5 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517754 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 294.35 | 394.6 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.1494 | 21.85 | 1.512 | 1.507 | 431.81 | 441.77 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.5988 | 22.43 | 1.514 | 1.511 | 441.77 | 446.8 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517755 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 298.26 | 399.1 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.43825 | 21.85 | 1.512 | 1.507 | 436.41 | 446.45 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | 0.1996 | 22.43 | 1.514 | 1.511 | 446.45 | 451.55 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517756 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 298.84 | 399.25 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.0099602 | 21.85 | 1.512 | 1.507 | 436.41 | 446.4 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.2994 | 22.43 | 1.514 | 1.511 | 446.4 | 451.41 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517757 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 291.8 | 390.75 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -1.4442 | 21.85 | 1.512 | 1.507 | 427.44 | 437.36 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.998 | 22.43 | 1.514 | 1.511 | 437.36 | 442.38 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517758 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 295.25 | 394.8 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.84661 | 21.85 | 1.512 | 1.507 | 431.8 | 441.79 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.2994 | 22.43 | 1.514 | 1.511 | 441.79 | 446.84 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517759 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 296.38 | 396.58 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.1992 | 21.85 | 1.512 | 1.507 | 433.71 | 443.65 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.7984 | 22.43 | 1.514 | 1.511 | 443.65 | 448.66 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517760 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 296.69 | 397.23 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.13944 | 21.85 | 1.512 | 1.507 | 434.35 | 444.34 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.2994 | 22.43 | 1.514 | 1.511 | 444.34 | 449.4 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517761 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 293.3 | 393.74 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.039841 | 21.85 | 1.512 | 1.507 | 430.91 | 440.95 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | 0.1996 | 22.43 | 1.514 | 1.511 | 440.95 | 446.03 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
27 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517762 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 298.77 | 399.58 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.40837 | 21.85 | 1.512 | 1.507 | 436.9 | 446.91 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -0.0998 | 22.43 | 1.514 | 1.511 | 446.91 | 451.98 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517763 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 297.25 | 397.6 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.049801 | 21.85 | 1.512 | 1.507 | 434.71 | 444.75 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | 0.1996 | 22.43 | 1.514 | 1.511 | 444.75 | 449.81 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517764 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 295.1 | 395.25 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -0.249 | 21.85 | 1.512 | 1.507 | 432.46 | 442.53 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | 0.499 | 22.43 | 1.514 | 1.511 | 442.53 | 447.6 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-20 | 2025-09-20 | 5 | 0 |  | PA25517765 | 25.6 | 0 | 21.75 | 7.643 | 0.517 | 146.22 | 246.66 | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | 0.039841 | 21.85 | 1.512 | 1.507 | 283.88 | 293.92 | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | 0.1996 | 22.43 | 1.514 | 1.511 | 293.92 | 299.01 | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | 1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
31 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
32 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
33 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
34 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
35 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
36 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
37 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
38 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
39 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
40 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
41 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
42 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
43 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
44 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
45 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
46 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
47 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
48 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
49 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
50 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
51 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
52 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
53 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
54 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
55 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
56 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
57 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
58 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
59 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
60 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
61 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
62 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
63 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
64 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
65 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
66 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
67 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
68 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
69 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
70 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
71 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
72 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
73 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
74 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
75 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
76 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
77 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
78 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
79 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
80 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
81 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
82 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
83 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
84 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
85 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
86 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
87 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
88 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
89 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
90 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
91 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
92 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
93 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
94 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
95 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
96 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
97 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
98 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
99 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
100 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
101 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
102 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
103 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
104 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
105 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
106 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
107 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
108 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
109 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
110 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
111 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
112 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
113 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
114 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
115 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
116 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
117 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
118 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
119 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
120 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
121 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-20 | 2025-09-20 | 5 | 0 |  |  | 25.6 | 0 | 21.75 | 7.643 | 0.517 |  |  | 100.4 | 2,31,960 | 2 | 31,960 | 0.042211 | 2378.5207 | -100 | 21.85 | 1.512 | 1.507 |  |  | 10.02 | 22,43,730 | 22 | 43,730 | 0.37881 | 26.451 | -100 | 22.43 | 1.514 | 1.511 | 0 |  | 5.015 | 18,20,640 | 18 | 20,640 | 0.30573 | 16.4032 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 5 | PA25517736 | 0 | -0.40634 | -0.18964 | 0.13964 | CONFORME | Medición
2 | 5 | PA25517737 | 0 | -1.2785 | -0.58888 | 0.73808 | CONFORME | Medición
3 | 5 | PA25517738 | 0 | -1.4569 | 0.2096 | 1.935 | CONFORME | Medición
4 | 5 | PA25517739 | 0 | -0.72349 | -0.88831 | -0.25933 | CONFORME | Medición
5 | 5 | PA25517740 | 0 | -1.0803 | -0.7885 | 0.5386 | CONFORME | Medición
6 | 5 | PA25517741 | 0 | -0.79286 | -0.98812 | 0.33912 | CONFORME | Medición
7 | 5 | PA25517742 | 0 | -1.3875 | -0.68869 | 0.5386 | CONFORME | Medición
8 | 5 | PA25517743 | 0 | -0.86224 | -0.089829 | 1.3365 | CONFORME | Medición
9 | 5 | PA25517744 | 0 | -0.39643 | -0.089829 | 0.73808 | CONFORME | Medición
10 | 5 | PA25517745 | 0 | -1.219 | -0.18964 | 0.73808 | CONFORME | Medición
11 | 5 | PA25517746 | 0 | -1.0605 | -0.48907 | 0.33912 | CONFORME | Medición
12 | 5 | PA25517747 | 0 | -1.1893 | -0.38926 | 0.5386 | CONFORME | Medición
13 | 5 | PA25517748 | 0 | -0.12884 | 0.30941 | 1.137 | CONFORME | Medición
14 | 5 | PA25517749 | 0 | -1.1298 | -0.88831 | -0.45881 | CONFORME | Medición
15 | 5 | PA25517750 | 0 | 0.059465 | 0.80846 | 1.137 | CONFORME | Medición
16 | 5 | PA25517751 | 0 | 0.1494 | 0.1996 | 1.0967 | CONFORME | Medición
17 | 5 | PA25517752 | 0 | -0.069721 | -0.1996 | 0.69791 | CONFORME | Medición
18 | 5 | PA25517753 | 0 | -0.50797 | -0.0998 | 0.69791 | CONFORME | Medición
19 | 5 | PA25517754 | 0 | -0.1494 | -0.5988 | 0.2991 | CONFORME | Medición
20 | 5 | PA25517755 | 0 | 0.43825 | 0.1996 | 1.6949 | CONFORME | Medición
21 | 5 | PA25517756 | 0 | 0.0099602 | -0.2994 | -0.099701 | CONFORME | Medición
22 | 5 | PA25517757 | 0 | -1.4442 | -0.998 | 0.099701 | CONFORME | Medición
23 | 5 | PA25517758 | 0 | -0.84661 | -0.2994 | 0.69791 | CONFORME | Medición
24 | 5 | PA25517759 | 0 | -0.1992 | -0.7984 | -0.099701 | CONFORME | Medición
25 | 5 | PA25517760 | 0 | 0.13944 | -0.2994 | 0.89731 | CONFORME | Medición
26 | 5 | PA25517761 | 0 | 0.039841 | 0.1996 | 1.2961 | CONFORME | Medición
27 | 5 | PA25517762 | 0 | 0.40837 | -0.0998 | 1.0967 | CONFORME | Medición
28 | 5 | PA25517763 | 0 | -0.049801 | 0.1996 | 0.89731 | CONFORME | Medición
29 | 5 | PA25517764 | 0 | -0.249 | 0.499 | 1.0967 | CONFORME | Medición
30 | 5 | PA25517765 | 0 | 0.039841 | 0.1996 | 1.4955 | CONFORME | Medición
31 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
32 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
33 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
34 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
35 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
36 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
37 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
38 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
39 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
40 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
41 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
42 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
43 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
44 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
45 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
46 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
47 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
48 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
49 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
50 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
51 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
52 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
53 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
54 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
55 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
56 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
57 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
58 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
59 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
60 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
61 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
62 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
63 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
64 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
65 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
66 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
67 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
68 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
69 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
70 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
71 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
72 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
73 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
74 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
75 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
76 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
77 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
78 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
79 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
80 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
81 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
82 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
83 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
84 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
85 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
86 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
87 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
88 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
89 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
90 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
91 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
92 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
93 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
94 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
95 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
96 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
97 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
98 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
99 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
100 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
101 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
102 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
103 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
104 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
105 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
106 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
107 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
108 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
109 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
110 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
111 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
112 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
113 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
114 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
115 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
116 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
117 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
118 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
119 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
120 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
121 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 5 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición