## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517181 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 113.47 | 213.25 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.83482 | 20.4 | 1.507 | 1.504 | 224.8 | 234.72 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | -0.94858 | 20.3 | 1.509 | 1.504 | 234.72 | 239.76 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 0.5386 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517182 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 113.65 | 214.16 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.10932 | 20.4 | 1.507 | 1.504 | 225.78 | 235.75 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | -0.44933 | 20.3 | 1.509 | 1.504 | 235.75 | 240.81 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 0.93756 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517183 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 112.12 | 211 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -1.7293 | 20.4 | 1.507 | 1.504 | 223.25 | 233.3 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | 0.34948 | 20.3 | 1.509 | 1.504 | 233.3 | 238.37 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 1.137 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517184 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 109.9 | 210.24 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.27827 | 20.4 | 1.507 | 1.504 | 221.86 | 231.8 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | -0.74888 | 20.3 | 1.509 | 1.504 | 231.8 | 236.81 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | -0.059844 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
5 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517185 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 112.81 | 213.23 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.19877 | 20.4 | 1.507 | 1.504 | 224.85 | 234.9 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | 0.34948 | 20.3 | 1.509 | 1.504 | 234.9 | 240.01 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 1.935 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517186 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 110.77 | 211.25 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.13914 | 20.4 | 1.507 | 1.504 | 222.87 | 232.85 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | -0.34948 | 20.3 | 1.509 | 1.504 | 232.85 | 237.94 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 1.536 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517187 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 110.7 | 210.4 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.91433 | 20.4 | 1.507 | 1.504 | 221.94 | 231.89 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | -0.64903 | 20.3 | 1.509 | 1.504 | 231.89 | 236.91 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 0.13964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517188 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 112.1 | 212.35 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.36772 | 20.4 | 1.507 | 1.504 | 223.97 | 234 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | 0.14978 | 20.3 | 1.509 | 1.504 | 234 | 239.1 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 1.7355 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
9 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517189 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 111.1 | 211.75 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | 0.029815 | 20.4 | 1.507 | 1.504 | 223.4 | 233.45 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | 0.34948 | 20.3 | 1.509 | 1.504 | 233.45 | 238.5 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 0.73808 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517190 | 25.6 | 0 | 20.9 | 7.883 | 0.554 | 110.67 | 210.99 | 100.62 | 2,29,560 | 2 | 29,560 | 0.041544 | 2421.9845 | -0.29815 | 20.4 | 1.507 | 1.504 | 222.65 | 232.75 | 10.015 | 22,10,090 | 22 | 10,090 | 0.36947 | 27.1064 | 0.84873 | 20.3 | 1.509 | 1.504 | 232.75 | 237.84 | 5.013 | 17,59,050 | 17 | 59,050 | 0.29974 | 16.7247 | 1.536 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517191 | 25.6 | 0 | 21 | 7.892 | 0.557 | 105.07 | 205.25 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.51639 | 20.9 | 1.51 | 1.509 | 210.5 | 220.5 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.10988 | 20.8 | 1.51 | 1.506 | 220.5 | 225.54 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517192 | 25.6 | 0 | 21 | 7.892 | 0.557 | 108.3 | 207.9 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -1.0924 | 20.9 | 1.51 | 1.509 | 213.15 | 223.1 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.60933 | 20.8 | 1.51 | 1.506 | 223.1 | 228.16 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
13 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517193 | 25.6 | 0 | 21 | 7.892 | 0.557 | 107.35 | 207.2 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.84409 | 20.9 | 1.51 | 1.509 | 212.45 | 222.42 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.40955 | 20.8 | 1.51 | 1.506 | 222.42 | 227.45 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517194 | 25.6 | 0 | 21 | 7.892 | 0.557 | 104.58 | 204.4 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.87388 | 20.9 | 1.51 | 1.509 | 209.69 | 219.65 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.50944 | 20.8 | 1.51 | 1.506 | 219.65 | 224.75 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 1.7152 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517195 | 25.6 | 0 | 21 | 7.892 | 0.557 | 107.15 | 207.5 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.34757 | 20.9 | 1.51 | 1.509 | 212.78 | 222.75 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.40955 | 20.8 | 1.51 | 1.506 | 222.75 | 227.75 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | -0.27922 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
16 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517196 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110.62 | 211.43 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | 0.10924 | 20.9 | 1.51 | 1.509 | 216.72 | 226.75 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | 0.18979 | 20.8 | 1.51 | 1.506 | 226.75 | 231.89 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 2.513 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
17 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517197 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110.21 | 210.82 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.089374 | 20.9 | 1.51 | 1.509 | 216.07 | 226.05 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.30966 | 20.8 | 1.51 | 1.506 | 226.05 | 231.09 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
18 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517198 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110.4 | 210.95 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.14896 | 20.9 | 1.51 | 1.509 | 216.25 | 226.3 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | 0.38957 | 20.8 | 1.51 | 1.506 | 226.3 | 231.35 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517199 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110 | 210.37 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.32771 | 20.9 | 1.51 | 1.509 | 215.64 | 225.62 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.30966 | 20.8 | 1.51 | 1.506 | 225.62 | 230.65 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.31911 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
20 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517200 | 25.6 | 0 | 21 | 7.892 | 0.557 | 111.41 | 211.44 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.66534 | 20.9 | 1.51 | 1.509 | 217.2 | 227.2 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.10988 | 20.8 | 1.51 | 1.506 | 227.2 | 232.25 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.71799 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
21 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517201 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110.95 | 211.87 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | 0.21847 | 20.9 | 1.51 | 1.509 | 217.15 | 227.15 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.10988 | 20.8 | 1.51 | 1.506 | 227.15 | 232.21 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.91743 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517202 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110.78 | 211.17 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.30785 | 20.9 | 1.51 | 1.509 | 216.45 | 226.46 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.009989 | 20.8 | 1.51 | 1.506 | 226.46 | 231.4 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | -1.4759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
23 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517203 | 25.6 | 0 | 21 | 7.892 | 0.557 | 109.72 | 210.15 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.26812 | 20.9 | 1.51 | 1.509 | 215.33 | 225.32 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.20977 | 20.8 | 1.51 | 1.506 | 225.32 | 230.4 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 1.3163 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517204 | 25.6 | 0 | 21 | 7.892 | 0.557 | 109.9 | 209.9 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -0.69513 | 20.9 | 1.51 | 1.509 | 215.2 | 225.15 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | -0.60933 | 20.8 | 1.51 | 1.506 | 225.15 | 230.22 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 1.1169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517205 | 25.6 | 0 | 21 | 7.892 | 0.557 | 110.4 | 211.1 | 100.7 | 2,29,640 | 2 | 29,640 | 0.041567 | 2422.6143 | -1.4112e-14 | 20.9 | 1.51 | 1.509 | 216.25 | 226.28 | 10.011 | 22,11,640 | 22 | 11,640 | 0.3699 | 27.0641 | 0.18979 | 20.8 | 1.51 | 1.506 | 226.28 | 231.32 | 5.014 | 17,59,710 | 17 | 59,710 | 0.29992 | 16.7178 | 0.51855 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517206 | 25.6 | 0 | 21 | 7.897 | 0.557 | 113.46 | 213.16 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.99305 | 21.2 | 1.509 | 1.506 | 218.95 | 228.9 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.64903 | 21.1 | 1.509 | 1.504 | 228.9 | 233.92 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
27 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517207 | 25.6 | 0 | 21 | 7.897 | 0.557 | 114.1 | 214.11 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.6852 | 21.2 | 1.509 | 1.506 | 219.87 | 229.82 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.64903 | 21.1 | 1.509 | 1.504 | 229.82 | 234.85 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517208 | 25.6 | 0 | 21 | 7.897 | 0.557 | 117.12 | 218.15 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | 0.32771 | 21.2 | 1.509 | 1.506 | 223.95 | 233.95 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.14978 | 21.1 | 1.509 | 1.504 | 233.95 | 239 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517209 | 25.6 | 0 | 21 | 7.897 | 0.557 | 117.65 | 218.6 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | 0.24826 | 21.2 | 1.509 | 1.506 | 224.43 | 234.48 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | 0.34948 | 21.1 | 1.509 | 1.504 | 234.48 | 239.52 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
30 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517210 | 25.6 | 0 | 21 | 7.897 | 0.557 | 116.4 | 216.5 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.59583 | 21.2 | 1.509 | 1.506 | 222.3 | 232.27 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.44933 | 21.1 | 1.509 | 1.504 | 232.27 | 237.34 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517211 | 25.6 | 0 | 21 | 7.897 | 0.557 | 116.2 | 216.7 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.19861 | 21.2 | 1.509 | 1.506 | 222.51 | 232.54 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | 0.14978 | 21.1 | 1.509 | 1.504 | 232.54 | 237.62 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517212 | 25.6 | 0 | 21 | 7.897 | 0.557 | 112.06 | 211.35 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -1.4002 | 21.2 | 1.509 | 1.506 | 217.1 | 226.97 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -1.4478 | 21.1 | 1.509 | 1.504 | 226.97 | 231.95 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
33 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517213 | 25.6 | 0 | 21 | 7.897 | 0.557 | 116.65 | 216.85 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.49652 | 21.2 | 1.509 | 1.506 | 222.69 | 232.65 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.54918 | 21.1 | 1.509 | 1.504 | 232.65 | 237.72 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
34 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517214 | 25.6 | 0 | 21 | 7.897 | 0.557 | 117.25 | 217.6 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.34757 | 21.2 | 1.509 | 1.506 | 223.4 | 233.4 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.14978 | 21.1 | 1.509 | 1.504 | 233.4 | 238.46 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
35 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517215 | 25.6 | 0 | 21 | 7.897 | 0.557 | 115.65 | 216.2 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.14896 | 21.2 | 1.509 | 1.506 | 222 | 231.97 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.44933 | 21.1 | 1.509 | 1.504 | 231.97 | 237.03 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
36 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517216 | 25.6 | 0 | 21 | 7.897 | 0.557 | 116 | 216.31 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.38729 | 21.2 | 1.509 | 1.506 | 222.1 | 232.06 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.54918 | 21.1 | 1.509 | 1.504 | 232.06 | 237.13 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517217 | 25.6 | 0 | 21 | 7.897 | 0.557 | 115.2 | 215.7 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.19861 | 21.2 | 1.509 | 1.506 | 221.5 | 231.49 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.24963 | 21.1 | 1.509 | 1.504 | 231.49 | 236.52 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517218 | 25.6 | 0 | 21 | 7.897 | 0.557 | 114.2 | 215.25 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | 0.34757 | 21.2 | 1.509 | 1.506 | 221.06 | 231.04 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | -0.34948 | 21.1 | 1.509 | 1.504 | 231.04 | 236.06 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
39 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517219 | 25.6 | 0 | 21 | 7.897 | 0.557 | 117.66 | 218.3 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | -0.059583 | 21.2 | 1.509 | 1.506 | 224.1 | 234.15 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | 0.34948 | 21.1 | 1.509 | 1.504 | 234.15 | 239.21 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
40 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517220 | 25.6 | 0 | 21 | 7.897 | 0.557 | 113.37 | 214.15 | 100.7 | 2,29,810 | 2 | 29,810 | 0.041614 | 2419.8652 | 0.079444 | 21.2 | 1.509 | 1.506 | 219.95 | 230.02 | 10.015 | 22,21,390 | 22 | 21,390 | 0.37261 | 26.8781 | 0.54918 | 21.1 | 1.509 | 1.504 | 230.02 | 235.14 | 5.015 | 18,01,670 | 18 | 01,670 | 0.30046 | 16.6909 | 2.0937 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
41 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517221 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 108.04 | 208.16 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.47714 | 21.2 | 1.509 | 1.504 | 213.8 | 223.8 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.14978 | 21.2 | 1.508 | 1.507 | 223.8 | 228.87 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
42 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517222 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 103.25 | 202.85 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.99404 | 21.2 | 1.509 | 1.504 | 208.46 | 218.45 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.24963 | 21.2 | 1.508 | 1.507 | 218.45 | 223.49 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
43 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517223 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 103 | 202.75 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.84493 | 21.2 | 1.509 | 1.504 | 208.4 | 218.37 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.44933 | 21.2 | 1.508 | 1.507 | 218.37 | 223.32 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | -1.3158 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517224 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 106.45 | 206.46 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.58648 | 21.2 | 1.509 | 1.504 | 212.1 | 222.1 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.14978 | 21.2 | 1.508 | 1.507 | 222.1 | 227.15 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
45 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517225 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 105.64 | 205.56 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.67594 | 21.2 | 1.509 | 1.504 | 211.2 | 221.15 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.64903 | 21.2 | 1.508 | 1.507 | 221.15 | 226.16 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | -0.11962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
46 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517226 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 107.2 | 207.9 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | 0.099404 | 21.2 | 1.509 | 1.504 | 213.57 | 223.62 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | 0.34948 | 21.2 | 1.508 | 1.507 | 223.62 | 228.7 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 1.2759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
47 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517227 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 106.43 | 206.84 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.18887 | 21.2 | 1.509 | 1.504 | 212.51 | 222.51 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.14978 | 21.2 | 1.508 | 1.507 | 222.51 | 227.56 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 0.67783 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
48 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517228 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 107.05 | 207.25 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.39761 | 21.2 | 1.509 | 1.504 | 212.89 | 222.91 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | 0.049925 | 21.2 | 1.508 | 1.507 | 222.91 | 228.02 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 1.874 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517229 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 108.7 | 208.84 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.45726 | 21.2 | 1.509 | 1.504 | 214.5 | 224.52 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | 0.049925 | 21.2 | 1.508 | 1.507 | 224.52 | 229.6 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 1.2759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
50 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517230 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 104.71 | 205.2 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.10934 | 21.2 | 1.509 | 1.504 | 210.82 | 220.85 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | 0.14978 | 21.2 | 1.508 | 1.507 | 220.85 | 225.93 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 1.2759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
51 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517231 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 108.2 | 208.15 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.64612 | 21.2 | 1.509 | 1.504 | 213.75 | 223.65 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -1.1483 | 21.2 | 1.508 | 1.507 | 223.65 | 228.9 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 4.6651 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
52 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517232 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 106.2 | 205.68 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -1.1133 | 21.2 | 1.509 | 1.504 | 211.3 | 221.25 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.64903 | 21.2 | 1.508 | 1.507 | 221.25 | 226.26 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | -0.11962 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
53 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517233 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 108.94 | 208.73 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.80517 | 21.2 | 1.509 | 1.504 | 214.35 | 224.35 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | -0.14978 | 21.2 | 1.508 | 1.507 | 224.35 | 229.37 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 0.079745 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
54 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517234 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 107.35 | 207.6 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | -0.34791 | 21.2 | 1.509 | 1.504 | 213.27 | 223.3 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | 0.14978 | 21.2 | 1.508 | 1.507 | 223.3 | 228.37 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 1.0766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517235 | 25.6 | 0 | 21.1 | 7.893 | 0.558 | 106 | 206.61 | 100.6 | 2,29,460 | 2 | 29,460 | 0.041517 | 2423.1232 | 0.0099404 | 21.2 | 1.509 | 1.504 | 212.3 | 222.36 | 10.015 | 22,08,570 | 22 | 08,570 | 0.36905 | 27.1374 | 0.44933 | 21.2 | 1.508 | 1.507 | 222.36 | 227.4 | 5.016 | 17,57,310 | 17 | 57,310 | 0.29925 | 16.7617 | 0.47847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
56 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517236 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 106.25 | 206.15 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.69583 | 21.2 | 1.508 | 1.504 | 214.03 | 224 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -0.3996 | 21 | 1.509 | 1.507 | 224 | 229.02 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
57 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517237 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 103.8 | 203.12 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -1.2724 | 21.2 | 1.508 | 1.504 | 210.96 | 220.85 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -1.1988 | 21 | 1.509 | 1.507 | 220.85 | 225.85 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
58 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517238 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 103.62 | 203.74 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.47714 | 21.2 | 1.508 | 1.504 | 211.65 | 221.7 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 0.3996 | 21 | 1.509 | 1.507 | 221.7 | 226.78 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
59 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517239 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 108.6 | 208.55 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.64612 | 21.2 | 1.508 | 1.504 | 216.39 | 226.42 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 0.1998 | 21 | 1.509 | 1.507 | 226.42 | 231.46 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
60 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517240 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 107.7 | 207.95 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.34791 | 21.2 | 1.508 | 1.504 | 215.85 | 225.71 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -1.4985 | 21 | 1.509 | 1.507 | 225.71 | 230.65 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | -1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
61 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517241 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 107.4 | 207.56 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.43738 | 21.2 | 1.508 | 1.504 | 215.45 | 225.45 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -0.0999 | 21 | 1.509 | 1.507 | 225.45 | 230.46 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
62 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517242 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 108.46 | 208.5 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.55666 | 21.2 | 1.508 | 1.504 | 216.36 | 226.35 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -0.1998 | 21 | 1.509 | 1.507 | 226.35 | 231.37 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
63 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517243 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 150.6 | 250.5 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.69583 | 21.2 | 1.508 | 1.504 | 258.3 | 268.32 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 0.0999 | 21 | 1.509 | 1.507 | 268.32 | 273.37 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
64 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517244 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 108.64 | 208.63 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.60636 | 21.2 | 1.508 | 1.504 | 216.5 | 226.49 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -0.1998 | 21 | 1.509 | 1.507 | 226.49 | 231.55 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
65 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517245 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 105.22 | 205.8 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.019881 | 21.2 | 1.508 | 1.504 | 213.69 | 223.64 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -0.5994 | 21 | 1.509 | 1.507 | 223.64 | 228.6 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | -1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
66 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517246 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 109.12 | 209.35 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.36779 | 21.2 | 1.508 | 1.504 | 217.2 | 227.24 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 0.2997 | 21 | 1.509 | 1.507 | 227.24 | 232.3 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
67 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517247 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 104.06 | 204.05 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.60636 | 21.2 | 1.508 | 1.504 | 211.89 | 221.86 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | -0.3996 | 21 | 1.509 | 1.507 | 221.86 | 226.91 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
68 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517248 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 105.4 | 205.08 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.91451 | 21.2 | 1.508 | 1.504 | 213.54 | 223.55 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 1.952e-13 | 21 | 1.509 | 1.507 | 223.55 | 228.6 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
69 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517249 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 111.75 | 211.6 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | -0.74553 | 21.2 | 1.508 | 1.504 | 219.43 | 229.55 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 1.0989 | 21 | 1.509 | 1.507 | 229.55 | 234.65 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
70 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517250 | 25.6 | 0 | 21.3 | 7.889 | 0.556 | 110.45 | 211.3 | 100.6 | 2,30,980 | 2 | 30,980 | 0.041939 | 2398.7283 | 0.24851 | 21.2 | 1.508 | 1.504 | 219.12 | 229.3 | 10.01 | 22,12,050 | 22 | 12,050 | 0.37001 | 27.053 | 1.6983 | 21 | 1.509 | 1.507 | 229.3 | 234.43 | 5.015 | 17,58,370 | 17 | 58,370 | 0.29955 | 16.7419 | 2.2931 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
71 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517251 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 146.35 | 246.52 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.625 | 21.3 | 1.509 | 1.504 | 253.64 | 263.66 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.089901 | 21.2 | 1.506 | 1.503 | 263.66 | 268.74 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
72 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517252 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 102.8 | 202.95 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.64484 | 21.3 | 1.509 | 1.504 | 210.05 | 220.05 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -0.10988 | 21.2 | 1.506 | 1.503 | 220.05 | 225.06 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
73 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517253 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 107.7 | 207.55 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.94246 | 21.3 | 1.509 | 1.504 | 214.64 | 224.57 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -0.80911 | 21.2 | 1.506 | 1.503 | 224.57 | 229.6 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
74 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517254 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 105.8 | 206.12 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.47619 | 21.3 | 1.509 | 1.504 | 213.25 | 223.26 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -0.009989 | 21.2 | 1.506 | 1.503 | 223.26 | 228.32 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
75 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517255 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 106.5 | 206.25 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -1.0417 | 21.3 | 1.509 | 1.504 | 213.17 | 223.2 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.18979 | 21.2 | 1.506 | 1.503 | 223.2 | 228.25 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
76 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517256 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 105.9 | 205.82 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.87302 | 21.3 | 1.509 | 1.504 | 212.94 | 222.9 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -0.50944 | 21.2 | 1.506 | 1.503 | 222.9 | 227.95 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
77 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517257 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 106.5 | 206.65 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.64484 | 21.3 | 1.509 | 1.504 | 213.75 | 223.84 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.78913 | 21.2 | 1.506 | 1.503 | 223.84 | 228.96 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 1.992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
78 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517258 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 106.86 | 206.05 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -1.5972 | 21.3 | 1.509 | 1.504 | 213.15 | 223.17 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.089901 | 21.2 | 1.506 | 1.503 | 223.17 | 228.27 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 1.5936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
79 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517259 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 149.12 | 250.25 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | 0.32738 | 21.3 | 1.509 | 1.504 | 257.37 | 267.45 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.68924 | 21.2 | 1.506 | 1.503 | 267.45 | 272.52 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.99602 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
80 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517260 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 109.85 | 210.6 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.049603 | 21.3 | 1.509 | 1.504 | 217.65 | 227.7 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.38957 | 21.2 | 1.506 | 1.503 | 227.7 | 232.85 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 2.5896 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
81 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517261 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 108.8 | 208.7 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.89286 | 21.3 | 1.509 | 1.504 | 215.72 | 225.75 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.18979 | 21.2 | 1.506 | 1.503 | 225.75 | 230.86 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 1.7928 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
82 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517262 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 109.31 | 209.67 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.43651 | 21.3 | 1.509 | 1.504 | 216.72 | 226.7 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -0.30966 | 21.2 | 1.506 | 1.503 | 226.7 | 231.66 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
83 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517263 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 110.13 | 210.87 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.059524 | 21.3 | 1.509 | 1.504 | 217.94 | 228 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.48946 | 21.2 | 1.506 | 1.503 | 228 | 233.05 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
84 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517264 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 145.96 | 245.95 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.80357 | 21.3 | 1.509 | 1.504 | 252.9 | 262.9 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -0.10988 | 21.2 | 1.506 | 1.503 | 262.9 | 267.95 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
85 | 2025-09-19 | 2025-09-19 | 6 | 15 |  | PA25517265 | 25.6 | 0 | 21.4 | 7.897 | 0.564 | 185.3 | 285.35 | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -0.74405 | 21.3 | 1.509 | 1.504 | 292.37 | 302.44 | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | 0.58935 | 21.2 | 1.506 | 1.503 | 302.44 | 307.52 | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | 1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
86 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
87 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
88 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
89 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
90 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
91 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
92 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
93 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
94 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
95 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
96 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
97 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
98 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
99 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
100 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
101 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
102 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
103 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
104 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
105 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
106 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
107 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
108 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
109 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
110 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
111 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
112 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
113 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
114 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
115 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
116 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
117 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
118 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
119 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
120 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
121 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
122 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
123 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
124 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
125 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
126 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
128 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
129 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
130 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
131 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
132 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
133 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
134 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
135 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
136 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
137 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
138 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
139 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
140 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
141 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
142 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
143 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
144 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
145 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
146 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
147 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
148 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
150 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
151 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
152 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
153 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
154 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
155 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
156 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
157 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
158 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
159 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
160 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
161 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
162 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
163 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
164 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
165 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
166 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
167 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
168 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
169 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
170 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
171 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
172 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
173 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
174 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
175 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
176 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
177 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
178 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-19 | 2025-09-19 | 6 | 15 |  |  | 25.6 | 0 | 21.4 | 7.897 | 0.564 |  |  | 100.8 | 2,29,840 | 2 | 29,840 | 0.041622 | 2421.7832 | -100 | 21.3 | 1.509 | 1.504 |  |  | 10.011 | 22,14,040 | 22 | 14,040 | 0.37057 | 27.0154 | -100 | 21.2 | 1.506 | 1.503 | 0 |  | 5.02 | 17,59,600 | 17 | 59,600 | 0.29989 | 16.7395 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 6 | PA25517181 | 0 | -0.83482 | -0.94858 | 0.5386 | CONFORME | Medición
2 | 6 | PA25517182 | 0 | -0.10932 | -0.44933 | 0.93756 | CONFORME | Medición
3 | 6 | PA25517183 | 0 | -1.7293 | 0.34948 | 1.137 | CONFORME | Medición
4 | 6 | PA25517184 | 0 | -0.27827 | -0.74888 | -0.059844 | CONFORME | Medición
5 | 6 | PA25517185 | 0 | -0.19877 | 0.34948 | 1.935 | CONFORME | Medición
6 | 6 | PA25517186 | 0 | -0.13914 | -0.34948 | 1.536 | CONFORME | Medición
7 | 6 | PA25517187 | 0 | -0.91433 | -0.64903 | 0.13964 | CONFORME | Medición
8 | 6 | PA25517188 | 0 | -0.36772 | 0.14978 | 1.7355 | CONFORME | Medición
9 | 6 | PA25517189 | 0 | 0.029815 | 0.34948 | 0.73808 | CONFORME | Medición
10 | 6 | PA25517190 | 0 | -0.29815 | 0.84873 | 1.536 | CONFORME | Medición
11 | 6 | PA25517191 | 0 | -0.51639 | -0.10988 | 0.51855 | CONFORME | Medición
12 | 6 | PA25517192 | 0 | -1.0924 | -0.60933 | 0.91743 | CONFORME | Medición
13 | 6 | PA25517193 | 0 | -0.84409 | -0.40955 | 0.31911 | CONFORME | Medición
14 | 6 | PA25517194 | 0 | -0.87388 | -0.50944 | 1.7152 | CONFORME | Medición
15 | 6 | PA25517195 | 0 | -0.34757 | -0.40955 | -0.27922 | CONFORME | Medición
16 | 6 | PA25517196 | 0 | 0.10924 | 0.18979 | 2.513 | CONFORME | Medición
17 | 6 | PA25517197 | 0 | -0.089374 | -0.30966 | 0.51855 | CONFORME | Medición
18 | 6 | PA25517198 | 0 | -0.14896 | 0.38957 | 0.71799 | CONFORME | Medición
19 | 6 | PA25517199 | 0 | -0.32771 | -0.30966 | 0.31911 | CONFORME | Medición
20 | 6 | PA25517200 | 0 | -0.66534 | -0.10988 | 0.71799 | CONFORME | Medición
21 | 6 | PA25517201 | 0 | 0.21847 | -0.10988 | 0.91743 | CONFORME | Medición
22 | 6 | PA25517202 | 0 | -0.30785 | -0.009989 | -1.4759 | CONFORME | Medición
23 | 6 | PA25517203 | 0 | -0.26812 | -0.20977 | 1.3163 | CONFORME | Medición
24 | 6 | PA25517204 | 0 | -0.69513 | -0.60933 | 1.1169 | CONFORME | Medición
25 | 6 | PA25517205 | 0 | -1.4112e-14 | 0.18979 | 0.51855 | CONFORME | Medición
26 | 6 | PA25517206 | 0 | -0.99305 | -0.64903 | 0.099701 | CONFORME | Medición
27 | 6 | PA25517207 | 0 | -0.6852 | -0.64903 | 0.2991 | CONFORME | Medición
28 | 6 | PA25517208 | 0 | 0.32771 | -0.14978 | 0.69791 | CONFORME | Medición
29 | 6 | PA25517209 | 0 | 0.24826 | 0.34948 | 0.4985 | CONFORME | Medición
30 | 6 | PA25517210 | 0 | -0.59583 | -0.44933 | 1.0967 | CONFORME | Medición
31 | 6 | PA25517211 | 0 | -0.19861 | 0.14978 | 1.2961 | CONFORME | Medición
32 | 6 | PA25517212 | 0 | -1.4002 | -1.4478 | -0.69791 | CONFORME | Medición
33 | 6 | PA25517213 | 0 | -0.49652 | -0.54918 | 1.0967 | CONFORME | Medición
34 | 6 | PA25517214 | 0 | -0.34757 | -0.14978 | 0.89731 | CONFORME | Medición
35 | 6 | PA25517215 | 0 | -0.14896 | -0.44933 | 0.89731 | CONFORME | Medición
36 | 6 | PA25517216 | 0 | -0.38729 | -0.54918 | 1.0967 | CONFORME | Medición
37 | 6 | PA25517217 | 0 | -0.19861 | -0.24963 | 0.2991 | CONFORME | Medición
38 | 6 | PA25517218 | 0 | 0.34757 | -0.34948 | 0.099701 | CONFORME | Medición
39 | 6 | PA25517219 | 0 | -0.059583 | 0.34948 | 0.89731 | CONFORME | Medición
40 | 6 | PA25517220 | 0 | 0.079444 | 0.54918 | 2.0937 | CONFORME | Medición
41 | 6 | PA25517221 | 0 | -0.47714 | -0.14978 | 1.0766 | CONFORME | Medición
42 | 6 | PA25517222 | 0 | -0.99404 | -0.24963 | 0.47847 | CONFORME | Medición
43 | 6 | PA25517223 | 0 | -0.84493 | -0.44933 | -1.3158 | CONFORME | Medición
44 | 6 | PA25517224 | 0 | -0.58648 | -0.14978 | 0.67783 | CONFORME | Medición
45 | 6 | PA25517225 | 0 | -0.67594 | -0.64903 | -0.11962 | CONFORME | Medición
46 | 6 | PA25517226 | 0 | 0.099404 | 0.34948 | 1.2759 | CONFORME | Medición
47 | 6 | PA25517227 | 0 | -0.18887 | -0.14978 | 0.67783 | CONFORME | Medición
48 | 6 | PA25517228 | 0 | -0.39761 | 0.049925 | 1.874 | CONFORME | Medición
49 | 6 | PA25517229 | 0 | -0.45726 | 0.049925 | 1.2759 | CONFORME | Medición
50 | 6 | PA25517230 | 0 | -0.10934 | 0.14978 | 1.2759 | CONFORME | Medición
51 | 6 | PA25517231 | 0 | -0.64612 | -1.1483 | 4.6651 | CONFORME | Medición
52 | 6 | PA25517232 | 0 | -1.1133 | -0.64903 | -0.11962 | CONFORME | Medición
53 | 6 | PA25517233 | 0 | -0.80517 | -0.14978 | 0.079745 | CONFORME | Medición
54 | 6 | PA25517234 | 0 | -0.34791 | 0.14978 | 1.0766 | CONFORME | Medición
55 | 6 | PA25517235 | 0 | 0.0099404 | 0.44933 | 0.47847 | CONFORME | Medición
56 | 6 | PA25517236 | 0 | -0.69583 | -0.3996 | 0.099701 | CONFORME | Medición
57 | 6 | PA25517237 | 0 | -1.2724 | -1.1988 | -0.2991 | CONFORME | Medición
58 | 6 | PA25517238 | 0 | -0.47714 | 0.3996 | 1.2961 | CONFORME | Medición
59 | 6 | PA25517239 | 0 | -0.64612 | 0.1998 | 0.4985 | CONFORME | Medición
60 | 6 | PA25517240 | 0 | -0.34791 | -1.4985 | -1.4955 | CONFORME | Medición
61 | 6 | PA25517241 | 0 | -0.43738 | -0.0999 | -0.099701 | CONFORME | Medición
62 | 6 | PA25517242 | 0 | -0.55666 | -0.1998 | 0.099701 | CONFORME | Medición
63 | 6 | PA25517243 | 0 | -0.69583 | 0.0999 | 0.69791 | CONFORME | Medición
64 | 6 | PA25517244 | 0 | -0.60636 | -0.1998 | 0.89731 | CONFORME | Medición
65 | 6 | PA25517245 | 0 | -0.019881 | -0.5994 | -1.0967 | CONFORME | Medición
66 | 6 | PA25517246 | 0 | -0.36779 | 0.2997 | 0.89731 | CONFORME | Medición
67 | 6 | PA25517247 | 0 | -0.60636 | -0.3996 | 0.69791 | CONFORME | Medición
68 | 6 | PA25517248 | 0 | -0.91451 | 1.952e-13 | 0.69791 | CONFORME | Medición
69 | 6 | PA25517249 | 0 | -0.74553 | 1.0989 | 1.6949 | CONFORME | Medición
70 | 6 | PA25517250 | 0 | 0.24851 | 1.6983 | 2.2931 | CONFORME | Medición
71 | 6 | PA25517251 | 0 | -0.625 | 0.089901 | 1.1952 | CONFORME | Medición
72 | 6 | PA25517252 | 0 | -0.64484 | -0.10988 | -0.1992 | CONFORME | Medición
73 | 6 | PA25517253 | 0 | -0.94246 | -0.80911 | 0.1992 | CONFORME | Medición
74 | 6 | PA25517254 | 0 | -0.47619 | -0.009989 | 0.79681 | CONFORME | Medición
75 | 6 | PA25517255 | 0 | -1.0417 | 0.18979 | 0.59761 | CONFORME | Medición
76 | 6 | PA25517256 | 0 | -0.87302 | -0.50944 | 0.59761 | CONFORME | Medición
77 | 6 | PA25517257 | 0 | -0.64484 | 0.78913 | 1.992 | CONFORME | Medición
78 | 6 | PA25517258 | 0 | -1.5972 | 0.089901 | 1.5936 | CONFORME | Medición
79 | 6 | PA25517259 | 0 | 0.32738 | 0.68924 | 0.99602 | CONFORME | Medición
80 | 6 | PA25517260 | 0 | -0.049603 | 0.38957 | 2.5896 | CONFORME | Medición
81 | 6 | PA25517261 | 0 | -0.89286 | 0.18979 | 1.7928 | CONFORME | Medición
82 | 6 | PA25517262 | 0 | -0.43651 | -0.30966 | -1.1952 | CONFORME | Medición
83 | 6 | PA25517263 | 0 | -0.059524 | 0.48946 | 0.59761 | CONFORME | Medición
84 | 6 | PA25517264 | 0 | -0.80357 | -0.10988 | 0.59761 | CONFORME | Medición
85 | 6 | PA25517265 | 0 | -0.74405 | 0.58935 | 1.1952 | CONFORME | Medición
86 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
87 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
88 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
89 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
90 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
91 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
92 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
93 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
94 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
95 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
96 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
97 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
98 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
99 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
100 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
101 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
102 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
103 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
104 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
105 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
106 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
107 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
108 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
109 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
110 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
111 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
112 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
113 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
114 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
115 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
116 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
117 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
118 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
119 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
120 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
121 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
122 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
123 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
124 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
125 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
126 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
127 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
128 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
129 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
130 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
131 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
132 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
133 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
134 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
135 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
136 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
137 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
138 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
139 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
140 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
141 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
142 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
143 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
144 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
145 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
146 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
147 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
148 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
149 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
150 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
151 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
152 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
153 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
154 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
155 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
156 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
157 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
158 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
159 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
160 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
161 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
162 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
163 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
164 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
165 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
166 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
167 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
168 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
169 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
170 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
171 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
172 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
173 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
174 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
175 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
176 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
177 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
178 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 6 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición