## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517169 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 104.86 | 204.45 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.95475 | 21.4 | 1.092 | 1.084 | 212.32 | 222.31 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.34913 | 21.1 | 1.076 | 1.066 | 222.31 | 227.3 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517170 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 103.1 | 202.7 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.9448 | 21.4 | 1.092 | 1.084 | 210.51 | 220.48 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.54863 | 21.1 | 1.076 | 1.066 | 220.48 | 225.49 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 3.9002e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517171 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 106.55 | 206.51 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.58677 | 21.4 | 1.092 | 1.084 | 214.35 | 224.33 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.44888 | 21.1 | 1.076 | 1.066 | 224.33 | 229.35 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517172 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 108.16 | 208.05 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.65639 | 21.4 | 1.092 | 1.084 | 215.95 | 225.9 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.74813 | 21.1 | 1.076 | 1.066 | 225.9 | 230.95 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517173 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 105.9 | 205.72 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.72601 | 21.4 | 1.092 | 1.084 | 213.56 | 223.55 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.34913 | 21.1 | 1.076 | 1.066 | 223.55 | 228.58 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517174 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 105.95 | 206.05 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.44754 | 21.4 | 1.092 | 1.084 | 213.93 | 223.89 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.64838 | 21.1 | 1.076 | 1.066 | 223.89 | 228.9 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 3.9002e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517175 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 106.7 | 206.71 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.53705 | 21.4 | 1.092 | 1.084 | 214.6 | 224.56 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.64838 | 21.1 | 1.076 | 1.066 | 224.56 | 229.6 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517176 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 107.64 | 207.75 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.43759 | 21.4 | 1.092 | 1.084 | 215.62 | 225.56 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.84788 | 21.1 | 1.076 | 1.066 | 225.56 | 230.48 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | -1.7964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517177 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 109 | 209.36 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.18896 | 21.4 | 1.092 | 1.084 | 217.26 | 227.25 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.34913 | 21.1 | 1.076 | 1.066 | 227.25 | 232.3 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517178 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 108.96 | 209.15 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.35803 | 21.4 | 1.092 | 1.084 | 217.05 | 227.06 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.14963 | 21.1 | 1.076 | 1.066 | 227.06 | 232.12 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517179 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 107.88 | 207.66 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.76579 | 21.4 | 1.092 | 1.084 | 215.54 | 225.54 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.24938 | 21.1 | 1.076 | 1.066 | 225.54 | 230.55 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 3.9002e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25517180 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 108.3 | 208.69 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.15912 | 21.4 | 1.092 | 1.084 | 216.6 | 226.63 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | 0.049875 | 21.1 | 1.076 | 1.066 | 226.63 | 231.66 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516701 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 107.85 | 207.88 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.51716 | 21.4 | 1.092 | 1.084 | 215.8 | 225.85 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | 0.24938 | 21.1 | 1.076 | 1.066 | 225.85 | 230.9 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516702 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 105 | 204.71 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.83541 | 21.4 | 1.092 | 1.084 | 212.55 | 222.55 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.24938 | 21.1 | 1.076 | 1.066 | 222.55 | 227.58 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516703 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 106.34 | 205.84 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -1.0443 | 21.4 | 1.092 | 1.084 | 213.74 | 223.75 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | -0.14963 | 21.1 | 1.076 | 1.066 | 223.75 | 228.82 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516704 | 25.6 | 0 | 22.4 | 8.209 | 0.588 | 106.29 | 206.1 | 100.55 | 2,32,330 | 2 | 32,330 | 0.042314 | 2376.2883 | -0.73595 | 21.4 | 1.092 | 1.084 | 214 | 224.08 | 10.025 | 22,23,380 | 22 | 23,380 | 0.37316 | 26.8651 | 0.54863 | 21.1 | 1.076 | 1.066 | 224.08 | 229.15 | 5.01 | 18,28,500 | 18 | 28,500 | 0.30792 | 16.2706 | 1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516705 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 172.95 | 272.85 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.68595 | 21.3 | 1.073 | 1.062 | 278.4 | 288.36 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.40996 | 21.3 | 1.073 | 1.065 | 288.36 | 293.31 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -1.0989 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516706 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 170.69 | 269.9 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -1.3719 | 21.3 | 1.073 | 1.062 | 275.4 | 285.35 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.50995 | 21.3 | 1.073 | 1.065 | 285.35 | 290.36 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516707 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 172.61 | 271.92 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -1.2725 | 21.3 | 1.073 | 1.062 | 277.36 | 287.29 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.70993 | 21.3 | 1.073 | 1.065 | 287.29 | 292.22 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -1.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
20 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516708 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 173.05 | 272.84 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.79531 | 21.3 | 1.073 | 1.062 | 278.36 | 288.32 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.40996 | 21.3 | 1.073 | 1.065 | 288.32 | 293.3 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516709 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 173.06 | 272.64 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -1.0041 | 21.3 | 1.073 | 1.062 | 278.15 | 288.12 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.30997 | 21.3 | 1.073 | 1.065 | 288.12 | 293.15 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516710 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 172.75 | 272.38 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.95437 | 21.3 | 1.073 | 1.062 | 277.87 | 287.9 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | 0.28997 | 21.3 | 1.073 | 1.065 | 287.9 | 292.92 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516711 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 281.36 | 381.44 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.50701 | 21.3 | 1.073 | 1.062 | 386.95 | 396.92 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.30997 | 21.3 | 1.073 | 1.065 | 396.92 | 401.9 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516712 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 283.25 | 382.85 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.98419 | 21.3 | 1.073 | 1.062 | 388.36 | 398.26 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -1.0099 | 21.3 | 1.073 | 1.065 | 398.26 | 403.3 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516713 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 171.19 | 271.44 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.33801 | 21.3 | 1.073 | 1.062 | 276.96 | 286.99 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | 0.28997 | 21.3 | 1.073 | 1.065 | 286.99 | 291.95 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -0.8991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516714 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 168.24 | 266.91 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -1.9087 | 21.3 | 1.073 | 1.062 | 272.35 | 281.85 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -5.0095 | 21.3 | 1.073 | 1.065 | 281.85 | 286.29 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -11.2887 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
27 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516715 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 173.86 | 274.15 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.29824 | 21.3 | 1.073 | 1.062 | 279.71 | 289.71 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.009999 | 21.3 | 1.073 | 1.065 | 289.71 | 294.71 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
28 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516716 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 171.25 | 271.34 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.49707 | 21.3 | 1.073 | 1.062 | 276.86 | 286.85 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.10999 | 21.3 | 1.073 | 1.065 | 286.85 | 291.89 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.6993 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
29 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516717 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 175.36 | 275.73 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.21871 | 21.3 | 1.073 | 1.062 | 281.26 | 291.25 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.10999 | 21.3 | 1.073 | 1.065 | 291.25 | 296.28 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.4995 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516718 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 171.66 | 271.35 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.89472 | 21.3 | 1.073 | 1.062 | 276.86 | 286.82 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.40996 | 21.3 | 1.073 | 1.065 | 286.82 | 291.81 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | -0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
31 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516719 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 171.78 | 271.84 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.52689 | 21.3 | 1.073 | 1.062 | 277.35 | 287.33 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.20998 | 21.3 | 1.073 | 1.065 | 287.33 | 292.35 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.2997 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
32 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516720 | 25.6 | 0 | 21.1 | 8.215 | 0.618 | 170.8 | 270.76 | 100.59 | 2,32,840 | 2 | 32,840 | 0.042456 | 2369.3012 | -0.6263 | 21.3 | 1.073 | 1.062 | 276.3 | 286.29 | 10.001 | 22,22,750 | 22 | 22,750 | 0.37299 | 26.8133 | -0.10999 | 21.3 | 1.073 | 1.065 | 286.29 | 291.3 | 5.005 | 18,27,050 | 18 | 27,050 | 0.30751 | 16.2757 | 0.0999 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
33 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516721 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 82.68 | 182.04 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -1.1835 | 21.5 | 1.071 | 1.061 | 188.35 | 198.35 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.39841 | 21.3 | 1.071 | 1.062 | 198.35 | 203.35 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.019996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
34 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516722 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 123.65 | 223.7 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.49727 | 21.5 | 1.071 | 1.061 | 230.06 | 240.06 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.39841 | 21.3 | 1.071 | 1.062 | 240.06 | 245.06 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.019996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
35 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516723 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 82.8 | 182.44 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.90502 | 21.5 | 1.071 | 1.061 | 188.71 | 198.56 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -1.8924 | 21.3 | 1.071 | 1.062 | 198.56 | 203.54 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.41992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
36 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516724 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 83.95 | 184.4 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.099453 | 21.5 | 1.071 | 1.061 | 190.75 | 200.8 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | 0.099602 | 21.3 | 1.071 | 1.062 | 200.8 | 205.82 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | 0.37992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
37 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516725 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 83.75 | 183.8 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.49727 | 21.5 | 1.071 | 1.061 | 190.14 | 200.19 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | 0.099602 | 21.3 | 1.071 | 1.062 | 200.19 | 205.19 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.019996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516726 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 85 | 184.9 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.64644 | 21.5 | 1.071 | 1.061 | 191.18 | 201.16 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.59761 | 21.3 | 1.071 | 1.062 | 201.16 | 206.15 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.21996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
39 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516727 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 158.49 | 258.2 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.83541 | 21.5 | 1.071 | 1.061 | 264.48 | 274.47 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.49801 | 21.3 | 1.071 | 1.062 | 274.47 | 279.46 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.21996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
40 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516728 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 81.45 | 181.11 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.88513 | 21.5 | 1.071 | 1.061 | 187.42 | 197.42 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.39841 | 21.3 | 1.071 | 1.062 | 197.42 | 202.41 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.21996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
41 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516729 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 82.96 | 182.3 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -1.2034 | 21.5 | 1.071 | 1.061 | 188.65 | 198.7 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | 0.099602 | 21.3 | 1.071 | 1.062 | 198.7 | 203.72 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | 0.37992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
42 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516730 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 82 | 181.81 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.73595 | 21.5 | 1.071 | 1.061 | 188.1 | 198.06 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.79681 | 21.3 | 1.071 | 1.062 | 198.06 | 203.01 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -1.0198 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
43 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516731 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 83.5 | 183.27 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.77573 | 21.5 | 1.071 | 1.061 | 189.58 | 199.56 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.59761 | 21.3 | 1.071 | 1.062 | 199.56 | 204.55 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.21996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516732 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 84.09 | 184.12 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.51716 | 21.5 | 1.071 | 1.061 | 190.45 | 200.47 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.1992 | 21.3 | 1.071 | 1.062 | 200.47 | 205.45 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.41992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
45 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516733 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 123 | 222.8 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.7459 | 21.5 | 1.071 | 1.061 | 229.15 | 239.16 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.2988 | 21.3 | 1.071 | 1.062 | 239.16 | 244.09 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -1.4197 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
46 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516734 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 86.2 | 186.44 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.3083 | 21.5 | 1.071 | 1.061 | 192.75 | 202.78 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.099602 | 21.3 | 1.071 | 1.062 | 202.78 | 207.78 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.019996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
48 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516735 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 85.26 | 185.48 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.32819 | 21.5 | 1.071 | 1.061 | 191.79 | 201.84 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | 0.099602 | 21.3 | 1.071 | 1.062 | 201.84 | 206.86 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | 0.37992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
49 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516736 | 25.6 | 0 | 21.5 | 8.209 | 0.561 | 84.95 | 185 | 100.55 | 2,32,090 | 2 | 32,090 | 0.042247 | 2380.0381 | -0.49727 | 21.5 | 1.071 | 1.061 | 191.31 | 201.34 | 10.04 | 22,29,940 | 22 | 29,940 | 0.37498 | 26.7745 | -0.099602 | 21.3 | 1.071 | 1.062 | 201.34 | 206.3 | 5.001 | 18,26,550 | 18 | 26,550 | 0.30738 | 16.27 | -0.81984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
50 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516737 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 88.29 | 187.9 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.9841 | 21.3 | 1.074 | 1.065 | 193.07 | 203.05 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.2997 | 21.4 | 1.075 | 1.065 | 203.05 | 208.1 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
51 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516738 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 87.8 | 187.78 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.6163 | 21.3 | 1.074 | 1.065 | 192.95 | 202.91 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.4995 | 21.4 | 1.075 | 1.065 | 202.91 | 207.94 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
52 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516739 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 85.39 | 184.86 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -1.1233 | 21.3 | 1.074 | 1.065 | 190 | 200.05 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | 0.3996 | 21.4 | 1.075 | 1.065 | 200.05 | 205.15 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
53 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516740 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 87.8 | 188.01 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.38767 | 21.3 | 1.074 | 1.065 | 193.2 | 203.16 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.4995 | 21.4 | 1.075 | 1.065 | 203.16 | 208.16 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
54 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516741 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 89.75 | 189.09 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -1.2525 | 21.3 | 1.074 | 1.065 | 194.25 | 204.21 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.4995 | 21.4 | 1.075 | 1.065 | 204.21 | 209.24 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516742 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 88.65 | 188.44 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.80517 | 21.3 | 1.074 | 1.065 | 193.55 | 203.56 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -8.8729e-14 | 21.4 | 1.075 | 1.065 | 203.56 | 208.64 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
56 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516743 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 236.6 | 336.36 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.83499 | 21.3 | 1.074 | 1.065 | 341.51 | 351.49 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.2997 | 21.4 | 1.075 | 1.065 | 351.49 | 356.54 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
57 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516744 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 86.79 | 186.8 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.58648 | 21.3 | 1.074 | 1.065 | 192 | 202.02 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | 0.0999 | 21.4 | 1.075 | 1.065 | 202.02 | 207.07 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
58 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516745 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 85.4 | 185.08 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.91451 | 21.3 | 1.074 | 1.065 | 190.22 | 200.2 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.2997 | 21.4 | 1.075 | 1.065 | 200.2 | 205.21 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
59 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516746 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 87.95 | 188.71 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | 0.15905 | 21.3 | 1.074 | 1.065 | 193.91 | 203.95 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | 0.2997 | 21.4 | 1.075 | 1.065 | 203.95 | 209 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
60 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516747 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 89.85 | 190.3 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.14911 | 21.3 | 1.074 | 1.065 | 195.48 | 205.43 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.5994 | 21.4 | 1.075 | 1.065 | 205.43 | 210.39 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
61 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516748 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 91.05 | 190.79 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.85487 | 21.3 | 1.074 | 1.065 | 195.94 | 205.86 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.8991 | 21.4 | 1.075 | 1.065 | 205.86 | 210.77 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -2.0937 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
62 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516749 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 90.94 | 190.85 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.68588 | 21.3 | 1.074 | 1.065 | 196 | 205.94 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.6993 | 21.4 | 1.075 | 1.065 | 205.94 | 210.92 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
63 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516750 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 92.7 | 193.25 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.049702 | 21.3 | 1.074 | 1.065 | 198.45 | 208.42 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.3996 | 21.4 | 1.075 | 1.065 | 208.42 | 213.39 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
64 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516751 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 91.89 | 191.85 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.63618 | 21.3 | 1.074 | 1.065 | 197 | 206.92 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.8991 | 21.4 | 1.075 | 1.065 | 206.92 | 211.85 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
65 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516752 | 25.6 | 0 | 21.3 | 8.212 | 0.595 | 92.25 | 192.2 | 100.6 | 2,32,530 | 2 | 32,530 | 0.042369 | 2374.3526 | -0.64612 | 21.3 | 1.074 | 1.065 | 197.36 | 207.35 | 10.01 | 22,26,350 | 22 | 26,350 | 0.37399 | 26.7657 | -0.1998 | 21.4 | 1.075 | 1.065 | 207.35 | 212.36 | 5.015 | 18,20,900 | 18 | 20,900 | 0.30581 | 16.3993 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
66 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516753 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 88.1 | 187.58 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -1.1624 | 20.9 | 1.074 | 1.062 | 193.18 | 203.12 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.64968 | 21 | 1.074 | 1.062 | 203.12 | 208.06 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -1.4169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
67 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516754 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.66 | 187.05 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -1.2519 | 20.9 | 1.074 | 1.062 | 192.6 | 202.55 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.54973 | 21 | 1.074 | 1.062 | 202.55 | 207.56 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.019956 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
68 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516755 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 88.01 | 187.72 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.93393 | 20.9 | 1.074 | 1.062 | 193.35 | 203.36 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | 0.049975 | 21 | 1.074 | 1.062 | 203.36 | 208.45 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | 1.5765 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
69 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516756 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 88.14 | 187.81 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.97367 | 20.9 | 1.074 | 1.062 | 193.44 | 203.38 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.64968 | 21 | 1.074 | 1.062 | 203.38 | 208.36 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.61864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
70 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516757 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 86.3 | 186.06 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.88425 | 20.9 | 1.074 | 1.062 | 191.7 | 201.59 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -1.1494 | 21 | 1.074 | 1.062 | 201.59 | 206.53 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -1.4169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
71 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516758 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 86.19 | 185.78 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -1.0532 | 20.9 | 1.074 | 1.062 | 191.39 | 201.36 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.34983 | 21 | 1.074 | 1.062 | 201.36 | 206.35 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
72 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516759 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.6 | 186.84 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -1.4009 | 20.9 | 1.074 | 1.062 | 192.54 | 202.46 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.84958 | 21 | 1.074 | 1.062 | 202.46 | 207.45 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
73 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516760 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 86.88 | 186.7 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.82464 | 20.9 | 1.074 | 1.062 | 192.3 | 202.26 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.44978 | 21 | 1.074 | 1.062 | 202.26 | 207.25 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
74 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516761 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.3 | 187.06 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.88425 | 20.9 | 1.074 | 1.062 | 192.65 | 202.56 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.94953 | 21 | 1.074 | 1.062 | 202.56 | 207.49 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -1.6164 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
75 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516762 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.7 | 187.86 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.48684 | 20.9 | 1.074 | 1.062 | 193.51 | 203.46 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.54973 | 21 | 1.074 | 1.062 | 203.46 | 208.45 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
76 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516763 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 127.84 | 227.49 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.99354 | 20.9 | 1.074 | 1.062 | 233.06 | 243 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.64968 | 21 | 1.074 | 1.062 | 243 | 248 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
77 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516764 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 86.55 | 186.31 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.88425 | 20.9 | 1.074 | 1.062 | 191.9 | 201.9 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.049975 | 21 | 1.074 | 1.062 | 201.9 | 206.9 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.21952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
78 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516765 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.24 | 187.36 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.52658 | 20.9 | 1.074 | 1.062 | 193.01 | 202.96 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.54973 | 21 | 1.074 | 1.062 | 202.96 | 207.95 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -0.41908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
79 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516766 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.86 | 187.72 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.7849 | 20.9 | 1.074 | 1.062 | 193.31 | 203.19 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -1.2494 | 21 | 1.074 | 1.062 | 203.19 | 208.07 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -2.6142 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
80 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516767 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 87.04 | 187.01 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -0.67561 | 20.9 | 1.074 | 1.062 | 192.65 | 202.58 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.74963 | 21 | 1.074 | 1.062 | 202.58 | 207.52 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -1.4169 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
81 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516768 | 25.6 | 0 | 21.2 | 8.208 | 0.578 | 86.26 | 185.72 | 100.65 | 2,32,500 | 2 | 32,500 | 0.042361 | 2376 | -1.1823 | 20.9 | 1.074 | 1.062 | 191.35 | 201.26 | 10.005 | 22,28,810 | 22 | 28,810 | 0.37467 | 26.7035 | -0.94953 | 21 | 1.074 | 1.062 | 201.26 | 206.22 | 5.011 | 18,22,110 | 18 | 22,110 | 0.30614 | 16.3682 | -1.0178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
82 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516769 | 25.6 | 0 | 21 | 8.217 | 0.558 | 127.56 | 227.42 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.68623 | 21 | 1.069 | 1.06 | 233.53 | 243.46 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.8982 | 21 | 1.07 | 1.062 | 243.46 | 248.5 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
83 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516770 | 25.6 | 0 | 21 | 8.217 | 0.558 | 90.1 | 189.65 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.99453 | 21 | 1.069 | 1.06 | 195.75 | 205.7 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.6986 | 21 | 1.07 | 1.062 | 205.7 | 210.7 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
84 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516771 | 25.6 | 0 | 21 | 8.217 | 0.558 | 91.5 | 191.75 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.29836 | 21 | 1.069 | 1.06 | 197.88 | 207.9 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | 1.0637e-13 | 21 | 1.07 | 1.062 | 207.9 | 212.94 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
85 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516772 | 25.6 | 0 | 21 | 8.217 | 0.558 | 90.49 | 190.04 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.99453 | 21 | 1.069 | 1.06 | 196.15 | 206.15 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.1996 | 21 | 1.07 | 1.062 | 206.15 | 211.16 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
86 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516773 | 25.6 | 0 | 21 | 8.217 | 0.558 | 90.75 | 190.11 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -1.1835 | 21 | 1.069 | 1.06 | 196.2 | 206.14 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.7984 | 21 | 1.07 | 1.062 | 206.14 | 211.12 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
87 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516774 | 25.6 | 0 | 21 | 8.217 | 0.558 | 89.8 | 189.39 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.95475 | 21 | 1.069 | 1.06 | 195.47 | 205.4 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.8982 | 21 | 1.07 | 1.062 | 205.4 | 210.33 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -1.5968 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
88 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516775 | 25.6 | 0 | 21 | 8.217 | 0.558 | 90.65 | 190.4 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.79562 | 21 | 1.069 | 1.06 | 196.51 | 206.49 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.3992 | 21 | 1.07 | 1.062 | 206.49 | 211.46 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
89 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516776 | 25.6 | 0 | 21 | 8.217 | 0.558 | 90.5 | 190.5 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.54699 | 21 | 1.069 | 1.06 | 196.63 | 206.62 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.2994 | 21 | 1.07 | 1.062 | 206.62 | 211.65 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | 0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
90 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516777 | 25.6 | 0 | 21 | 8.217 | 0.558 | 91.2 | 190.28 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -1.462 | 21 | 1.069 | 1.06 | 196.35 | 206.3 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.6986 | 21 | 1.07 | 1.062 | 206.3 | 211.3 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
91 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516778 | 25.6 | 0 | 21 | 8.217 | 0.558 | 91.06 | 190.17 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -1.4321 | 21 | 1.069 | 1.06 | 196.25 | 206.19 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.7984 | 21 | 1.07 | 1.062 | 206.19 | 211.13 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
92 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516779 | 25.6 | 0 | 21 | 8.217 | 0.558 | 91.15 | 190.95 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.7459 | 21 | 1.069 | 1.06 | 197.06 | 207.1 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | 0.1996 | 21 | 1.07 | 1.062 | 207.1 | 212.12 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
93 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516780 | 25.6 | 0 | 21 | 8.217 | 0.558 | 92.06 | 191.96 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.64644 | 21 | 1.069 | 1.06 | 198.06 | 208 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.7984 | 21 | 1.07 | 1.062 | 208 | 213.01 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
94 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516781 | 25.6 | 0 | 21 | 8.217 | 0.558 | 93.75 | 193.65 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.64644 | 21 | 1.069 | 1.06 | 199.75 | 209.75 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.1996 | 21 | 1.07 | 1.062 | 209.75 | 214.71 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
95 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516782 | 25.6 | 0 | 21 | 8.217 | 0.558 | 91.28 | 190.7 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -1.1238 | 21 | 1.069 | 1.06 | 196.8 | 206.77 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.499 | 21 | 1.07 | 1.062 | 206.77 | 211.78 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
96 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516783 | 25.6 | 0 | 21 | 8.217 | 0.558 | 91.82 | 191.62 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.7459 | 21 | 1.069 | 1.06 | 197.74 | 207.64 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -1.1976 | 21 | 1.07 | 1.062 | 207.64 | 212.56 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | -1.7964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
97 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516784 | 25.6 | 0 | 21 | 8.217 | 0.558 | 92.45 | 192.2 | 100.55 | 2,32,290 | 2 | 32,290 | 0.042303 | 2376.9125 | -0.79562 | 21 | 1.069 | 1.06 | 198.29 | 208.23 | 10.02 | 22,28,180 | 22 | 28,180 | 0.37449 | 26.7561 | -0.7984 | 21 | 1.07 | 1.062 | 208.23 | 213.25 | 5.01 | 18,21,540 | 18 | 21,540 | 0.30598 | 16.3734 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
98 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516785 | 25.6 | 0 | 21 | 8.209 | 0.551 | 154 | 253.6 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.84619 | 20.9 | 1.076 | 1.067 | 259.51 | 269.46 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.55966 | 21 | 1.074 | 1.068 | 269.46 | 274.4 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -1.3775 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
99 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516786 | 25.6 | 0 | 21 | 8.209 | 0.551 | 77.3 | 176.61 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.1349 | 20.9 | 1.076 | 1.067 | 182.54 | 192.49 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.55966 | 21 | 1.074 | 1.068 | 192.49 | 197.46 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -0.7786 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
100 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516787 | 25.6 | 0 | 21 | 8.209 | 0.551 | 77.19 | 176.95 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.68691 | 20.9 | 1.076 | 1.067 | 182.9 | 192.86 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.45972 | 21 | 1.074 | 1.068 | 192.86 | 197.88 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | 0.2196 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
101 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516788 | 25.6 | 0 | 21 | 8.209 | 0.551 | 77.39 | 176.39 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.4435 | 20.9 | 1.076 | 1.067 | 182.3 | 192.25 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.55966 | 21 | 1.074 | 1.068 | 192.25 | 197.26 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | 0.019964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
102 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516789 | 25.6 | 0 | 21 | 8.209 | 0.551 | 78.59 | 179 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.039821 | 20.9 | 1.076 | 1.067 | 184.98 | 195 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | 0.13992 | 21 | 1.074 | 1.068 | 195 | 200 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -0.17968 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
103 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516790 | 25.6 | 0 | 21 | 8.209 | 0.551 | 77.91 | 178.1 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.25884 | 20.9 | 1.076 | 1.067 | 184.05 | 194.03 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.25984 | 21 | 1.074 | 1.068 | 194.03 | 199.04 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | 0.019964 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
104 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516791 | 25.6 | 0 | 21 | 8.209 | 0.551 | 78 | 178.2 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.24888 | 20.9 | 1.076 | 1.067 | 184.19 | 194.13 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.6596 | 21 | 1.074 | 1.068 | 194.13 | 199.12 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
105 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516792 | 25.6 | 0 | 21 | 8.209 | 0.551 | 78.18 | 178.28 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.34843 | 20.9 | 1.076 | 1.067 | 184.2 | 194.07 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -1.3592 | 21 | 1.074 | 1.068 | 194.07 | 198.98 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -1.9764 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
106 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516793 | 25.6 | 0 | 21 | 8.209 | 0.551 | 80.04 | 179.09 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.3937 | 20.9 | 1.076 | 1.067 | 185 | 194.95 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.55966 | 21 | 1.074 | 1.068 | 194.95 | 199.94 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
107 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516794 | 25.6 | 0 | 21 | 8.209 | 0.551 | 79.62 | 178.86 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.2046 | 20.9 | 1.076 | 1.067 | 184.76 | 194.76 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.059964 | 21 | 1.074 | 1.068 | 194.76 | 199.75 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -0.37932 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
108 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516795 | 25.6 | 0 | 21 | 8.209 | 0.551 | 79.9 | 178.65 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.6924 | 20.9 | 1.076 | 1.067 | 184.55 | 194.44 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -1.1593 | 21 | 1.074 | 1.068 | 194.44 | 199.4 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -0.97824 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
109 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516796 | 25.6 | 0 | 21 | 8.209 | 0.551 | 80 | 179.35 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.0951 | 20.9 | 1.076 | 1.067 | 185.26 | 195.22 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.45972 | 21 | 1.074 | 1.068 | 195.22 | 200.29 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | 1.2178 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
110 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516797 | 25.6 | 0 | 21 | 8.209 | 0.551 | 81.28 | 181.04 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.68691 | 20.9 | 1.076 | 1.067 | 186.95 | 196.88 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.75954 | 21 | 1.074 | 1.068 | 196.88 | 201.83 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -1.1779 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
111 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516798 | 25.6 | 0 | 21 | 8.209 | 0.551 | 78.8 | 179.4 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | 0.14933 | 20.9 | 1.076 | 1.067 | 185.35 | 195.4 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | 0.43974 | 21 | 1.074 | 1.068 | 195.4 | 200.45 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | 0.81853 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
112 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516799 | 25.6 | 0 | 21 | 8.209 | 0.551 | 83 | 182.3 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -1.1448 | 20.9 | 1.076 | 1.067 | 188.21 | 198.12 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.95942 | 21 | 1.074 | 1.068 | 198.12 | 203.04 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | -1.7768 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
113 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516800 | 25.6 | 0 | 21 | 8.209 | 0.551 | 79 | 179.22 | 100.45 | 2,31,460 | 2 | 31,460 | 0.042072 | 2387.5611 | -0.22897 | 20.9 | 1.076 | 1.067 | 185.2 | 195.2 | 10.006 | 22,28,700 | 22 | 28,700 | 0.37464 | 26.7084 | -0.059964 | 21 | 1.074 | 1.068 | 195.2 | 200.24 | 5.009 | 18,22,780 | 18 | 22,780 | 0.30633 | 16.3518 | 0.61889 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
114 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516801 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 103.44 | 203 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -1.0436 | 20.9 | 1.071 | 1.062 | 208.35 | 218.32 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.34983 | 21 | 1.07 | 1.062 | 218.32 | 223.3 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
115 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516802 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 102.45 | 202.55 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.50691 | 20.9 | 1.071 | 1.062 | 207.95 | 217.91 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.44978 | 21 | 1.07 | 1.062 | 217.91 | 222.9 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
116 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516803 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 102.25 | 201.94 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.91442 | 20.9 | 1.071 | 1.062 | 207.29 | 217.22 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.74963 | 21 | 1.07 | 1.062 | 217.22 | 222.19 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.7984 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
117 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516804 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 102.78 | 202.26 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -1.1231 | 20.9 | 1.071 | 1.062 | 207.6 | 217.58 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.24988 | 21 | 1.07 | 1.062 | 217.58 | 222.58 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
118 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516805 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 101.64 | 201.05 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -1.1927 | 20.9 | 1.071 | 1.062 | 206.4 | 216.36 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.44978 | 21 | 1.07 | 1.062 | 216.36 | 221.32 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
119 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516806 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 99.39 | 199.49 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.50691 | 20.9 | 1.071 | 1.062 | 204.85 | 214.76 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.94953 | 21 | 1.07 | 1.062 | 214.76 | 219.72 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
120 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516807 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 142.29 | 242.27 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.62618 | 20.9 | 1.071 | 1.062 | 247.66 | 257.69 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | 0.24988 | 21 | 1.07 | 1.062 | 257.69 | 262.75 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | 0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
121 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516808 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 101.25 | 201.25 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.6063 | 20.9 | 1.071 | 1.062 | 206.6 | 216.55 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.54973 | 21 | 1.07 | 1.062 | 216.55 | 221.51 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
122 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516809 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 142.9 | 242.77 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.73551 | 20.9 | 1.071 | 1.062 | 248.06 | 257.96 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -1.0495 | 21 | 1.07 | 1.062 | 257.96 | 262.84 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -2.5948 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
123 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516810 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 102.31 | 202.51 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.40751 | 20.9 | 1.071 | 1.062 | 207.9 | 217.82 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.84958 | 21 | 1.07 | 1.062 | 217.82 | 222.8 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
124 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516811 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 101.3 | 201.15 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.75539 | 20.9 | 1.071 | 1.062 | 206.54 | 216.46 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.84958 | 21 | 1.07 | 1.062 | 216.46 | 221.41 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
125 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516812 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 101.45 | 201.2 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.85479 | 20.9 | 1.071 | 1.062 | 206.55 | 216.46 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.94953 | 21 | 1.07 | 1.062 | 216.46 | 221.42 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
126 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516813 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 101.6 | 201.65 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.5566 | 20.9 | 1.071 | 1.062 | 207.01 | 217.03 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | 0.14993 | 21 | 1.07 | 1.062 | 217.03 | 222.05 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | 0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
127 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516814 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 141.16 | 241.36 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.40751 | 20.9 | 1.071 | 1.062 | 246.74 | 256.66 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.84958 | 21 | 1.07 | 1.062 | 256.66 | 261.62 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
128 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516815 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 139.05 | 238.31 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -1.3418 | 20.9 | 1.071 | 1.062 | 243.65 | 253.52 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -1.3493 | 21 | 1.07 | 1.062 | 253.52 | 258.46 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
129 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516816 | 25.6 | 0 | 20.8 | 8.212 | 0.562 | 101.55 | 201.39 | 100.61 | 2,32,230 | 2 | 32,230 | 0.042286 | 2379.2682 | -0.76533 | 20.9 | 1.071 | 1.062 | 206.75 | 216.71 | 10.005 | 22,30,400 | 22 | 30,400 | 0.37511 | 26.6721 | -0.44978 | 21 | 1.07 | 1.062 | 216.71 | 221.66 | 5.01 | 18,24,550 | 18 | 24,550 | 0.30682 | 16.3288 | -1.1976 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
130 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516817 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 123.38 | 222.65 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -1.3221 | 20.8 | 1.07 | 1.064 | 228.71 | 238.66 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.67878 | 20.9 | 1.068 | 1.064 | 238.66 | 243.65 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
131 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516818 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 123.26 | 223.79 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.069583 | 20.8 | 1.07 | 1.064 | 229.9 | 239.88 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.37932 | 20.9 | 1.068 | 1.064 | 239.88 | 244.86 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
132 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516819 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 122.65 | 222.56 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.68588 | 20.8 | 1.07 | 1.064 | 228.6 | 238.6 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.17968 | 20.9 | 1.068 | 1.064 | 238.6 | 243.61 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
133 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516820 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 82.84 | 182.84 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.59642 | 20.8 | 1.07 | 1.064 | 188.9 | 198.82 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.97824 | 20.9 | 1.068 | 1.064 | 198.82 | 203.78 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
134 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516821 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 82.5 | 183 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.099404 | 20.8 | 1.07 | 1.064 | 189.09 | 199.1 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.079856 | 20.9 | 1.068 | 1.064 | 199.1 | 204.1 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
135 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516822 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 120.6 | 220.19 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -1.004 | 20.8 | 1.07 | 1.064 | 226.21 | 236.17 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.57896 | 20.9 | 1.068 | 1.064 | 236.17 | 241.17 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
136 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516823 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 83.54 | 183.4 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.73559 | 20.8 | 1.07 | 1.064 | 189.45 | 199.36 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -1.0781 | 20.9 | 1.068 | 1.064 | 199.36 | 204.31 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
137 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516824 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 84.08 | 184.35 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.32803 | 20.8 | 1.07 | 1.064 | 190.44 | 200.45 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.079856 | 20.9 | 1.068 | 1.064 | 200.45 | 205.51 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
138 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516825 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 84 | 184.28 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.31809 | 20.8 | 1.07 | 1.064 | 190.33 | 200.3 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.47914 | 20.9 | 1.068 | 1.064 | 200.3 | 205.32 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
139 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516826 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 83.65 | 184.05 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.19881 | 20.8 | 1.07 | 1.064 | 190.11 | 200.11 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.17968 | 20.9 | 1.068 | 1.064 | 200.11 | 205.14 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
140 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516827 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 83.49 | 183.85 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.23857 | 20.8 | 1.07 | 1.064 | 189.95 | 199.95 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.17968 | 20.9 | 1.068 | 1.064 | 199.95 | 205 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
141 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516828 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 83.55 | 183.39 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.75547 | 20.8 | 1.07 | 1.064 | 189.44 | 199.41 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.47914 | 20.9 | 1.068 | 1.064 | 199.41 | 204.4 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
142 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516829 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 83.9 | 184.15 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.34791 | 20.8 | 1.07 | 1.064 | 190.25 | 200.21 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.57896 | 20.9 | 1.068 | 1.064 | 200.21 | 205.2 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
143 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516830 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 82.45 | 182.71 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.33797 | 20.8 | 1.07 | 1.064 | 188.75 | 198.72 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.47914 | 20.9 | 1.068 | 1.064 | 198.72 | 203.72 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
144 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516831 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 82.9 | 182.65 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.84493 | 20.8 | 1.07 | 1.064 | 188.71 | 198.63 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.97824 | 20.9 | 1.068 | 1.064 | 198.63 | 203.55 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -1.8943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
145 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516832 | 25.6 | 0 | 21.4 | 8.202 | 0.56 | 83.52 | 183.65 | 100.6 | 2,32,100 | 2 | 32,100 | 0.04225 | 2381.0651 | -0.4672 | 20.8 | 1.07 | 1.064 | 189.71 | 199.72 | 10.018 | 22,25,120 | 22 | 25,120 | 0.37364 | 26.8116 | -0.079856 | 20.9 | 1.068 | 1.064 | 199.72 | 204.71 | 5.015 | 18,25,660 | 18 | 25,660 | 0.30713 | 16.3287 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
146 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516833 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 77.84 | 177.25 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -1.1829 | 20.9 | 1.07 | 1.062 | 183.48 | 193.36 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -1.2888 | 20.9 | 1.066 | 1.063 | 193.36 | 198.36 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
147 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516834 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 76.45 | 175.8 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -1.2425 | 20.9 | 1.07 | 1.062 | 182.06 | 192.01 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.58947 | 20.9 | 1.066 | 1.063 | 192.01 | 196.96 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
148 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516835 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 76.66 | 176.74 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.5169 | 20.9 | 1.07 | 1.062 | 183 | 192.86 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -1.4887 | 20.9 | 1.066 | 1.063 | 192.86 | 197.45 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -8.4746 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
149 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516836 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 76.45 | 176.46 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.58648 | 20.9 | 1.07 | 1.062 | 182.77 | 192.75 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.28974 | 20.9 | 1.066 | 1.063 | 192.75 | 197.7 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
150 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516837 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 75.98 | 175.26 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -1.3121 | 20.9 | 1.07 | 1.062 | 181.5 | 191.39 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -1.1889 | 20.9 | 1.066 | 1.063 | 191.39 | 196.36 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
151 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516838 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 76.8 | 176.7 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.69583 | 20.9 | 1.07 | 1.062 | 183 | 193 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.089919 | 20.9 | 1.066 | 1.063 | 193 | 198.05 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
152 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516839 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 78.8 | 179.24 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.15905 | 20.9 | 1.07 | 1.062 | 185.55 | 195.51 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.48956 | 20.9 | 1.066 | 1.063 | 195.51 | 200.49 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
153 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516840 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 78.15 | 178 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.74553 | 20.9 | 1.07 | 1.062 | 184.31 | 194.26 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.58947 | 20.9 | 1.066 | 1.063 | 194.26 | 199.26 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
154 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516841 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 76.26 | 175.61 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -1.2425 | 20.9 | 1.07 | 1.062 | 181.91 | 191.85 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.68938 | 20.9 | 1.066 | 1.063 | 191.85 | 196.81 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
155 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516842 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 77.35 | 178.05 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | 0.099404 | 20.9 | 1.07 | 1.062 | 184.34 | 194.33 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.18983 | 20.9 | 1.066 | 1.063 | 194.33 | 199.36 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | 0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
156 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516843 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 76.95 | 176.98 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.5666 | 20.9 | 1.07 | 1.062 | 183.25 | 193.21 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.48956 | 20.9 | 1.066 | 1.063 | 193.21 | 198.2 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
157 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516844 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 77.1 | 177.31 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.38767 | 20.9 | 1.07 | 1.062 | 183.62 | 193.57 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.58947 | 20.9 | 1.066 | 1.063 | 193.57 | 198.58 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
158 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516845 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 77.55 | 178.05 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.099404 | 20.9 | 1.07 | 1.062 | 184.35 | 194.26 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.98911 | 20.9 | 1.066 | 1.063 | 194.26 | 199.22 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -1.0967 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
159 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516846 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 78.88 | 179.09 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.38767 | 20.9 | 1.07 | 1.062 | 185.4 | 195.35 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.58947 | 20.9 | 1.066 | 1.063 | 195.35 | 200.3 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
160 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516847 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 78.9 | 179 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.49702 | 20.9 | 1.07 | 1.062 | 185.26 | 195.21 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.58947 | 20.9 | 1.066 | 1.063 | 195.21 | 200.22 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
161 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516848 | 25.6 | 0 | 21.2 | 8.21 | 0.564 | 77.55 | 177.6 | 100.6 | 2,32,140 | 2 | 32,140 | 0.042261 | 2380.4391 | -0.54672 | 20.9 | 1.07 | 1.062 | 183.9 | 193.86 | 10.009 | 22,30,670 | 22 | 30,670 | 0.37519 | 26.6774 | -0.48956 | 20.9 | 1.066 | 1.063 | 193.86 | 198.85 | 5.015 | 18,27,010 | 18 | 27,010 | 0.3075 | 16.3088 | -0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
162 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516849 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 78.39 | 179.1 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.31876 | 20.9 | 1.071 | 1.065 | 183.55 | 193.55 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.0999 | 20.9 | 1.074 | 1.064 | 193.55 | 198.56 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
163 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516850 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 78.55 | 179.55 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.60763 | 20.9 | 1.071 | 1.065 | 184 | 193.96 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.4995 | 20.9 | 1.074 | 1.064 | 193.96 | 199 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
164 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516851 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 79.31 | 180.15 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.44825 | 20.9 | 1.071 | 1.065 | 184.6 | 194.61 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | 1.952e-13 | 20.9 | 1.074 | 1.064 | 194.61 | 199.59 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | -1 | 11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
165 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516852 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 75.24 | 175 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -0.62755 | 20.9 | 1.071 | 1.065 | 179.4 | 189.36 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.4995 | 20.9 | 1.074 | 1.064 | 189.36 | 194.32 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
166 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516853 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 76.16 | 176.29 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -0.25899 | 20.9 | 1.071 | 1.065 | 180.69 | 190.62 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.7992 | 20.9 | 1.074 | 1.064 | 190.62 | 195.6 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
167 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516854 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 76.56 | 177.1 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.14942 | 20.9 | 1.071 | 1.065 | 181.55 | 191.52 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.3996 | 20.9 | 1.074 | 1.064 | 191.52 | 196.48 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.998 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
168 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516855 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 76.25 | 176.45 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -0.18926 | 20.9 | 1.071 | 1.065 | 180.9 | 190.86 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.4995 | 20.9 | 1.074 | 1.064 | 190.86 | 195.85 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
169 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516856 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 75.81 | 176.1 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -0.099612 | 20.9 | 1.071 | 1.065 | 180.55 | 190.42 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -1.3986 | 20.9 | 1.074 | 1.064 | 190.42 | 195.36 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
170 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516857 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 76.9 | 177.46 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.16934 | 20.9 | 1.071 | 1.065 | 181.9 | 191.9 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.0999 | 20.9 | 1.074 | 1.064 | 191.9 | 196.94 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
171 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516858 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 76.75 | 177.38 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.23907 | 20.9 | 1.071 | 1.065 | 181.8 | 191.82 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | 0.0999 | 20.9 | 1.074 | 1.064 | 191.82 | 196.86 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | 0.5988 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
172 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516859 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 116.06 | 215.35 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -1.0957 | 20.9 | 1.071 | 1.065 | 219.75 | 229.7 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.5994 | 20.9 | 1.074 | 1.064 | 229.7 | 234.7 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
173 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516860 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 77.47 | 178.04 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.1793 | 20.9 | 1.071 | 1.065 | 182.49 | 192.39 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -1.0989 | 20.9 | 1.074 | 1.064 | 192.39 | 197.33 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -1.3972 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
174 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516861 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 116.49 | 216.66 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -0.21915 | 20.9 | 1.071 | 1.065 | 221.1 | 231.05 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -0.5994 | 20.9 | 1.074 | 1.064 | 231.05 | 236.05 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.1996 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
175 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516862 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 117.33 | 218.2 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.47814 | 20.9 | 1.071 | 1.065 | 222.65 | 232.29 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -3.6963 | 20.9 | 1.074 | 1.064 | 232.29 | 236.79 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -10.1796 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | NO CONFORME
176 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516863 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 117.5 | 218.25 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | 0.3586 | 20.9 | 1.071 | 1.065 | 222.7 | 232.74 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | 0.2997 | 20.9 | 1.074 | 1.064 | 232.74 | 237.75 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -1.7728e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | -1 | 11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
177 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516864 | 25.6 | 0 | 21.2 | 8.209 | 0.529 | 306 | 406.28 | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -0.10957 | 20.9 | 1.071 | 1.065 | 410.7 | 420.71 | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -8.8729e-14 | 20.9 | 1.074 | 1.064 | 420.71 | 425.7 | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -0.3992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
178 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516865 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
179 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516866 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
180 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516867 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
181 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516868 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
182 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516869 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
183 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516870 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
184 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516871 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
185 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516872 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
186 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516873 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
187 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516874 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
188 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516875 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
189 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516876 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
190 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516877 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
191 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516878 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
192 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516879 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
193 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516880 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
194 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516881 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
195 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516882 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
196 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516883 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
197 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516884 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
198 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516885 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
199 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516886 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
200 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516887 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
201 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516888 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
202 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516889 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
203 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516890 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
204 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516891 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
205 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516892 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
206 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516893 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
207 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516894 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
208 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516895 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
209 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516896 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
210 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516897 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
211 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516898 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
212 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516899 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
213 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516900 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
214 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516901 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
215 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516902 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
216 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516903 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
217 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516904 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
218 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516905 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
219 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516906 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
220 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516907 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
221 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516908 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
222 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516909 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
223 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516910 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
224 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516911 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
225 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516912 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
226 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516913 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
227 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516914 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
228 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516915 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
229 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516916 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
230 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516917 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
231 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516918 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
232 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516919 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
233 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516920 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
234 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516921 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
235 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516922 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
236 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516923 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
237 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516924 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
238 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516925 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
239 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516926 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
240 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516927 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
241 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516928 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
242 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516929 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
243 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516930 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
244 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516931 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
245 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516932 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
246 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516933 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
247 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516934 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
248 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516935 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
249 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516936 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
250 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516937 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
251 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516938 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
252 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516939 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
253 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516940 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
254 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516941 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
255 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516942 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
256 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516943 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
257 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516944 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
258 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516945 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
259 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516946 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
260 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516947 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
261 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516948 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
262 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516949 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
263 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516950 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
264 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516951 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
265 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516952 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
266 | 2025-09-20 | 2025-09-20 | 3 | 11 |  | PA25516953 | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
267 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
268 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
269 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
270 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
271 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
272 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
273 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
274 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
275 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
276 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
277 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
278 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
279 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
280 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
281 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
282 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
283 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
284 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
285 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
286 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
287 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
288 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
289 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
290 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
291 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
292 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
293 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
294 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
295 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
296 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
297 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
298 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
299 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
300 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
301 | 2025-09-20 | 2025-09-20 | 3 | 11 |  |  | 25.6 | 0 | 21.2 | 8.209 | 0.529 |  |  | 100.39 | 2,32,090 | 2 | 32,090 | 0.042247 | 2376.2509 | -100 | 20.9 | 1.071 | 1.065 |  |  | 10.01 | 22,21,980 | 22 | 21,980 | 0.37277 | 26.8529 | -100 | 20.9 | 1.074 | 1.064 | 0 |  | 5.01 | 18,24,480 | 18 | 24,480 | 0.3068 | 16.3299 | -100 | NO CONFORME | TRUE | FALSE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | NO CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
127
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 3 | PA25517169 | 0 | -0.95475 | -0.34913 | -0.3992 | CONFORME | Medición
2 | 3 | PA25517170 | 0 | -0.9448 | -0.54863 | 3.9002e-13 | CONFORME | Medición
3 | 3 | PA25517171 | 0 | -0.58677 | -0.44888 | 0.1996 | CONFORME | Medición
4 | 3 | PA25517172 | 0 | -0.65639 | -0.74813 | 0.7984 | CONFORME | Medición
5 | 3 | PA25517173 | 0 | -0.72601 | -0.34913 | 0.3992 | CONFORME | Medición
6 | 3 | PA25517174 | 0 | -0.44754 | -0.64838 | 3.9002e-13 | CONFORME | Medición
7 | 3 | PA25517175 | 0 | -0.53705 | -0.64838 | 0.5988 | CONFORME | Medición
8 | 3 | PA25517176 | 0 | -0.43759 | -0.84788 | -1.7964 | CONFORME | Medición
9 | 3 | PA25517177 | 0 | -0.18896 | -0.34913 | 0.7984 | CONFORME | Medición
10 | 3 | PA25517178 | 0 | -0.35803 | -0.14963 | 0.998 | CONFORME | Medición
11 | 3 | PA25517179 | 0 | -0.76579 | -0.24938 | 3.9002e-13 | CONFORME | Medición
12 | 3 | PA25517180 | 0 | -0.15912 | 0.049875 | 0.3992 | CONFORME | Medición
13 | 3 | PA25516701 | 0 | -0.51716 | 0.24938 | 0.7984 | CONFORME | Medición
14 | 3 | PA25516702 | 0 | -0.83541 | -0.24938 | 0.3992 | CONFORME | Medición
15 | 3 | PA25516703 | 0 | -1.0443 | -0.14963 | 1.1976 | CONFORME | Medición
16 | 3 | PA25516704 | 0 | -0.73595 | 0.54863 | 1.1976 | CONFORME | Medición
17 | 3 | PA25516705 | 0 | -0.68595 | -0.40996 | -1.0989 | CONFORME | Medición
18 | 3 | PA25516706 | 0 | -1.3719 | -0.50995 | 0.0999 | CONFORME | Medición
19 | 3 | PA25516707 | 0 | -1.2725 | -0.70993 | -1.4985 | CONFORME | Medición
20 | 3 | PA25516708 | 0 | -0.79531 | -0.40996 | -0.4995 | CONFORME | Medición
21 | 3 | PA25516709 | 0 | -1.0041 | -0.30997 | 0.4995 | CONFORME | Medición
22 | 3 | PA25516710 | 0 | -0.95437 | 0.28997 | 0.2997 | CONFORME | Medición
23 | 3 | PA25516711 | 0 | -0.50701 | -0.30997 | -0.4995 | CONFORME | Medición
24 | 3 | PA25516712 | 0 | -0.98419 | -1.0099 | 0.6993 | CONFORME | Medición
25 | 3 | PA25516713 | 0 | -0.33801 | 0.28997 | -0.8991 | CONFORME | Medición
26 | 3 | PA25516714 | 0 | -1.9087 | -5.0095 | -11.2887 | NO CONFORME | Medición
27 | 3 | PA25516715 | 0 | -0.29824 | -0.009999 | -0.0999 | CONFORME | Medición
28 | 3 | PA25516716 | 0 | -0.49707 | -0.10999 | 0.6993 | CONFORME | Medición
29 | 3 | PA25516717 | 0 | -0.21871 | -0.10999 | 0.4995 | CONFORME | Medición
30 | 3 | PA25516718 | 0 | -0.89472 | -0.40996 | -0.2997 | CONFORME | Medición
31 | 3 | PA25516719 | 0 | -0.52689 | -0.20998 | 0.2997 | CONFORME | Medición
32 | 3 | PA25516720 | 0 | -0.6263 | -0.10999 | 0.0999 | CONFORME | Medición
33 | 3 | PA25516721 | 0 | -1.1835 | -0.39841 | -0.019996 | CONFORME | Medición
34 | 3 | PA25516722 | 0 | -0.49727 | -0.39841 | -0.019996 | CONFORME | Medición
35 | 3 | PA25516723 | 0 | -0.90502 | -1.8924 | -0.41992 | CONFORME | Medición
36 | 3 | PA25516724 | 0 | -0.099453 | 0.099602 | 0.37992 | CONFORME | Medición
37 | 3 | PA25516725 | 0 | -0.49727 | 0.099602 | -0.019996 | CONFORME | Medición
38 | 3 | PA25516726 | 0 | -0.64644 | -0.59761 | -0.21996 | CONFORME | Medición
39 | 3 | PA25516727 | 0 | -0.83541 | -0.49801 | -0.21996 | CONFORME | Medición
40 | 3 | PA25516728 | 0 | -0.88513 | -0.39841 | -0.21996 | CONFORME | Medición
41 | 3 | PA25516729 | 0 | -1.2034 | 0.099602 | 0.37992 | CONFORME | Medición
42 | 3 | PA25516730 | 0 | -0.73595 | -0.79681 | -1.0198 | CONFORME | Medición
43 | 3 | PA25516731 | 0 | -0.77573 | -0.59761 | -0.21996 | CONFORME | Medición
44 | 3 | PA25516732 | 0 | -0.51716 | -0.1992 | -0.41992 | CONFORME | Medición
45 | 3 | PA25516733 | 0 | -0.7459 | -0.2988 | -1.4197 | CONFORME | Medición
46 | 3 | PA25516734 | 0 | -0.3083 | -0.099602 | -0.019996 | CONFORME | Medición
48 | 3 | PA25516735 | 0 | -0.32819 | 0.099602 | 0.37992 | CONFORME | Medición
49 | 3 | PA25516736 | 0 | -0.49727 | -0.099602 | -0.81984 | CONFORME | Medición
50 | 3 | PA25516737 | 0 | -0.9841 | -0.2997 | 0.69791 | CONFORME | Medición
51 | 3 | PA25516738 | 0 | -0.6163 | -0.4995 | 0.2991 | CONFORME | Medición
52 | 3 | PA25516739 | 0 | -1.1233 | 0.3996 | 1.6949 | CONFORME | Medición
53 | 3 | PA25516740 | 0 | -0.38767 | -0.4995 | -0.2991 | CONFORME | Medición
54 | 3 | PA25516741 | 0 | -1.2525 | -0.4995 | 0.2991 | CONFORME | Medición
55 | 3 | PA25516742 | 0 | -0.80517 | -8.8729e-14 | 1.2961 | CONFORME | Medición
56 | 3 | PA25516743 | 0 | -0.83499 | -0.2997 | 0.69791 | CONFORME | Medición
57 | 3 | PA25516744 | 0 | -0.58648 | 0.0999 | 0.69791 | CONFORME | Medición
58 | 3 | PA25516745 | 0 | -0.91451 | -0.2997 | -0.099701 | CONFORME | Medición
59 | 3 | PA25516746 | 0 | 0.15905 | 0.2997 | 0.69791 | CONFORME | Medición
60 | 3 | PA25516747 | 0 | -0.14911 | -0.5994 | -1.0967 | CONFORME | Medición
61 | 3 | PA25516748 | 0 | -0.85487 | -0.8991 | -2.0937 | CONFORME | Medición
62 | 3 | PA25516749 | 0 | -0.68588 | -0.6993 | -0.69791 | CONFORME | Medición
63 | 3 | PA25516750 | 0 | -0.049702 | -0.3996 | -0.89731 | CONFORME | Medición
64 | 3 | PA25516751 | 0 | -0.63618 | -0.8991 | -1.6949 | CONFORME | Medición
65 | 3 | PA25516752 | 0 | -0.64612 | -0.1998 | -0.099701 | CONFORME | Medición
66 | 3 | PA25516753 | 0 | -1.1624 | -0.64968 | -1.4169 | CONFORME | Medición
67 | 3 | PA25516754 | 0 | -1.2519 | -0.54973 | -0.019956 | CONFORME | Medición
68 | 3 | PA25516755 | 0 | -0.93393 | 0.049975 | 1.5765 | CONFORME | Medición
69 | 3 | PA25516756 | 0 | -0.97367 | -0.64968 | -0.61864 | CONFORME | Medición
70 | 3 | PA25516757 | 0 | -0.88425 | -1.1494 | -1.4169 | CONFORME | Medición
71 | 3 | PA25516758 | 0 | -1.0532 | -0.34983 | -0.41908 | CONFORME | Medición
72 | 3 | PA25516759 | 0 | -1.4009 | -0.84958 | -0.41908 | CONFORME | Medición
73 | 3 | PA25516760 | 0 | -0.82464 | -0.44978 | -0.41908 | CONFORME | Medición
74 | 3 | PA25516761 | 0 | -0.88425 | -0.94953 | -1.6164 | CONFORME | Medición
75 | 3 | PA25516762 | 0 | -0.48684 | -0.54973 | -0.41908 | CONFORME | Medición
76 | 3 | PA25516763 | 0 | -0.99354 | -0.64968 | -0.21952 | CONFORME | Medición
77 | 3 | PA25516764 | 0 | -0.88425 | -0.049975 | -0.21952 | CONFORME | Medición
78 | 3 | PA25516765 | 0 | -0.52658 | -0.54973 | -0.41908 | CONFORME | Medición
79 | 3 | PA25516766 | 0 | -0.7849 | -1.2494 | -2.6142 | CONFORME | Medición
80 | 3 | PA25516767 | 0 | -0.67561 | -0.74963 | -1.4169 | CONFORME | Medición
81 | 3 | PA25516768 | 0 | -1.1823 | -0.94953 | -1.0178 | CONFORME | Medición
82 | 3 | PA25516769 | 0 | -0.68623 | -0.8982 | 0.5988 | CONFORME | Medición
83 | 3 | PA25516770 | 0 | -0.99453 | -0.6986 | -0.1996 | CONFORME | Medición
84 | 3 | PA25516771 | 0 | -0.29836 | 1.0637e-13 | 0.5988 | CONFORME | Medición
85 | 3 | PA25516772 | 0 | -0.99453 | -0.1996 | -1.7728e-13 | CONFORME | Medición
86 | 3 | PA25516773 | 0 | -1.1835 | -0.7984 | -0.5988 | CONFORME | Medición
87 | 3 | PA25516774 | 0 | -0.95475 | -0.8982 | -1.5968 | CONFORME | Medición
88 | 3 | PA25516775 | 0 | -0.79562 | -0.3992 | -0.7984 | CONFORME | Medición
89 | 3 | PA25516776 | 0 | -0.54699 | -0.2994 | 0.3992 | CONFORME | Medición
90 | 3 | PA25516777 | 0 | -1.462 | -0.6986 | -0.1996 | CONFORME | Medición
91 | 3 | PA25516778 | 0 | -1.4321 | -0.7984 | -1.3972 | CONFORME | Medición
92 | 3 | PA25516779 | 0 | -0.7459 | 0.1996 | 0.1996 | CONFORME | Medición
93 | 3 | PA25516780 | 0 | -0.64644 | -0.7984 | -1.7728e-13 | CONFORME | Medición
94 | 3 | PA25516781 | 0 | -0.64644 | -0.1996 | -0.998 | CONFORME | Medición
95 | 3 | PA25516782 | 0 | -1.1238 | -0.499 | -1.7728e-13 | CONFORME | Medición
96 | 3 | PA25516783 | 0 | -0.7459 | -1.1976 | -1.7964 | CONFORME | Medición
97 | 3 | PA25516784 | 0 | -0.79562 | -0.7984 | 0.1996 | CONFORME | Medición
98 | 3 | PA25516785 | 0 | -0.84619 | -0.55966 | -1.3775 | CONFORME | Medición
99 | 3 | PA25516786 | 0 | -1.1349 | -0.55966 | -0.7786 | CONFORME | Medición
100 | 3 | PA25516787 | 0 | -0.68691 | -0.45972 | 0.2196 | CONFORME | Medición
101 | 3 | PA25516788 | 0 | -1.4435 | -0.55966 | 0.019964 | CONFORME | Medición
102 | 3 | PA25516789 | 0 | -0.039821 | 0.13992 | -0.17968 | CONFORME | Medición
103 | 3 | PA25516790 | 0 | -0.25884 | -0.25984 | 0.019964 | CONFORME | Medición
104 | 3 | PA25516791 | 0 | -0.24888 | -0.6596 | -0.37932 | CONFORME | Medición
105 | 3 | PA25516792 | 0 | -0.34843 | -1.3592 | -1.9764 | CONFORME | Medición
106 | 3 | PA25516793 | 0 | -1.3937 | -0.55966 | -0.37932 | CONFORME | Medición
107 | 3 | PA25516794 | 0 | -1.2046 | -0.059964 | -0.37932 | CONFORME | Medición
108 | 3 | PA25516795 | 0 | -1.6924 | -1.1593 | -0.97824 | CONFORME | Medición
109 | 3 | PA25516796 | 0 | -1.0951 | -0.45972 | 1.2178 | CONFORME | Medición
110 | 3 | PA25516797 | 0 | -0.68691 | -0.75954 | -1.1779 | CONFORME | Medición
111 | 3 | PA25516798 | 0 | 0.14933 | 0.43974 | 0.81853 | CONFORME | Medición
112 | 3 | PA25516799 | 0 | -1.1448 | -0.95942 | -1.7768 | CONFORME | Medición
113 | 3 | PA25516800 | 0 | -0.22897 | -0.059964 | 0.61889 | CONFORME | Medición
114 | 3 | PA25516801 | 0 | -1.0436 | -0.34983 | -0.5988 | CONFORME | Medición
115 | 3 | PA25516802 | 0 | -0.50691 | -0.44978 | -0.3992 | CONFORME | Medición
116 | 3 | PA25516803 | 0 | -0.91442 | -0.74963 | -0.7984 | CONFORME | Medición
117 | 3 | PA25516804 | 0 | -1.1231 | -0.24988 | -0.1996 | CONFORME | Medición
118 | 3 | PA25516805 | 0 | -1.1927 | -0.44978 | -0.998 | CONFORME | Medición
119 | 3 | PA25516806 | 0 | -0.50691 | -0.94953 | -0.998 | CONFORME | Medición
120 | 3 | PA25516807 | 0 | -0.62618 | 0.24988 | 0.998 | CONFORME | Medición
121 | 3 | PA25516808 | 0 | -0.6063 | -0.54973 | -0.998 | CONFORME | Medición
122 | 3 | PA25516809 | 0 | -0.73551 | -1.0495 | -2.5948 | CONFORME | Medición
123 | 3 | PA25516810 | 0 | -0.40751 | -0.84958 | -0.5988 | CONFORME | Medición
124 | 3 | PA25516811 | 0 | -0.75539 | -0.84958 | -1.1976 | CONFORME | Medición
125 | 3 | PA25516812 | 0 | -0.85479 | -0.94953 | -0.998 | CONFORME | Medición
126 | 3 | PA25516813 | 0 | -0.5566 | 0.14993 | 0.1996 | CONFORME | Medición
127 | 3 | PA25516814 | 0 | -0.40751 | -0.84958 | -0.998 | CONFORME | Medición
128 | 3 | PA25516815 | 0 | -1.3418 | -1.3493 | -1.3972 | CONFORME | Medición
129 | 3 | PA25516816 | 0 | -0.76533 | -0.44978 | -1.1976 | CONFORME | Medición
130 | 3 | PA25516817 | 0 | -1.3221 | -0.67878 | -0.4985 | CONFORME | Medición
131 | 3 | PA25516818 | 0 | -0.069583 | -0.37932 | -0.69791 | CONFORME | Medición
132 | 3 | PA25516819 | 0 | -0.68588 | -0.17968 | -0.099701 | CONFORME | Medición
133 | 3 | PA25516820 | 0 | -0.59642 | -0.97824 | -1.0967 | CONFORME | Medición
134 | 3 | PA25516821 | 0 | -0.099404 | -0.079856 | -0.2991 | CONFORME | Medición
135 | 3 | PA25516822 | 0 | -1.004 | -0.57896 | -0.2991 | CONFORME | Medición
136 | 3 | PA25516823 | 0 | -0.73559 | -1.0781 | -1.2961 | CONFORME | Medición
137 | 3 | PA25516824 | 0 | -0.32803 | -0.079856 | 0.89731 | CONFORME | Medición
138 | 3 | PA25516825 | 0 | -0.31809 | -0.47914 | 0.099701 | CONFORME | Medición
139 | 3 | PA25516826 | 0 | -0.19881 | -0.17968 | 0.2991 | CONFORME | Medición
140 | 3 | PA25516827 | 0 | -0.23857 | -0.17968 | 0.69791 | CONFORME | Medición
141 | 3 | PA25516828 | 0 | -0.75547 | -0.47914 | -0.4985 | CONFORME | Medición
142 | 3 | PA25516829 | 0 | -0.34791 | -0.57896 | -0.4985 | CONFORME | Medición
143 | 3 | PA25516830 | 0 | -0.33797 | -0.47914 | -0.2991 | CONFORME | Medición
144 | 3 | PA25516831 | 0 | -0.84493 | -0.97824 | -1.8943 | CONFORME | Medición
145 | 3 | PA25516832 | 0 | -0.4672 | -0.079856 | -0.4985 | CONFORME | Medición
146 | 3 | PA25516833 | 0 | -1.1829 | -1.2888 | -0.2991 | CONFORME | Medición
147 | 3 | PA25516834 | 0 | -1.2425 | -0.58947 | -1.2961 | CONFORME | Medición
148 | 3 | PA25516835 | 0 | -0.5169 | -1.4887 | -8.4746 | NO CONFORME | Medición
149 | 3 | PA25516836 | 0 | -0.58648 | -0.28974 | -1.2961 | CONFORME | Medición
150 | 3 | PA25516837 | 0 | -1.3121 | -1.1889 | -0.89731 | CONFORME | Medición
151 | 3 | PA25516838 | 0 | -0.69583 | -0.089919 | 0.69791 | CONFORME | Medición
152 | 3 | PA25516839 | 0 | -0.15905 | -0.48956 | -0.69791 | CONFORME | Medición
153 | 3 | PA25516840 | 0 | -0.74553 | -0.58947 | -0.2991 | CONFORME | Medición
154 | 3 | PA25516841 | 0 | -1.2425 | -0.68938 | -1.0967 | CONFORME | Medición
155 | 3 | PA25516842 | 0 | 0.099404 | -0.18983 | 0.2991 | CONFORME | Medición
156 | 3 | PA25516843 | 0 | -0.5666 | -0.48956 | -0.4985 | CONFORME | Medición
157 | 3 | PA25516844 | 0 | -0.38767 | -0.58947 | -0.099701 | CONFORME | Medición
158 | 3 | PA25516845 | 0 | -0.099404 | -0.98911 | -1.0967 | CONFORME | Medición
159 | 3 | PA25516846 | 0 | -0.38767 | -0.58947 | -1.2961 | CONFORME | Medición
160 | 3 | PA25516847 | 0 | -0.49702 | -0.58947 | -0.099701 | CONFORME | Medición
161 | 3 | PA25516848 | 0 | -0.54672 | -0.48956 | -0.4985 | CONFORME | Medición
162 | 3 | PA25516849 | 0 | 0.31876 | -0.0999 | -1.7728e-13 | CONFORME | Medición
163 | 3 | PA25516850 | 0 | 0.60763 | -0.4995 | 0.5988 | CONFORME | Medición
164 | 3 | PA25516851 | 0 | 0.44825 | 1.952e-13 | -0.5988 | CONFORME | Medición
165 | 3 | PA25516852 | 0 | -0.62755 | -0.4995 | -0.998 | CONFORME | Medición
166 | 3 | PA25516853 | 0 | -0.25899 | -0.7992 | -0.5988 | CONFORME | Medición
167 | 3 | PA25516854 | 0 | 0.14942 | -0.3996 | -0.998 | CONFORME | Medición
168 | 3 | PA25516855 | 0 | -0.18926 | -0.4995 | -0.3992 | CONFORME | Medición
169 | 3 | PA25516856 | 0 | -0.099612 | -1.3986 | -1.3972 | CONFORME | Medición
170 | 3 | PA25516857 | 0 | 0.16934 | -0.0999 | 0.5988 | CONFORME | Medición
171 | 3 | PA25516858 | 0 | 0.23907 | 0.0999 | 0.5988 | CONFORME | Medición
172 | 3 | PA25516859 | 0 | -1.0957 | -0.5994 | -0.1996 | CONFORME | Medición
173 | 3 | PA25516860 | 0 | 0.1793 | -1.0989 | -1.3972 | CONFORME | Medición
174 | 3 | PA25516861 | 0 | -0.21915 | -0.5994 | -0.1996 | CONFORME | Medición
175 | 3 | PA25516862 | 0 | 0.47814 | -3.6963 | -10.1796 | NO CONFORME | Medición
176 | 3 | PA25516863 | 0 | 0.3586 | 0.2997 | -1.7728e-13 | CONFORME | Medición
177 | 3 | PA25516864 | 0 | -0.10957 | -8.8729e-14 | -0.3992 | CONFORME | Medición
178 | 3 | PA25516865 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
179 | 3 | PA25516866 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
180 | 3 | PA25516867 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
181 | 3 | PA25516868 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
182 | 3 | PA25516869 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
183 | 3 | PA25516870 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
184 | 3 | PA25516871 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
185 | 3 | PA25516872 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
186 | 3 | PA25516873 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
187 | 3 | PA25516874 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
188 | 3 | PA25516875 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
189 | 3 | PA25516876 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
190 | 3 | PA25516877 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
191 | 3 | PA25516878 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
192 | 3 | PA25516879 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
193 | 3 | PA25516880 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
194 | 3 | PA25516881 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
195 | 3 | PA25516882 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
196 | 3 | PA25516883 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
197 | 3 | PA25516884 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
198 | 3 | PA25516885 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
199 | 3 | PA25516886 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
200 | 3 | PA25516887 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
201 | 3 | PA25516888 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
202 | 3 | PA25516889 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
203 | 3 | PA25516890 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
204 | 3 | PA25516891 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
205 | 3 | PA25516892 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
206 | 3 | PA25516893 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
207 | 3 | PA25516894 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
208 | 3 | PA25516895 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
209 | 3 | PA25516896 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
210 | 3 | PA25516897 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
211 | 3 | PA25516898 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
212 | 3 | PA25516899 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
213 | 3 | PA25516900 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
214 | 3 | PA25516901 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
215 | 3 | PA25516902 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
216 | 3 | PA25516903 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
217 | 3 | PA25516904 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
218 | 3 | PA25516905 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
219 | 3 | PA25516906 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
220 | 3 | PA25516907 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
221 | 3 | PA25516908 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
222 | 3 | PA25516909 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
223 | 3 | PA25516910 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
224 | 3 | PA25516911 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
225 | 3 | PA25516912 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
226 | 3 | PA25516913 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
227 | 3 | PA25516914 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
228 | 3 | PA25516915 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
229 | 3 | PA25516916 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
230 | 3 | PA25516917 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
231 | 3 | PA25516918 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
232 | 3 | PA25516919 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
233 | 3 | PA25516920 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
234 | 3 | PA25516921 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
235 | 3 | PA25516922 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
236 | 3 | PA25516923 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
237 | 3 | PA25516924 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
238 | 3 | PA25516925 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
239 | 3 | PA25516926 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
240 | 3 | PA25516927 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
241 | 3 | PA25516928 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
242 | 3 | PA25516929 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
243 | 3 | PA25516930 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
244 | 3 | PA25516931 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
245 | 3 | PA25516932 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
246 | 3 | PA25516933 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
247 | 3 | PA25516934 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
248 | 3 | PA25516935 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
249 | 3 | PA25516936 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
250 | 3 | PA25516937 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
251 | 3 | PA25516938 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
252 | 3 | PA25516939 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
253 | 3 | PA25516940 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
254 | 3 | PA25516941 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
255 | 3 | PA25516942 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
256 | 3 | PA25516943 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
257 | 3 | PA25516944 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
258 | 3 | PA25516945 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
259 | 3 | PA25516946 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
260 | 3 | PA25516947 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
261 | 3 | PA25516948 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
262 | 3 | PA25516949 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
263 | 3 | PA25516950 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
264 | 3 | PA25516951 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
265 | 3 | PA25516952 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
266 | 3 | PA25516953 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
267 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
268 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
269 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
270 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
271 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
272 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
273 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
274 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
275 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
276 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
277 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
278 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
279 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
280 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
281 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
282 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
283 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
284 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
285 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
286 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
287 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
288 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
289 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
290 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
291 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
292 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
293 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
294 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
295 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
296 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
297 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
298 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
299 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
300 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
301 | 3 | 0 | 0 | -100 | -100 | -100 | NO CONFORME | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición