## ERROR FINAL
|  |  |  | FORMATO | [36 empty] | Código: F05-GA-PR-01 | [8 empty] | Q3/Q1 | 100 | 125 | 160 | 200 | 400 | 500
|  |  |  | INGRESO DE DATOS DE LA VERIFICACIÓN INICIAL DE MEDIDORES EN LA COMPUTADORA | [36 empty] | Versión: 08 | [8 empty] | Q3 (m3/h): | 1.6 | 2.5 | 4 | 6.3
Fecha de Vigencia: 2024-05-14
Q3 (m3/h): |  |  |  | 2.5 |  |  |  |  | Alcance Q3/Q1: |  |  |  |  | 160 |  |  |  | MAXIMO ERROR PERMISIBLE (EMP): | [12 empty] | Q3 |  | Q2 |  |  | Q1 |  |  | Revisado por: |  | - |  | Fecha: |  | -
Presión Estática: 0=conforme; 1 = Daño Físico; 2 = Fuga; 3 = Picado; 4 = Desfasado; 5 = Paralizado | [30 empty] | (±) 2% |  | (±) 2% |  |  | (±) 5% |  |  | Aprobado por: |  | - |  | Fecha: |  | -
Item | Fecha de Ensayo presion estatica | Fecha de Ensayo errores de indicacion | N° de banco de ensayo | 1' Técnico | Nro de ceritificado | # Medidor | Presión Estática |  | 2500 | [11 empty] | 25 | [11 empty] | 15.625 | [11 empty] | Estado
[9 empty] | Q3 | [11 empty] | Q2 | [11 empty] | Q1 | [16 empty] | RANGOS ACEPTABLES DE CAUDAL
[7 empty] | Presión (bar) | Estado | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q3 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q2 | E % | Temperatura | Presión Entrada | Presión Salida | L.I. | L.F. | Vol. P | Tiempo |  |  |  | Q1 | E % |  |  |  |  | Caudal | Q1 | Q2 | Q3 | Ley Signos y Errores
1 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517601 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132.58 | 232.65 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.49717 | 20.8 | 1.51 | 1.505 | 251 | 261 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.49751 | 20.7 | 1.51 | 1.505 | 261 | 266.09 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 1.1727 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
2 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517602 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132 | 231.95 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.61649 | 20.8 | 1.51 | 1.505 | 250.23 | 260.16 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -1.194 | 20.7 | 1.51 | 1.505 | 260.16 | 265.17 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -0.41741 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
3 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517603 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 133.6 | 234.2 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | 0.02983 | 20.8 | 1.51 | 1.505 | 252.6 | 262.63 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.199 | 20.7 | 1.51 | 1.505 | 262.63 | 267.69 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 0.57643 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
4 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517604 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 133.04 | 233.27 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.33807 | 20.8 | 1.51 | 1.505 | 251.69 | 261.69 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.49751 | 20.7 | 1.51 | 1.505 | 261.69 | 266.72 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -0.019877 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
5 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517605 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132 | 232.14 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.42756 | 20.8 | 1.51 | 1.505 | 250.47 | 260.45 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.69652 | 20.7 | 1.51 | 1.505 | 260.45 | 265.5 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
6 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517606 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132.62 | 232.75 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.43751 | 20.8 | 1.51 | 1.505 | 251.1 | 261.1 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.49751 | 20.7 | 1.51 | 1.505 | 261.1 | 266.15 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
7 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517607 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132.3 | 232.71 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.15909 | 20.8 | 1.51 | 1.505 | 251.05 | 261.05 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.49751 | 20.7 | 1.51 | 1.505 | 261.05 | 266.1 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
8 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517608 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132 | 232.09 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.47728 | 20.8 | 1.51 | 1.505 | 250.4 | 260.37 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.79602 | 20.7 | 1.51 | 1.505 | 260.37 | 265.4 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -0.019877 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
9 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517609 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 132.55 | 232.55 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.56677 | 20.8 | 1.51 | 1.505 | 250.88 | 260.8 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -1.2935 | 20.7 | 1.51 | 1.505 | 260.8 | 265.75 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -1.61 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
10 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517610 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 133.9 | 234.47 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | 0 | 20.8 | 1.51 | 1.505 | 252.85 | 262.77 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -1.2935 | 20.7 | 1.51 | 1.505 | 262.77 | 267.8 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -0.019877 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 0 | -1 | -1 | 0-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
11 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517611 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 121.7 | 221.3 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.9645 | 20.8 | 1.51 | 1.505 | 239.55 | 249.5 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.99502 | 20.7 | 1.51 | 1.505 | 249.5 | 254.33 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -3.9952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
12 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517612 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 122 | 222.3 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.26847 | 20.8 | 1.51 | 1.505 | 240.64 | 250.65 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.39801 | 20.7 | 1.51 | 1.505 | 250.65 | 255.75 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 1.3715 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
13 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517613 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 119.17 | 218.9 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.83524 | 20.8 | 1.51 | 1.505 | 237.16 | 247.2 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.099502 | 20.7 | 1.51 | 1.505 | 247.2 | 252.25 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
14 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517614 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 122.9 | 223.4 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.069603 | 20.8 | 1.51 | 1.505 | 241.7 | 251.7 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.49751 | 20.7 | 1.51 | 1.505 | 251.7 | 256.76 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | 0.57643 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
15 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517615 | 25.6 | 0 | 21.1 | 8.087 | 0.583 | 122.49 | 222.7 | 100.57 | 2,27,860 | 2 | 27,860 | 0.041072 | 2448.6136 | -0.35796 | 20.8 | 1.51 | 1.505 | 241.09 | 251.13 | 10.05 | 22,43,410 | 22 | 43,410 | 0.37873 | 26.5364 | -0.099502 | 20.7 | 1.51 | 1.505 | 251.13 | 256.15 | 5.031 | 18,21,790 | 18 | 21,790 | 0.30605 | 16.4383 | -0.21864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
16 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517616 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 165.75 | 265.45 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.75652 | 21.6 | 1.503 | 1.505 | 276.23 | 286.15 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.1952 | 21.5 | 1.506 | 1.504 | 286.15 | 291.25 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.73079 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
17 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517617 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 165.9 | 265.75 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.60721 | 21.6 | 1.503 | 1.505 | 276.55 | 286.45 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.3944 | 21.5 | 1.506 | 1.504 | 286.45 | 291.5 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -0.25676 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
18 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517618 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 163.5 | 263.2 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.75652 | 21.6 | 1.503 | 1.505 | 273.99 | 283.88 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.494 | 21.5 | 1.506 | 1.504 | 283.88 | 288.9 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -0.8493 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
19 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517619 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 166.86 | 267.03 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.28867 | 21.6 | 1.503 | 1.505 | 277.85 | 287.77 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.1952 | 21.5 | 1.506 | 1.504 | 287.77 | 292.83 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -0.059253 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
20 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517620 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 165.85 | 265.51 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.79634 | 21.6 | 1.503 | 1.505 | 276.3 | 286.22 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.1952 | 21.5 | 1.506 | 1.504 | 286.22 | 291.3 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.33577 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
21 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517621 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 164.57 | 264.66 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.36831 | 21.6 | 1.503 | 1.505 | 275.48 | 285.41 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.0956 | 21.5 | 1.506 | 1.504 | 285.41 | 290.44 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -0.65179 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
22 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517622 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 166.18 | 265.51 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -1.1248 | 21.6 | 1.503 | 1.505 | 276.26 | 286.22 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -0.79681 | 21.5 | 1.506 | 1.504 | 286.22 | 291.35 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 1.3233 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
23 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517623 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 160.5 | 259.85 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -1.1049 | 21.6 | 1.503 | 1.505 | 270.61 | 280.58 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -0.69721 | 21.5 | 1.506 | 1.504 | 280.58 | 285.65 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.13826 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
24 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517624 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 162.4 | 262.15 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.70675 | 21.6 | 1.503 | 1.505 | 272.9 | 282.85 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -0.89641 | 21.5 | 1.506 | 1.504 | 282.85 | 287.94 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.53328 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
25 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517625 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 161.6 | 261.5 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.55744 | 21.6 | 1.503 | 1.505 | 272.3 | 282.25 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -0.89641 | 21.5 | 1.506 | 1.504 | 282.25 | 287.3 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -0.25676 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
26 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517626 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 163.09 | 263.45 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.099542 | 21.6 | 1.503 | 1.505 | 274.33 | 284.25 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.1952 | 21.5 | 1.506 | 1.504 | 284.25 | 289.35 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.73079 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
27 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517627 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 162.35 | 261.85 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.9556 | 21.6 | 1.503 | 1.505 | 272.65 | 282.65 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -0.39841 | 21.5 | 1.506 | 1.504 | 282.65 | 287.73 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.33577 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
28 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517628 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 162.5 | 262.49 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.46785 | 21.6 | 1.503 | 1.505 | 273.31 | 283.24 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -1.0956 | 21.5 | 1.506 | 1.504 | 283.24 | 288.3 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -0.059253 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
29 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517629 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 165.59 | 266 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.049771 | 21.6 | 1.503 | 1.505 | 276.85 | 286.88 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | -0.099602 | 21.5 | 1.506 | 1.504 | 286.88 | 291.99 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | 0.9283 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
30 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517630 | 25.6 | 0 | 21.1 | 8.08 | 0.553 | 161.4 | 260.87 | 100.46 | 2,27,590 | 2 | 27,590 | 0.040997 | 2450.4099 | -0.98547 | 21.6 | 1.503 | 1.505 | 271.65 | 281.85 | 10.04 | 23,03,190 | 23 | 03,190 | 0.38422 | 26.1309 | 1.5936 | 21.5 | 1.506 | 1.504 | 281.85 | 286.76 | 5.063 | 18,22,580 | 18 | 22,580 | 0.30627 | 16.531 | -3.0219 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | -1 | -11-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
31 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517631 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 156.44 | 256.58 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.29869 | 21.5 | 1.51 | 1.508 | 267.8 | 277.7 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -1.61 | 21.2 | 1.508 | 1.504 | 277.7 | 282.63 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -1.871 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
32 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517632 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 154.91 | 254.25 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -1.0952 | 21.5 | 1.51 | 1.508 | 265.4 | 275.26 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -2.0076 | 21.2 | 1.508 | 1.504 | 275.26 | 280.2 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -1.672 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
33 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517633 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 156.05 | 255.65 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.83632 | 21.5 | 1.51 | 1.508 | 266.88 | 276.85 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.91433 | 21.2 | 1.508 | 1.504 | 276.85 | 281.88 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | 0.11943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
34 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517634 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 155.95 | 255.25 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -1.135 | 21.5 | 1.51 | 1.508 | 266.47 | 276.45 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.81495 | 21.2 | 1.508 | 1.504 | 276.45 | 281.45 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -0.47771 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
35 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517635 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 155.8 | 255 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -1.2346 | 21.5 | 1.51 | 1.508 | 266.15 | 276.05 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -1.61 | 21.2 | 1.508 | 1.504 | 276.05 | 281.05 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -0.47771 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
36 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517636 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 157.35 | 257.67 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.11947 | 21.5 | 1.51 | 1.508 | 268.96 | 278.95 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.71556 | 21.2 | 1.508 | 1.504 | 278.95 | 283.95 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -0.47771 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
37 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517637 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 156.99 | 256.99 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.43807 | 21.5 | 1.51 | 1.508 | 268.25 | 278.22 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.91433 | 21.2 | 1.508 | 1.504 | 278.22 | 283.29 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | 0.91561 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
38 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517638 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 157.1 | 257.47 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.069693 | 21.5 | 1.51 | 1.508 | 268.75 | 278.75 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.61618 | 21.2 | 1.508 | 1.504 | 278.75 | 283.7 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -1.4729 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
39 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517639 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 155.96 | 255.47 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.92593 | 21.5 | 1.51 | 1.508 | 266.7 | 276.7 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.61618 | 21.2 | 1.508 | 1.504 | 276.7 | 281.7 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -0.47771 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
40 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517640 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 159.15 | 259.39 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.19912 | 21.5 | 1.51 | 1.508 | 270.8 | 280.82 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.41741 | 21.2 | 1.508 | 1.504 | 280.82 | 285.86 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | 0.31847 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
41 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517641 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 156.89 | 256.8 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.52768 | 21.5 | 1.51 | 1.508 | 268.05 | 278.03 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.81495 | 21.2 | 1.508 | 1.504 | 278.03 | 283.05 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -0.079618 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
42 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517642 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 158.4 | 258.2 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.6372 | 21.5 | 1.51 | 1.508 | 269.45 | 279.4 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -1.1131 | 21.2 | 1.508 | 1.504 | 279.4 | 284.43 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | 0.11943 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
43 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517643 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 158.1 | 257.57 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.96575 | 21.5 | 1.51 | 1.508 | 268.76 | 278.68 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -1.4113 | 21.2 | 1.508 | 1.504 | 278.68 | 283.65 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | -1.0748 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
44 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517644 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 158.7 | 258.73 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.4082 | 21.5 | 1.51 | 1.508 | 269.99 | 280.05 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.019877 | 21.2 | 1.508 | 1.504 | 280.05 | 285.15 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | 1.5127 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
45 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517645 | 25.6 | 0 | 21.1 | 8.075 | 0.601 | 157.45 | 257.5 | 100.44 | 2,28,140 | 2 | 28,140 | 0.04115 | 2440.8262 | -0.38829 | 21.5 | 1.51 | 1.508 | 268.74 | 278.8 | 10.062 | 22,39,370 | 22 | 39,370 | 0.3776 | 26.647 | -0.019877 | 21.2 | 1.508 | 1.504 | 278.8 | 283.9 | 5.024 | 18,30,710 | 18 | 30,710 | 0.30853 | 16.2836 | 1.5127 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
46 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517646 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 170.05 | 270.3 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.18917 | 21.4 | 1.506 | 1.506 | 284 | 293.96 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.1807 | 21.3 | 1.507 | 1.505 | 293.96 | 299 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
47 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517647 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 168.34 | 268.05 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.7268 | 21.4 | 1.506 | 1.506 | 281.75 | 291.69 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.3791 | 21.3 | 1.507 | 1.505 | 291.69 | 296.7 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
48 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517648 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 167.9 | 267.78 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.55755 | 21.4 | 1.506 | 1.506 | 281.47 | 291.45 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -0.98224 | 21.3 | 1.507 | 1.505 | 291.45 | 296.46 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
49 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517649 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 168.3 | 267.78 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.95579 | 21.4 | 1.506 | 1.506 | 281.45 | 291.4 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.2799 | 21.3 | 1.507 | 1.505 | 291.4 | 296.39 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.59761 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
50 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517650 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 168.05 | 267.47 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -1.0155 | 21.4 | 1.506 | 1.506 | 281.2 | 291.15 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.2799 | 21.3 | 1.507 | 1.505 | 291.15 | 296.15 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
51 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517651 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 168.58 | 268.15 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.86619 | 21.4 | 1.506 | 1.506 | 281.84 | 291.71 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -2.0736 | 21.3 | 1.507 | 1.505 | 291.71 | 296.7 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.59761 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
52 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517652 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 168.65 | 268.3 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.78654 | 21.4 | 1.506 | 1.506 | 282 | 291.95 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.2799 | 21.3 | 1.507 | 1.505 | 291.95 | 296.95 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
53 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517653 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 169.54 | 269.09 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.8861 | 21.4 | 1.506 | 1.506 | 282.8 | 292.77 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.0815 | 21.3 | 1.507 | 1.505 | 292.77 | 297.83 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
54 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517654 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 168.99 | 268.8 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.62724 | 21.4 | 1.506 | 1.506 | 282.53 | 292.5 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.0815 | 21.3 | 1.507 | 1.505 | 292.5 | 297.53 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | 0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
55 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517655 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 167.8 | 267.5 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.73676 | 21.4 | 1.506 | 1.506 | 281.2 | 291.09 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.8752 | 21.3 | 1.507 | 1.505 | 291.09 | 296.1 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
56 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517656 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 169.17 | 269.15 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.45798 | 21.4 | 1.506 | 1.506 | 282.89 | 292.85 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.1807 | 21.3 | 1.507 | 1.505 | 292.85 | 297.91 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | 0.79681 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
57 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517657 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 169.33 | 270.04 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | 0.26882 | 21.4 | 1.506 | 1.506 | 283.76 | 293.75 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -0.88302 | 21.3 | 1.507 | 1.505 | 293.75 | 298.76 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -0.1992 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | -1 | 1-1-1 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
58 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517658 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 169.85 | 270.16 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.12943 | 21.4 | 1.506 | 1.506 | 283.9 | 293.78 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -1.9744 | 21.3 | 1.507 | 1.505 | 293.78 | 298.74 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -1.1952 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
59 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517659 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 169.76 | 269.98 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.21904 | 21.4 | 1.506 | 1.506 | 283.65 | 293.66 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -0.68459 | 21.3 | 1.507 | 1.505 | 293.66 | 298.7 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | 0.39841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
60 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517660 | 25.6 | 0 | 21.2 | 8.085 | 0.585 | 171.05 | 271.41 | 100.44 | 2,27,960 | 2 | 27,960 | 0.0411 | 2443.7956 | -0.07965 | 21.4 | 1.506 | 1.506 | 285.15 | 295.18 | 10.079 | 23,12,360 | 23 | 12,360 | 0.38677 | 26.0596 | -0.48616 | 21.3 | 1.507 | 1.505 | 295.18 | 300.2 | 5.02 | 18,13,280 | 18 | 13,280 | 0.30369 | 16.5301 | -3.5386e-13 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
61 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517661 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 147.97 | 248.35 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.38702 | 21.7 | 1.509 | 1.507 | 257.99 | 267.93 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.7885 | 21.3 | 1.507 | 1.505 | 267.93 | 272.91 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
62 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517662 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 145.49 | 245.25 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -1.0023 | 21.7 | 1.509 | 1.507 | 254.8 | 264.67 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.4872 | 21.3 | 1.507 | 1.505 | 264.67 | 269.6 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -1.6949 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
63 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517663 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 148.3 | 248.81 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.25801 | 21.7 | 1.509 | 1.507 | 258.45 | 268.35 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.1877 | 21.3 | 1.507 | 1.505 | 268.35 | 273.35 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.2991 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
64 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517664 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 149.43 | 249.6 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.59542 | 21.7 | 1.509 | 1.507 | 259.2 | 269.07 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.4872 | 21.3 | 1.507 | 1.505 | 269.07 | 274.05 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
65 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517665 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 146.39 | 246.51 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.64503 | 21.7 | 1.509 | 1.507 | 256.1 | 265.99 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.2876 | 21.3 | 1.507 | 1.505 | 265.99 | 270.97 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
66 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517666 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 146.75 | 247.15 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.36717 | 21.7 | 1.509 | 1.507 | 256.75 | 266.68 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.88831 | 21.3 | 1.507 | 1.505 | 266.68 | 271.69 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
67 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517667 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 150 | 249.95 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.81373 | 21.7 | 1.509 | 1.507 | 259.52 | 269.44 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.98812 | 21.3 | 1.507 | 1.505 | 269.44 | 274.45 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
68 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517668 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 149.6 | 250.39 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | 0.019847 | 21.7 | 1.509 | 1.507 | 260.01 | 269.95 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.7885 | 21.3 | 1.507 | 1.505 | 269.95 | 275 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | 0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
69 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517669 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 147.35 | 247.59 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.52595 | 21.7 | 1.509 | 1.507 | 257.19 | 267.06 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.4872 | 21.3 | 1.507 | 1.505 | 267.06 | 272 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -1.4955 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
70 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517670 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 147.8 | 247.85 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.7145 | 21.7 | 1.509 | 1.507 | 257.46 | 267.35 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.2876 | 21.3 | 1.507 | 1.505 | 267.35 | 272.3 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -1.2961 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
71 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517671 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 147.11 | 247.25 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.62519 | 21.7 | 1.509 | 1.507 | 256.85 | 266.7 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.6868 | 21.3 | 1.507 | 1.505 | 266.7 | 271.67 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
72 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517672 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 149.85 | 251 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | 0.3771 | 21.7 | 1.509 | 1.507 | 260.7 | 270.7 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.18964 | 21.3 | 1.507 | 1.505 | 270.7 | 275.76 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | 0.89731 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
73 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517673 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 148.13 | 248.2 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.69465 | 21.7 | 1.509 | 1.507 | 257.8 | 267.67 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -1.4872 | 21.3 | 1.507 | 1.505 | 267.67 | 272.65 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | -0.69791 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
74 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517674 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 149.5 | 249.63 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.63511 | 21.7 | 1.509 | 1.507 | 259.21 | 269.13 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.98812 | 21.3 | 1.507 | 1.505 | 269.13 | 274.15 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | 0.099701 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
75 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517675 | 25.6 | 0 | 21.2 | 8.077 | 0.562 | 146.1 | 245.95 | 100.77 | 2,28,180 | 2 | 28,180 | 0.041161 | 2448.1846 | -0.91297 | 21.7 | 1.509 | 1.507 | 255.5 | 265.45 | 10.019 | 22,40,500 | 22 | 40,500 | 0.37792 | 26.5111 | -0.68869 | 21.3 | 1.507 | 1.505 | 265.45 | 270.49 | 5.015 | 18,21,840 | 18 | 21,840 | 0.30607 | 16.3853 | 0.4985 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
76 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517676 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 139 | 238.55 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.94527 | 21.2 | 1.507 | 1.508 | 245.85 | 255.9 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | 0.16944 | 21.3 | 1.51 | 1.507 | 255.9 | 261 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 1.5329 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
77 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517677 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 139.1 | 239.15 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.44776 | 21.2 | 1.507 | 1.508 | 246.49 | 256.47 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.52826 | 21.3 | 1.51 | 1.507 | 256.47 | 261.5 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.13936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
78 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517678 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 136.35 | 236.2 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.64677 | 21.2 | 1.507 | 1.508 | 243.47 | 253.4 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -1.0266 | 21.3 | 1.51 | 1.507 | 253.4 | 258.42 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -0.059725 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
79 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517679 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 139.85 | 239.68 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.66667 | 21.2 | 1.507 | 1.508 | 247 | 256.94 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.92694 | 21.3 | 1.51 | 1.507 | 256.94 | 261.97 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.13936 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
80 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517680 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 139.35 | 238.8 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -1.0448 | 21.2 | 1.507 | 1.508 | 246.07 | 255.97 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -1.3256 | 21.3 | 1.51 | 1.507 | 255.97 | 260.99 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -0.059725 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
81 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517681 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 138.35 | 238.89 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | 0.039801 | 21.2 | 1.507 | 1.508 | 246.25 | 256.24 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.42859 | 21.3 | 1.51 | 1.507 | 256.24 | 261.31 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.9357 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
82 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517682 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 137.75 | 237.15 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -1.0945 | 21.2 | 1.507 | 1.508 | 244.42 | 254.35 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -1.0266 | 21.3 | 1.51 | 1.507 | 254.35 | 259.35 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -0.45789 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
83 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517683 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 136.4 | 235.89 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -1.005 | 21.2 | 1.507 | 1.508 | 243.15 | 253.05 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -1.3256 | 21.3 | 1.51 | 1.507 | 253.05 | 258.07 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -0.059725 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
84 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517684 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 138.35 | 238.35 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.49751 | 21.2 | 1.507 | 1.508 | 245.63 | 255.6 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.62793 | 21.3 | 1.51 | 1.507 | 255.6 | 260.6 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -0.45789 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
85 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517685 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 137.98 | 237.3 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -1.1741 | 21.2 | 1.507 | 1.508 | 244.58 | 254.58 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.32891 | 21.3 | 1.51 | 1.507 | 254.58 | 259.64 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.73661 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
86 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517686 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 139 | 239.09 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.40796 | 21.2 | 1.507 | 1.508 | 246.4 | 256.36 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.7276 | 21.3 | 1.51 | 1.507 | 256.36 | 261.4 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.33844 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
87 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517687 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 138 | 238.1 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.39801 | 21.2 | 1.507 | 1.508 | 245.4 | 255.25 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -1.824 | 21.3 | 1.51 | 1.507 | 255.25 | 260.2 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -1.4533 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
88 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517688 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 138 | 237.08 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -1.4129 | 21.2 | 1.507 | 1.508 | 244.34 | 254.15 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -2.2227 | 21.3 | 1.51 | 1.507 | 254.15 | 259.1 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | -1.4533 | NO CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | NO CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | NO CONFORME
89 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517689 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 138.85 | 238.73 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.61692 | 21.2 | 1.507 | 1.508 | 246.04 | 255.96 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -1.1263 | 21.3 | 1.51 | 1.507 | 255.96 | 261 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.33844 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
90 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517690 | 25.6 | 0 | 21.2 | 8.095 | 0.524 | 138.65 | 238.4 | 100.5 | 2,27,250 | 2 | 27,250 | 0.040903 | 2457.0458 | -0.74627 | 21.2 | 1.507 | 1.508 | 245.66 | 255.6 | 10.033 | 23,07,200 | 23 | 07,200 | 0.38533 | 26.0372 | -0.92694 | 21.3 | 1.51 | 1.507 | 255.6 | 260.65 | 5.023 | 18,25,290 | 18 | 25,290 | 0.30702 | 16.3602 | 0.53753 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
91 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517691 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 149.71 | 249.46 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.82521 | 21.3 | 1.508 | 1.502 | 258.61 | 268.55 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.7984 | 20.9 | 1.508 | 1.505 | 268.55 | 273.55 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.833 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
92 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517692 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 149.65 | 249.27 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.95446 | 21.3 | 1.508 | 1.502 | 258.36 | 268.2 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.7964 | 20.9 | 1.508 | 1.505 | 268.2 | 273.2 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.833 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
93 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517693 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 145.94 | 244.8 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -1.7101 | 21.3 | 1.508 | 1.502 | 253.85 | 263.7 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.6966 | 20.9 | 1.508 | 1.505 | 263.7 | 268.7 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.833 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
94 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517694 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 147.6 | 247.37 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.80533 | 21.3 | 1.508 | 1.502 | 256.5 | 266.51 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.0998 | 20.9 | 1.508 | 1.505 | 266.51 | 271.63 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 1.547 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
95 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517695 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 147.29 | 246.8 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -1.0638 | 21.3 | 1.508 | 1.502 | 255.94 | 265.82 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.3972 | 20.9 | 1.508 | 1.505 | 265.82 | 270.87 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 0.15867 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
96 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517696 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 147.88 | 247.8 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.65619 | 21.3 | 1.508 | 1.502 | 256.91 | 266.84 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.8982 | 20.9 | 1.508 | 1.505 | 266.84 | 271.89 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 0.15867 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
97 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517697 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 147.27 | 247.29 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.55677 | 21.3 | 1.508 | 1.502 | 256.5 | 266.45 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.6986 | 20.9 | 1.508 | 1.505 | 266.45 | 271.53 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 0.75367 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
98 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517698 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 148.14 | 248.2 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.517 | 21.3 | 1.508 | 1.502 | 257.35 | 267.22 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.497 | 20.9 | 1.508 | 1.505 | 267.22 | 272.25 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.238 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
99 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517699 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 149.25 | 249.6 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.22867 | 21.3 | 1.508 | 1.502 | 258.8 | 268.8 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.1996 | 20.9 | 1.508 | 1.505 | 268.8 | 273.9 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 1.1503 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
100 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517700 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 150.1 | 250.27 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.40764 | 21.3 | 1.508 | 1.502 | 259.41 | 269.34 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.8982 | 20.9 | 1.508 | 1.505 | 269.34 | 274.36 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.43633 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
101 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517701 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 150.95 | 250.94 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.5866 | 21.3 | 1.508 | 1.502 | 260.1 | 270 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.1976 | 20.9 | 1.508 | 1.505 | 270 | 275.04 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.039667 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
102 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517702 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 148.9 | 248.79 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.68602 | 21.3 | 1.508 | 1.502 | 257.92 | 267.85 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.8982 | 20.9 | 1.508 | 1.505 | 267.85 | 272.9 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 0.15867 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
103 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517703 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 148.95 | 248.95 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.57666 | 21.3 | 1.508 | 1.502 | 258.14 | 268.04 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.1976 | 20.9 | 1.508 | 1.505 | 268.04 | 273.05 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | -0.63467 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
104 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517704 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 150.73 | 250.8 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | -0.50706 | 21.3 | 1.508 | 1.502 | 260 | 269.9 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -1.1976 | 20.9 | 1.508 | 1.505 | 269.9 | 274.95 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 0.15867 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
105 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517705 | 25.6 | 0 | 21.3 | 8.094 | 0.684 | 150.09 | 251 | 100.58 | 2,28,850 | 2 | 28,850 | 0.041347 | 2432.5697 | 0.3281 | 21.3 | 1.508 | 1.502 | 260.16 | 270.15 | 10.02 | 23,00,100 | 23 | 00,100 | 0.38336 | 26.1372 | -0.2994 | 20.9 | 1.508 | 1.505 | 270.15 | 275.25 | 5.042 | 18,24,310 | 18 | 24,310 | 0.30675 | 16.4367 | 1.1503 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | -1 | 1 | 1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
106 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517706 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 120.08 | 220.15 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.41795 | 21.1 | 1.505 | 1.502 | 226.25 | 236.2 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.84704 | 20.9 | 1.51 | 1.503 | 236.2 | 241.2 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | -0.61618 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
107 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517707 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 117.65 | 217.2 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.93542 | 21.1 | 1.505 | 1.502 | 223.27 | 233.2 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -1.0463 | 20.9 | 1.51 | 1.503 | 233.2 | 238.2 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | -0.61618 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
108 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517708 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 118.05 | 218.07 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.46771 | 21.1 | 1.505 | 1.502 | 224.18 | 234.15 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.64773 | 20.9 | 1.51 | 1.503 | 234.15 | 239.2 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
109 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517709 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 115.36 | 214.99 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.85581 | 21.1 | 1.505 | 1.502 | 221.05 | 231 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.84704 | 20.9 | 1.51 | 1.503 | 231 | 236.02 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | -0.21864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
110 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517710 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 114.65 | 213.7 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -1.433 | 21.1 | 1.505 | 1.502 | 219.76 | 229.7 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.94669 | 20.9 | 1.51 | 1.503 | 229.7 | 234.75 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
111 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517711 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 118.82 | 218 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -1.3036 | 21.1 | 1.505 | 1.502 | 224.05 | 234 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.84704 | 20.9 | 1.51 | 1.503 | 234 | 239 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | -0.61618 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
112 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517712 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 116.47 | 216.4 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.55727 | 21.1 | 1.505 | 1.502 | 222.5 | 232.45 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.84704 | 20.9 | 1.51 | 1.503 | 232.45 | 237.47 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | -0.21864 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
113 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517713 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 116 | 215.4 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -1.0847 | 21.1 | 1.505 | 1.502 | 221.49 | 231.45 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.74738 | 20.9 | 1.51 | 1.503 | 231.45 | 236.48 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | -0.019877 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
114 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517714 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 116.95 | 217 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.43785 | 21.1 | 1.505 | 1.502 | 223.09 | 233.02 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -1.0463 | 20.9 | 1.51 | 1.503 | 233.02 | 238.07 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
115 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517715 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 117.1 | 217.35 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.23883 | 21.1 | 1.505 | 1.502 | 223.45 | 233.4 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.84704 | 20.9 | 1.51 | 1.503 | 233.4 | 238.45 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
116 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517716 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 117.3 | 217.2 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.58712 | 21.1 | 1.505 | 1.502 | 223.3 | 233.26 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.74738 | 20.9 | 1.51 | 1.503 | 233.26 | 238.3 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.17889 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
117 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517717 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 117.09 | 216.85 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.72644 | 21.1 | 1.505 | 1.502 | 222.95 | 232.91 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.74738 | 20.9 | 1.51 | 1.503 | 232.91 | 237.98 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.77519 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
118 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517718 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 116.33 | 216.3 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.51746 | 21.1 | 1.505 | 1.502 | 222.4 | 232.35 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.84704 | 20.9 | 1.51 | 1.503 | 232.35 | 237.4 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
119 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517719 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 116.7 | 215.95 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -1.234 | 21.1 | 1.505 | 1.502 | 222.02 | 231.95 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -1.0463 | 20.9 | 1.51 | 1.503 | 231.95 | 237 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.37766 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
120 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517720 | 25.6 | 0 | 21.2 | 8.075 | 0.65 | 115.65 | 215.85 | 100.49 | 2,28,470 | 2 | 28,470 | 0.041242 | 2436.6135 | -0.28859 | 21.1 | 1.505 | 1.502 | 221.95 | 231.97 | 10.035 | 22,38,890 | 22 | 38,890 | 0.37747 | 26.5849 | -0.14948 | 20.9 | 1.51 | 1.503 | 231.97 | 237.05 | 5.031 | 18,24,320 | 18 | 24,320 | 0.30676 | 16.4007 | 0.97396 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
121 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517721 | 25.6 | 0 | 20 | 7.91 | 0.554 | 110.15 | 210.5 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.38714 | 19.9 | 1.498 | 1.493 | 216.25 | 226.21 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.48956 | 19.4 | 1.499 | 1.497 | 226.21 | 231.3 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 1.3742 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
122 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517722 | 25.6 | 0 | 20 | 7.91 | 0.554 | 110.2 | 209.85 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -1.082 | 19.9 | 1.498 | 1.493 | 215.55 | 225.45 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -1.089 | 19.4 | 1.499 | 1.497 | 225.45 | 230.5 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.57757 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
123 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517723 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.3 | 211.65 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.38714 | 19.9 | 1.498 | 1.493 | 217.38 | 227.38 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.089919 | 19.4 | 1.499 | 1.497 | 227.38 | 232.45 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.9759 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
124 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517724 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.95 | 213.15 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | 0.45662 | 19.9 | 1.498 | 1.493 | 218.9 | 228.96 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | 0.50954 | 19.4 | 1.499 | 1.497 | 228.96 | 234.09 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 2.1709 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | 1 | 1 | 1 | 111 | SI | SD | TRUE | SIGIGUALES | CONFORME
125 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517725 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.3 | 211.86 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.17868 | 19.9 | 1.498 | 1.493 | 217.6 | 227.55 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.58947 | 19.4 | 1.499 | 1.497 | 227.55 | 232.6 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.57757 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
126 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517726 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.45 | 211.46 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.72464 | 19.9 | 1.498 | 1.493 | 217.2 | 227.2 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.089919 | 19.4 | 1.499 | 1.497 | 227.2 | 232.25 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.57757 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
127 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517727 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.31 | 211.45 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.59559 | 19.9 | 1.498 | 1.493 | 217.2 | 227.15 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.58947 | 19.4 | 1.499 | 1.497 | 227.15 | 232.2 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.57757 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
128 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517728 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.55 | 211.98 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.30772 | 19.9 | 1.498 | 1.493 | 217.7 | 227.65 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.58947 | 19.4 | 1.499 | 1.497 | 227.65 | 232.69 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.37841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
129 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517729 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.5 | 211.5 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.73456 | 19.9 | 1.498 | 1.493 | 217.24 | 227.24 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.089919 | 19.4 | 1.499 | 1.497 | 227.24 | 232.3 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.77674 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
130 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517730 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.21 | 211.7 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.24816 | 19.9 | 1.498 | 1.493 | 217.45 | 227.39 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.68938 | 19.4 | 1.499 | 1.497 | 227.39 | 232.4 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | -0.21908 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
131 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517731 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.05 | 210.75 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -1.0324 | 19.9 | 1.498 | 1.493 | 216.48 | 226.41 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.78929 | 19.4 | 1.499 | 1.497 | 226.41 | 231.45 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.37841 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
132 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517732 | 25.6 | 0 | 20 | 7.91 | 0.554 | 110.6 | 210.75 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.58567 | 19.9 | 1.498 | 1.493 | 216.5 | 226.4 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -1.089 | 19.4 | 1.499 | 1.497 | 226.4 | 231.45 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.57757 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
133 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517733 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.6 | 211.9 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.43677 | 19.9 | 1.498 | 1.493 | 217.64 | 227.51 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -1.3888 | 19.4 | 1.499 | 1.497 | 227.51 | 232.5 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | -0.61741 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | -1 | -1-1-1 | SD | SI | TRUE | SIGIGUALES | CONFORME
134 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517734 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.9 | 212.45 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.1886 | 19.9 | 1.498 | 1.493 | 218.2 | 228.2 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | -0.089919 | 19.4 | 1.499 | 1.497 | 228.2 | 233.25 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 0.57757 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | -1 | 1 | -1-11 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
135 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517735 | 25.6 | 0 | 20 | 7.91 | 0.554 | 111.45 | 211.81 | 100.74 | 2,29,820 | 2 | 29,820 | 0.041617 | 2420.6648 | -0.37721 | 19.9 | 1.498 | 1.493 | 217.56 | 227.6 | 10.009 | 22,43,470 | 22 | 43,470 | 0.37874 | 26.427 | 0.30972 | 19.4 | 1.499 | 1.497 | 227.6 | 232.7 | 5.021 | 18,35,460 | 18 | 35,460 | 0.30985 | 16.2046 | 1.5734 | CONFORME | TRUE | TRUE | TRUE | Qcorrecto | aceptable | aceptable | aceptable | CONFORME | CONFORME | -1 | 1 | 1 | -111 | SD | SD | FALSE | SIGDIFERENTES | CONFORME
136 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517736 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
137 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517737 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
138 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517738 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
139 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517739 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
140 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517740 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
141 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517741 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
142 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517742 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
143 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517743 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
144 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517744 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
145 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517745 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
146 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517746 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
147 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517747 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
148 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517748 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
149 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517749 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
150 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517750 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
151 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517751 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
152 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517752 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
153 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517753 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
154 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517754 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
155 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517755 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
156 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517756 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
157 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517757 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
158 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517758 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
159 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517759 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
160 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517760 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
161 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517761 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
162 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517762 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
163 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517763 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
164 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517764 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
165 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517765 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
166 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517766 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
167 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517767 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
168 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517768 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
169 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517769 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
170 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517770 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
171 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517771 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
172 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517772 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
173 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517773 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
174 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517774 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
175 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517775 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
176 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517776 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
177 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517777 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
178 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517778 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
179 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517779 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
180 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517780 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
181 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517781 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
182 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517782 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
183 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517783 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
184 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517784 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
185 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517785 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
186 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517786 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
187 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517787 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
188 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517788 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
189 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517789 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
190 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517790 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
191 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517791 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
192 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517792 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
193 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517793 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
194 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517794 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
195 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517795 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
196 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517796 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
197 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517797 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
198 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517798 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
199 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517799 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
200 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517800 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
201 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517801 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
202 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517802 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
203 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517803 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
204 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517804 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
205 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517805 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
206 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517806 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
207 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517807 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
208 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517808 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
209 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517809 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
210 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517810 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
211 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517811 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
212 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517812 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
213 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517813 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
214 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517814 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
215 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517815 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
216 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517816 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
217 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517817 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
218 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517818 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
219 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517819 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
220 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517820 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
221 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517821 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
222 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517822 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
223 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517823 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
224 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517824 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
225 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517825 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
226 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517826 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
227 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517827 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
228 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517828 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
229 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517829 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
230 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517830 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
231 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517831 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
232 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517832 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
233 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517833 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
234 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517834 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
235 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517835 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
236 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517836 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
237 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517837 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
238 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517838 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
239 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517839 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
240 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517840 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
241 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517841 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
242 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517842 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
243 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517843 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
244 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517844 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
245 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517845 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
246 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517846 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
247 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517847 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
248 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517848 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
249 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517849 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
250 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517850 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
251 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517851 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
252 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517852 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
253 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517853 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
254 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517854 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
255 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517855 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
256 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517856 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
257 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517857 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
258 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517858 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
259 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517859 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
260 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517860 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
261 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517861 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
262 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517862 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
263 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517863 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
264 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517864 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
265 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517865 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
266 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517866 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
267 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517867 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
268 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517868 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
269 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517869 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
270 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517870 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
271 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517871 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
272 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517872 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
273 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517873 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
274 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517874 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
275 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517875 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
276 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517876 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
277 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517877 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
278 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517878 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
279 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517879 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
280 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517880 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
281 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517881 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
282 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517882 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
283 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517883 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
284 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517884 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
285 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517885 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
286 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517886 | 25.6 | 0 | [9 empty] | #VALUE! | #VALUE! | #DIV/0! | [9 empty] | #VALUE! | #VALUE! | #DIV/0! |  |  |  | 0 | [5 empty] | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
287 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517887 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
288 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517888 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
289 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517889 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
290 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517890 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
291 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517891 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
292 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517892 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
293 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517893 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
294 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517894 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
295 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517895 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
296 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517896 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
297 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517897 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
298 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517898 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
299 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517899 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
300 | 2025-09-20 | 2025-09-20 | 5 | 10 |  | PA25517900 | 25.6 | 0 | 0 | 0 | 0 |  |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 0 | 0 | 0 | 0 |  | 0 | 0:00 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | TRUE | FALSE | TRUE | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2
#REF!

## Hoja1
LO VERIFICADO | VIMA
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE
|  | TRUE | 000000-2024 | FALSE

## Hoja2 (2)
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición
1 | 5 | PA25517601 | 0 | -0.49717 | -0.49751 | 1.1727 | CONFORME | Medición
2 | 5 | PA25517602 | 0 | -0.61649 | -1.194 | -0.41741 | CONFORME | Medición
3 | 5 | PA25517603 | 0 | 0.02983 | -0.199 | 0.57643 | CONFORME | Medición
4 | 5 | PA25517604 | 0 | -0.33807 | -0.49751 | -0.019877 | CONFORME | Medición
5 | 5 | PA25517605 | 0 | -0.42756 | -0.69652 | 0.37766 | CONFORME | Medición
6 | 5 | PA25517606 | 0 | -0.43751 | -0.49751 | 0.37766 | CONFORME | Medición
7 | 5 | PA25517607 | 0 | -0.15909 | -0.49751 | 0.37766 | CONFORME | Medición
8 | 5 | PA25517608 | 0 | -0.47728 | -0.79602 | -0.019877 | CONFORME | Medición
9 | 5 | PA25517609 | 0 | -0.56677 | -1.2935 | -1.61 | CONFORME | Medición
10 | 5 | PA25517610 | 0 | 0 | -1.2935 | -0.019877 | CONFORME | Medición
11 | 5 | PA25517611 | 0 | -0.9645 | -0.99502 | -3.9952 | CONFORME | Medición
12 | 5 | PA25517612 | 0 | -0.26847 | -0.39801 | 1.3715 | CONFORME | Medición
13 | 5 | PA25517613 | 0 | -0.83524 | -0.099502 | 0.37766 | CONFORME | Medición
14 | 5 | PA25517614 | 0 | -0.069603 | -0.49751 | 0.57643 | CONFORME | Medición
15 | 5 | PA25517615 | 0 | -0.35796 | -0.099502 | -0.21864 | CONFORME | Medición
16 | 5 | PA25517616 | 0 | -0.75652 | -1.1952 | 0.73079 | CONFORME | Medición
17 | 5 | PA25517617 | 0 | -0.60721 | -1.3944 | -0.25676 | CONFORME | Medición
18 | 5 | PA25517618 | 0 | -0.75652 | -1.494 | -0.8493 | CONFORME | Medición
19 | 5 | PA25517619 | 0 | -0.28867 | -1.1952 | -0.059253 | CONFORME | Medición
20 | 5 | PA25517620 | 0 | -0.79634 | -1.1952 | 0.33577 | CONFORME | Medición
21 | 5 | PA25517621 | 0 | -0.36831 | -1.0956 | -0.65179 | CONFORME | Medición
22 | 5 | PA25517622 | 0 | -1.1248 | -0.79681 | 1.3233 | CONFORME | Medición
23 | 5 | PA25517623 | 0 | -1.1049 | -0.69721 | 0.13826 | CONFORME | Medición
24 | 5 | PA25517624 | 0 | -0.70675 | -0.89641 | 0.53328 | CONFORME | Medición
25 | 5 | PA25517625 | 0 | -0.55744 | -0.89641 | -0.25676 | CONFORME | Medición
26 | 5 | PA25517626 | 0 | -0.099542 | -1.1952 | 0.73079 | CONFORME | Medición
27 | 5 | PA25517627 | 0 | -0.9556 | -0.39841 | 0.33577 | CONFORME | Medición
28 | 5 | PA25517628 | 0 | -0.46785 | -1.0956 | -0.059253 | CONFORME | Medición
29 | 5 | PA25517629 | 0 | -0.049771 | -0.099602 | 0.9283 | CONFORME | Medición
30 | 5 | PA25517630 | 0 | -0.98547 | 1.5936 | -3.0219 | CONFORME | Medición
31 | 5 | PA25517631 | 0 | -0.29869 | -1.61 | -1.871 | CONFORME | Medición
32 | 5 | PA25517632 | 0 | -1.0952 | -2.0076 | -1.672 | CONFORME | Medición
33 | 5 | PA25517633 | 0 | -0.83632 | -0.91433 | 0.11943 | CONFORME | Medición
34 | 5 | PA25517634 | 0 | -1.135 | -0.81495 | -0.47771 | CONFORME | Medición
35 | 5 | PA25517635 | 0 | -1.2346 | -1.61 | -0.47771 | CONFORME | Medición
36 | 5 | PA25517636 | 0 | -0.11947 | -0.71556 | -0.47771 | CONFORME | Medición
37 | 5 | PA25517637 | 0 | -0.43807 | -0.91433 | 0.91561 | CONFORME | Medición
38 | 5 | PA25517638 | 0 | -0.069693 | -0.61618 | -1.4729 | CONFORME | Medición
39 | 5 | PA25517639 | 0 | -0.92593 | -0.61618 | -0.47771 | CONFORME | Medición
40 | 5 | PA25517640 | 0 | -0.19912 | -0.41741 | 0.31847 | CONFORME | Medición
41 | 5 | PA25517641 | 0 | -0.52768 | -0.81495 | -0.079618 | CONFORME | Medición
42 | 5 | PA25517642 | 0 | -0.6372 | -1.1131 | 0.11943 | CONFORME | Medición
43 | 5 | PA25517643 | 0 | -0.96575 | -1.4113 | -1.0748 | CONFORME | Medición
44 | 5 | PA25517644 | 0 | -0.4082 | -0.019877 | 1.5127 | CONFORME | Medición
45 | 5 | PA25517645 | 0 | -0.38829 | -0.019877 | 1.5127 | CONFORME | Medición
46 | 5 | PA25517646 | 0 | -0.18917 | -1.1807 | 0.39841 | CONFORME | Medición
47 | 5 | PA25517647 | 0 | -0.7268 | -1.3791 | -0.1992 | CONFORME | Medición
48 | 5 | PA25517648 | 0 | -0.55755 | -0.98224 | -0.1992 | CONFORME | Medición
49 | 5 | PA25517649 | 0 | -0.95579 | -1.2799 | -0.59761 | CONFORME | Medición
50 | 5 | PA25517650 | 0 | -1.0155 | -1.2799 | -0.39841 | CONFORME | Medición
51 | 5 | PA25517651 | 0 | -0.86619 | -2.0736 | -0.59761 | NO CONFORME | Medición
52 | 5 | PA25517652 | 0 | -0.78654 | -1.2799 | -0.39841 | CONFORME | Medición
53 | 5 | PA25517653 | 0 | -0.8861 | -1.0815 | 0.79681 | CONFORME | Medición
54 | 5 | PA25517654 | 0 | -0.62724 | -1.0815 | 0.1992 | CONFORME | Medición
55 | 5 | PA25517655 | 0 | -0.73676 | -1.8752 | -0.1992 | CONFORME | Medición
56 | 5 | PA25517656 | 0 | -0.45798 | -1.1807 | 0.79681 | CONFORME | Medición
57 | 5 | PA25517657 | 0 | 0.26882 | -0.88302 | -0.1992 | CONFORME | Medición
58 | 5 | PA25517658 | 0 | -0.12943 | -1.9744 | -1.1952 | CONFORME | Medición
59 | 5 | PA25517659 | 0 | -0.21904 | -0.68459 | 0.39841 | CONFORME | Medición
60 | 5 | PA25517660 | 0 | -0.07965 | -0.48616 | -3.5386e-13 | CONFORME | Medición
61 | 5 | PA25517661 | 0 | -0.38702 | -0.7885 | -0.69791 | CONFORME | Medición
62 | 5 | PA25517662 | 0 | -1.0023 | -1.4872 | -1.6949 | CONFORME | Medición
63 | 5 | PA25517663 | 0 | -0.25801 | -1.1877 | -0.2991 | CONFORME | Medición
64 | 5 | PA25517664 | 0 | -0.59542 | -1.4872 | -0.69791 | CONFORME | Medición
65 | 5 | PA25517665 | 0 | -0.64503 | -1.2876 | -0.69791 | CONFORME | Medición
66 | 5 | PA25517666 | 0 | -0.36717 | -0.88831 | -0.099701 | CONFORME | Medición
67 | 5 | PA25517667 | 0 | -0.81373 | -0.98812 | -0.099701 | CONFORME | Medición
68 | 5 | PA25517668 | 0 | 0.019847 | -0.7885 | 0.69791 | CONFORME | Medición
69 | 5 | PA25517669 | 0 | -0.52595 | -1.4872 | -1.4955 | CONFORME | Medición
70 | 5 | PA25517670 | 0 | -0.7145 | -1.2876 | -1.2961 | CONFORME | Medición
71 | 5 | PA25517671 | 0 | -0.62519 | -1.6868 | -0.89731 | CONFORME | Medición
72 | 5 | PA25517672 | 0 | 0.3771 | -0.18964 | 0.89731 | CONFORME | Medición
73 | 5 | PA25517673 | 0 | -0.69465 | -1.4872 | -0.69791 | CONFORME | Medición
74 | 5 | PA25517674 | 0 | -0.63511 | -0.98812 | 0.099701 | CONFORME | Medición
75 | 5 | PA25517675 | 0 | -0.91297 | -0.68869 | 0.4985 | CONFORME | Medición
76 | 5 | PA25517676 | 0 | -0.94527 | 0.16944 | 1.5329 | CONFORME | Medición
77 | 5 | PA25517677 | 0 | -0.44776 | -0.52826 | 0.13936 | CONFORME | Medición
78 | 5 | PA25517678 | 0 | -0.64677 | -1.0266 | -0.059725 | CONFORME | Medición
79 | 5 | PA25517679 | 0 | -0.66667 | -0.92694 | 0.13936 | CONFORME | Medición
80 | 5 | PA25517680 | 0 | -1.0448 | -1.3256 | -0.059725 | CONFORME | Medición
81 | 5 | PA25517681 | 0 | 0.039801 | -0.42859 | 0.9357 | CONFORME | Medición
82 | 5 | PA25517682 | 0 | -1.0945 | -1.0266 | -0.45789 | CONFORME | Medición
83 | 5 | PA25517683 | 0 | -1.005 | -1.3256 | -0.059725 | CONFORME | Medición
84 | 5 | PA25517684 | 0 | -0.49751 | -0.62793 | -0.45789 | CONFORME | Medición
85 | 5 | PA25517685 | 0 | -1.1741 | -0.32891 | 0.73661 | CONFORME | Medición
86 | 5 | PA25517686 | 0 | -0.40796 | -0.7276 | 0.33844 | CONFORME | Medición
87 | 5 | PA25517687 | 0 | -0.39801 | -1.824 | -1.4533 | CONFORME | Medición
88 | 5 | PA25517688 | 0 | -1.4129 | -2.2227 | -1.4533 | NO CONFORME | Medición
89 | 5 | PA25517689 | 0 | -0.61692 | -1.1263 | 0.33844 | CONFORME | Medición
90 | 5 | PA25517690 | 0 | -0.74627 | -0.92694 | 0.53753 | CONFORME | Medición
91 | 5 | PA25517691 | 0 | -0.82521 | -0.7984 | -0.833 | CONFORME | Medición
92 | 5 | PA25517692 | 0 | -0.95446 | -1.7964 | -0.833 | CONFORME | Medición
93 | 5 | PA25517693 | 0 | -1.7101 | -1.6966 | -0.833 | CONFORME | Medición
94 | 5 | PA25517694 | 0 | -0.80533 | -0.0998 | 1.547 | CONFORME | Medición
95 | 5 | PA25517695 | 0 | -1.0638 | -1.3972 | 0.15867 | CONFORME | Medición
96 | 5 | PA25517696 | 0 | -0.65619 | -0.8982 | 0.15867 | CONFORME | Medición
97 | 5 | PA25517697 | 0 | -0.55677 | -0.6986 | 0.75367 | CONFORME | Medición
98 | 5 | PA25517698 | 0 | -0.517 | -1.497 | -0.238 | CONFORME | Medición
99 | 5 | PA25517699 | 0 | -0.22867 | -0.1996 | 1.1503 | CONFORME | Medición
100 | 5 | PA25517700 | 0 | -0.40764 | -0.8982 | -0.43633 | CONFORME | Medición
101 | 5 | PA25517701 | 0 | -0.5866 | -1.1976 | -0.039667 | CONFORME | Medición
102 | 5 | PA25517702 | 0 | -0.68602 | -0.8982 | 0.15867 | CONFORME | Medición
103 | 5 | PA25517703 | 0 | -0.57666 | -1.1976 | -0.63467 | CONFORME | Medición
104 | 5 | PA25517704 | 0 | -0.50706 | -1.1976 | 0.15867 | CONFORME | Medición
105 | 5 | PA25517705 | 0 | 0.3281 | -0.2994 | 1.1503 | CONFORME | Medición
106 | 5 | PA25517706 | 0 | -0.41795 | -0.84704 | -0.61618 | CONFORME | Medición
107 | 5 | PA25517707 | 0 | -0.93542 | -1.0463 | -0.61618 | CONFORME | Medición
108 | 5 | PA25517708 | 0 | -0.46771 | -0.64773 | 0.37766 | CONFORME | Medición
109 | 5 | PA25517709 | 0 | -0.85581 | -0.84704 | -0.21864 | CONFORME | Medición
110 | 5 | PA25517710 | 0 | -1.433 | -0.94669 | 0.37766 | CONFORME | Medición
111 | 5 | PA25517711 | 0 | -1.3036 | -0.84704 | -0.61618 | CONFORME | Medición
112 | 5 | PA25517712 | 0 | -0.55727 | -0.84704 | -0.21864 | CONFORME | Medición
113 | 5 | PA25517713 | 0 | -1.0847 | -0.74738 | -0.019877 | CONFORME | Medición
114 | 5 | PA25517714 | 0 | -0.43785 | -1.0463 | 0.37766 | CONFORME | Medición
115 | 5 | PA25517715 | 0 | -0.23883 | -0.84704 | 0.37766 | CONFORME | Medición
116 | 5 | PA25517716 | 0 | -0.58712 | -0.74738 | 0.17889 | CONFORME | Medición
117 | 5 | PA25517717 | 0 | -0.72644 | -0.74738 | 0.77519 | CONFORME | Medición
118 | 5 | PA25517718 | 0 | -0.51746 | -0.84704 | 0.37766 | CONFORME | Medición
119 | 5 | PA25517719 | 0 | -1.234 | -1.0463 | 0.37766 | CONFORME | Medición
120 | 5 | PA25517720 | 0 | -0.28859 | -0.14948 | 0.97396 | CONFORME | Medición
121 | 5 | PA25517721 | 0 | -0.38714 | -0.48956 | 1.3742 | CONFORME | Medición
122 | 5 | PA25517722 | 0 | -1.082 | -1.089 | 0.57757 | CONFORME | Medición
123 | 5 | PA25517723 | 0 | -0.38714 | -0.089919 | 0.9759 | CONFORME | Medición
124 | 5 | PA25517724 | 0 | 0.45662 | 0.50954 | 2.1709 | CONFORME | Medición
125 | 5 | PA25517725 | 0 | -0.17868 | -0.58947 | 0.57757 | CONFORME | Medición
126 | 5 | PA25517726 | 0 | -0.72464 | -0.089919 | 0.57757 | CONFORME | Medición
127 | 5 | PA25517727 | 0 | -0.59559 | -0.58947 | 0.57757 | CONFORME | Medición
128 | 5 | PA25517728 | 0 | -0.30772 | -0.58947 | 0.37841 | CONFORME | Medición
129 | 5 | PA25517729 | 0 | -0.73456 | -0.089919 | 0.77674 | CONFORME | Medición
130 | 5 | PA25517730 | 0 | -0.24816 | -0.68938 | -0.21908 | CONFORME | Medición
131 | 5 | PA25517731 | 0 | -1.0324 | -0.78929 | 0.37841 | CONFORME | Medición
132 | 5 | PA25517732 | 0 | -0.58567 | -1.089 | 0.57757 | CONFORME | Medición
133 | 5 | PA25517733 | 0 | -0.43677 | -1.3888 | -0.61741 | CONFORME | Medición
134 | 5 | PA25517734 | 0 | -0.1886 | -0.089919 | 0.57757 | CONFORME | Medición
135 | 5 | PA25517735 | 0 | -0.37721 | 0.30972 | 1.5734 | CONFORME | Medición
136 | 5 | PA25517736 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
137 | 5 | PA25517737 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
138 | 5 | PA25517738 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
139 | 5 | PA25517739 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
140 | 5 | PA25517740 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
141 | 5 | PA25517741 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
142 | 5 | PA25517742 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
143 | 5 | PA25517743 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
144 | 5 | PA25517744 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
145 | 5 | PA25517745 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
146 | 5 | PA25517746 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
147 | 5 | PA25517747 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
148 | 5 | PA25517748 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
149 | 5 | PA25517749 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
150 | 5 | PA25517750 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
151 | 5 | PA25517751 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
152 | 5 | PA25517752 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
153 | 5 | PA25517753 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
154 | 5 | PA25517754 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
155 | 5 | PA25517755 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
156 | 5 | PA25517756 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
157 | 5 | PA25517757 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
158 | 5 | PA25517758 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
159 | 5 | PA25517759 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
160 | 5 | PA25517760 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
161 | 5 | PA25517761 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
162 | 5 | PA25517762 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
163 | 5 | PA25517763 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
164 | 5 | PA25517764 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
165 | 5 | PA25517765 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
166 | 5 | PA25517766 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
167 | 5 | PA25517767 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
168 | 5 | PA25517768 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
169 | 5 | PA25517769 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
170 | 5 | PA25517770 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
171 | 5 | PA25517771 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
172 | 5 | PA25517772 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
173 | 5 | PA25517773 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
174 | 5 | PA25517774 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
175 | 5 | PA25517775 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
176 | 5 | PA25517776 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
177 | 5 | PA25517777 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
178 | 5 | PA25517778 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
179 | 5 | PA25517779 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
180 | 5 | PA25517780 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
181 | 5 | PA25517781 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
182 | 5 | PA25517782 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
183 | 5 | PA25517783 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
184 | 5 | PA25517784 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
185 | 5 | PA25517785 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
186 | 5 | PA25517786 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
187 | 5 | PA25517787 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
188 | 5 | PA25517788 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
189 | 5 | PA25517789 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
190 | 5 | PA25517790 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
191 | 5 | PA25517791 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
192 | 5 | PA25517792 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
193 | 5 | PA25517793 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
194 | 5 | PA25517794 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
195 | 5 | PA25517795 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
196 | 5 | PA25517796 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
197 | 5 | PA25517797 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
198 | 5 | PA25517798 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
199 | 5 | PA25517799 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
200 | 5 | PA25517800 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
201 | 5 | PA25517801 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
202 | 5 | PA25517802 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
203 | 5 | PA25517803 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
204 | 5 | PA25517804 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
205 | 5 | PA25517805 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
206 | 5 | PA25517806 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
207 | 5 | PA25517807 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
208 | 5 | PA25517808 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
209 | 5 | PA25517809 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
210 | 5 | PA25517810 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
211 | 5 | PA25517811 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
212 | 5 | PA25517812 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
213 | 5 | PA25517813 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
214 | 5 | PA25517814 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
215 | 5 | PA25517815 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
216 | 5 | PA25517816 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
217 | 5 | PA25517817 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
218 | 5 | PA25517818 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
219 | 5 | PA25517819 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
220 | 5 | PA25517820 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
221 | 5 | PA25517821 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
222 | 5 | PA25517822 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
223 | 5 | PA25517823 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
224 | 5 | PA25517824 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
225 | 5 | PA25517825 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
226 | 5 | PA25517826 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
227 | 5 | PA25517827 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
228 | 5 | PA25517828 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
229 | 5 | PA25517829 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
230 | 5 | PA25517830 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
231 | 5 | PA25517831 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
232 | 5 | PA25517832 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
233 | 5 | PA25517833 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
234 | 5 | PA25517834 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
235 | 5 | PA25517835 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
236 | 5 | PA25517836 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
237 | 5 | PA25517837 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
238 | 5 | PA25517838 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
239 | 5 | PA25517839 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
240 | 5 | PA25517840 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
241 | 5 | PA25517841 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
242 | 5 | PA25517842 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
243 | 5 | PA25517843 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
244 | 5 | PA25517844 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
245 | 5 | PA25517845 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
246 | 5 | PA25517846 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
247 | 5 | PA25517847 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
248 | 5 | PA25517848 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
249 | 5 | PA25517849 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
250 | 5 | PA25517850 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
251 | 5 | PA25517851 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
252 | 5 | PA25517852 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
253 | 5 | PA25517853 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
254 | 5 | PA25517854 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
255 | 5 | PA25517855 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
256 | 5 | PA25517856 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
257 | 5 | PA25517857 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
258 | 5 | PA25517858 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
259 | 5 | PA25517859 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
260 | 5 | PA25517860 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
261 | 5 | PA25517861 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
262 | 5 | PA25517862 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
263 | 5 | PA25517863 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
264 | 5 | PA25517864 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
265 | 5 | PA25517865 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
266 | 5 | PA25517866 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
267 | 5 | PA25517867 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
268 | 5 | PA25517868 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
269 | 5 | PA25517869 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
270 | 5 | PA25517870 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
271 | 5 | PA25517871 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
272 | 5 | PA25517872 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
273 | 5 | PA25517873 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
274 | 5 | PA25517874 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
275 | 5 | PA25517875 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
276 | 5 | PA25517876 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
277 | 5 | PA25517877 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
278 | 5 | PA25517878 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
279 | 5 | PA25517879 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
280 | 5 | PA25517880 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
281 | 5 | PA25517881 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
282 | 5 | PA25517882 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
283 | 5 | PA25517883 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
284 | 5 | PA25517884 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
285 | 5 | PA25517885 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
286 | 5 | PA25517886 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
287 | 5 | PA25517887 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
288 | 5 | PA25517888 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
289 | 5 | PA25517889 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
290 | 5 | PA25517890 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
291 | 5 | PA25517891 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
292 | 5 | PA25517892 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
293 | 5 | PA25517893 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
294 | 5 | PA25517894 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
295 | 5 | PA25517895 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
296 | 5 | PA25517896 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
297 | 5 | PA25517897 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
298 | 5 | PA25517898 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
299 | 5 | PA25517899 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
300 | 5 | PA25517900 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Medición
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
#REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Hoja2
|  |  |  | 0 | 0
Item | N° de banco de ensayo | # Medidor | Presión Estática (Estado) | Q3 E % | Q2 E % | Q1 E % | Estado | Condición